BED-HEAD UNITS & MEDICAL GAS OUTLETS

Technical Tender Document

Table of Contents

30. ICU-GRADE BED-HEAD UNIT	4

30.1 Clinical & Functional Specifications	4

30.2 Engineering Technical Specifications	4

30.3 Standard Accessories	4

30.4 Installation Requirements	5

30.5 Documentation includes gas-pipeline test certificates, electrical test certificates, as-built drawings; Training on gas-outlet cleaning and terminal replacement; Warranty ≥ 2 years.	5

31. HDU-GRADE BED-HEAD UNIT	6

31.1 Clinical & Functional Specifications	6

31.2 Engineering Technical Specifications	6

31.3 Standard Accessories	6

31.4 As per ICU BHU clauses; Warranty ≥ 2 years.	6

32. EMERGENCY-BAY BED-HEAD UNIT (HEAVY-DUTY)	7

32.1 Clinical & Functional Specifications	7

32.2 Engineering Technical Specifications	7

32.3 Standard Accessories	7

32.4 As per ICU BHU clauses; Warranty ≥ 2 years.	7

33. WARD-GRADE BED-HEAD UNIT	8

33.1 Clinical & Functional Specifications	8

33.2 Engineering Technical Specifications	8

33.3 Standard Accessories	8

33.4 Warranty ≥ 2 years.	8

34. PAEDIATRIC BED-HEAD UNIT	9

34.1 Clinical & Functional Specifications	9

34.2 Engineering Technical Specifications	9

34.3 Standard Accessories	9

34.4 Warranty ≥ 2 years.	9

35. OXYGEN FLOWMETER WITH HUMIDIFIER BOTTLE	10

35.1 Clinical & Functional Specifications	10

35.2 Engineering Technical Specifications	10

35.3 Standard Accessories	10

35.4 Warranty ≥ 2 years.	10

36. VACUUM REGULATOR WITH COLLECTION JAR	11

36.1 Clinical & Functional Specifications	11

36.2 Engineering Technical Specifications	11

36.3 Standard Accessories	11

36.4 Warranty ≥ 2 years.	11

37. NURSE-CALL SYSTEM	12

37.1 Clinical & Functional Specifications	12

37.2 Engineering Technical Specifications	12

37.3 Standard Accessories	12

37.4 Installation Requirements	13

37.5 Documentation	13

37.6 Training	13

37.7 Warranty	13

# 30. ICU-GRADE BED-HEAD UNIT

Quantity: 47

## 30.1 Clinical & Functional Specifications

Horizontal architectural bedhead console for ICU/CCU patient bays providing centralized delivery of medical gases, electrical power, low-voltage communication, and equipment mounting rails. Integrated three-channel service trunking with electrically and medically isolated service compartments.

## 30.2 Engineering Technical Specifications

- Housing: Extruded architectural aluminum alloy, powder-coated / anodized; antibacterial finish.

- Length: Minimum 1.8 m per bed.

- Height: 200—300 mm; wall-mounted at standardized height per hospital guideline.

- Compartments: Minimum three isolated channels (medical gas, low-voltage/data, mains electrical) with earthed metallic separation.

- Medical gas outlets: 2 × O₂, 2 × Medical Air (4 bar), 2 × Vacuum, 1 × AGSS; per HTM 02-01 / DIN / BS / NFPA compatible terminal system (specify country standard at PO).

- Electrical outlets: Minimum 8 universal 220—240 VAC, 16 A hospital-grade sockets with individual MCB/RCBO protection.

- Data ports: 2 × RJ45 Cat 6, 1 × Nurse-call interface.

- Equipment rails: Two continuous DIN-25 or 10 × 30 mm equipment rails (top and bottom).

- Lighting: Optional integrated indirect and examination LED lighting.

- Earthing: Equipotential bonding studs (≥ 4 per unit).

- Standards: IEC 60601-1, IEC 60601-2-52, ISO 7396-1 (medical gas pipelines), HTM 02-01 / NFPA 99, CE/FDA.

## 30.3 Standard Accessories

- Extruded architectural aluminum alloy profile casing trunking (minimum 1.8-meter length) — Extruded aluminum multi-channel trunking, powder-coated, antibacterial, ≥ 1.8 m.

- Medical gas outlet terminals (2× O₂, 2× Medical Air, 2× Vacuum, 1× AGSS) — Terminal units per specified national standard (DIN/BS/NIST), color-coded, quick-connect self-sealing, ISO 7396-1 compliant.

- Electrical power grid sockets (minimum 8 universal electrical outlets) — Eight hospital-grade universal 220—240 VAC 16 A sockets with individual protection.

- Low-voltage data port terminals (2× RJ45 network jacks, 1× Nurse call interface jack) — Two Cat 6 RJ45 jacks + one nurse-call jack, isolated compartment.

- Integrated equipment accessory mounting rail lines (top and bottom tracks) — Two continuous DIN-25 equipment rails for monitors, syringe pumps, etc.

- Separated isolated internal service containment channels for gas pipelines and electrical wires — Metal partitioned earthed compartments per IEC 60601-1 and ISO 7396-1.

## 30.4 Installation Requirements

Wall-mounted at standard bed-head height; medical gas connections from central pipeline system tested per ISO 7396-1; dedicated MCB per bed; earthing bonded to hospital equipotential bar; commissioning includes gas identity, pressure, purity, and flow test.

## 30.5 Documentation includes gas-pipeline test certificates, electrical test certificates, as-built drawings; Training on gas-outlet cleaning and terminal replacement; Warranty ≥ 2 years.

# 31. HDU-GRADE BED-HEAD UNIT

Quantity: 18

## 31.1 Clinical & Functional Specifications

Surface-mounted horizontal bedhead console for high-dependency wards, providing gases, power, and communication in a slimmer profile than the ICU unit.

## 31.2 Engineering Technical Specifications

- Housing: Anodized aluminum, powder-coated antibacterial finish.

- Length: Minimum 1.5 m per bed.

- Compartments: Minimum two isolated channels.

- Gases: 2 × O₂, 1 × Medical Air, 1 × Vacuum.

- Electrical outlets: ≥ 6 hospital-grade 220—240 VAC sockets.

- Data ports: 1 × RJ45, 1 × Nurse call.

- Lighting: Integrated indirect ambient + direct reading LED with electronic ballast/driver.

- Standards: IEC 60601-1, ISO 7396-1, HTM 02-01/NFPA 99, CE.

## 31.3 Standard Accessories

- Surface-mount anodized structural aluminum horizontal channel console (1.5-meter length) — Anodized aluminum, ≥ 1.5 m.

- Quick-connect medical gas outlet terminal valves (2× O₂, 1× Medical Air, 1× Vacuum) — Self-sealing terminals per specified standard.

- Universal electrical connection ports (minimum 6 power grid sockets) — Six hospital-grade universal sockets, individually protected.

- Data communication network interfaces (1× RJ45 jack, 1× Nurse call jack connection) — One RJ45 Cat 6 + one nurse-call jack.

- Integrated indirect background lighting fixture module with electronic ballast switch — LED indirect ambient + direct reading light, electronic driver, bed-side switch.

## 31.4 As per ICU BHU clauses; Warranty ≥ 2 years.

# 32. EMERGENCY-BAY BED-HEAD UNIT (HEAVY-DUTY)

Quantity: 10

## 32.1 Clinical & Functional Specifications

Impact-resistant heavy-duty bedhead console for emergency department resuscitation bays, providing high-flow medical gases and enhanced electrical capacity for concurrent life-support equipment.

## 32.2 Engineering Technical Specifications

- Housing: Impact-resistant reinforced aluminum construction.

- Gases: 2 × O₂ (high-flow), 2 × Vacuum, 1 × Medical Air at 4 bar.

- Electrical outlets: ≥ 8 hospital-grade sockets, individually protected.

- Equipotential studs: Minimum 4 grounding pins.

- Equipment mounting arm interface: Heavy-duty rail/arm interface for IV pole and monitor mounting.

- Standards: IEC 60601-1, ISO 7396-1, HTM 02-01/NFPA 99, CE.

## 32.3 Standard Accessories

- Impact-resistant reinforced structural aluminum equipment console trunking — Heavy-gauge aluminum trunking, impact-tested finish.

- High-flow medical gas terminal blocks (2× O₂, 2× Vacuum, 1× Medical Air 4-bar) — High-flow ISO 7396-1 compliant terminals.

- Heavy-duty universal electrical socket arrays (minimum 8 power sockets) — Eight hospital-grade sockets, individual MCB.

- Grounding connection potential equalization pin studs (set of 4 pins) — Four equipotential bonding studs per IEC 60601-1.

- Heavy-duty integrated IV pole or monitoring mounting bracket arm interface — Reinforced bracket/rail for IV pole and monitor arm mounting.

## 32.4 As per ICU BHU clauses; Warranty ≥ 2 years.

# 33. WARD-GRADE BED-HEAD UNIT

Quantity: 252

## 33.1 Clinical & Functional Specifications

Slimline horizontal bedhead unit for general wards, providing basic medical gas outlets, electrical sockets, and integrated lighting.

## 33.2 Engineering Technical Specifications

- Housing: Extruded aluminum, powder-coated, ≥ 1.2 m length.

- Gases: 1 × O₂, 1 × Vacuum.

- Electrical outlets: ≥ 4 hospital-grade 220—240 VAC sockets.

- Nurse call port.

- Lighting: Integrated twin T5 or LED direct reading + indirect ambient light.

- Standards: IEC 60601-1, ISO 7396-1, HTM 02-01/NFPA 99, CE.

## 33.3 Standard Accessories

- Slimline surface-mounted architectural horizontal aluminum trunking layout (1.2-meter length) — Slimline aluminum, powder-coated.

- Standard medical gas outlet pipeline terminals (1× O₂, 1× Vacuum) — ISO 7396-1 terminals.

- Universal power grid electrical outlets (minimum 4 sockets) — Four hospital-grade sockets, individually protected.

- Nurse-call system connector port panel plug interface — Nurse-call jack per hospital system.

- Integrated twin T5/LED direct reading and indirect room light fixtures — Twin LED/T5 direct reading + indirect ambient, bedside switch.

## 33.4 Warranty ≥ 2 years.

# 34. PAEDIATRIC BED-HEAD UNIT

Quantity: 10

## 34.1 Clinical & Functional Specifications

Bedhead unit for paediatric ward beds providing basic medical gases, tamper-resistant electrical outlets, and child-safe finish.

## 34.2 Engineering Technical Specifications

- Housing: Flush-mount aluminum profile with child-friendly decorative panel.

- Gases: 1 × O₂, 1 × Vacuum.

- Electrical outlets: ≥ 4 tamper-resistant shuttered sockets.

- Emergency call jack.

- Night light: Soft-LED module with low-lux ambient illumination.

- Standards: IEC 60601-1, ISO 7396-1, HTM 02-01/NFPA 99, CE.

## 34.3 Standard Accessories

- Flush-mount horizontal aluminum profile channel panel with child-friendly exterior decoration cover — Aluminum profile with decorative child-safe overlay.

- Standard safety medical gas terminal outlets (1× O₂, 1× Vacuum) — ISO 7396-1 terminals.

- Child-safe tamper-resistant internal shuttered electrical power sockets (minimum 4 sockets) — Four tamper-resistant shuttered sockets, hospital-grade.

- Integrated emergency bedside call light socket jack port interface — Emergency call jack.

- Soft LED night light illumination module component fixture — Low-lux LED night light, bedside switched.

## 34.4 Warranty ≥ 2 years.

# 35. OXYGEN FLOWMETER WITH HUMIDIFIER BOTTLE

Quantity: 337

## 35.1 Clinical & Functional Specifications

Wall/terminal-mounted oxygen therapy device providing precise flow-controlled delivery of humidified medical oxygen to the patient. Compatible with hospital medical gas pipeline terminal.

## 35.2 Engineering Technical Specifications

- Body: Anodized brass or chromed alloy regulator body.

- Flow scale: Variable-area glass tube (Thorpe tube), 0—15 L/min, resolution ≤ 0.5 L/min at low flows.

- Accuracy: ± 10% of indicated flow or better.

- Inlet: Quick-connect probe compatible with specified national standard terminal (DIN/BS/NIST/Ohmeda — to be confirmed at PO).

- Outlet: Standard 9/16" fir-tree nipple for oxygen tubing.

- Humidifier bottle: Autoclavable polycarbonate, 250—300 mL, with microporous diffuser stem, non-drip check valve.

- Working pressure: 4 bar (400 kPa) nominal.

- Standards: ISO 15002, ISO 7396-1 (terminal compatibility), CE.

## 35.3 Standard Accessories

- Anodized brass flowmeter regulating body with vertical variable-area glass tube scale (0—15 L/min) — Brass/chromed body, Thorpe glass tube 0—15 L/min.

- Quick-connect probe adapter plug (matching specific country-standard terminal type) — Probe per specified national standard (to be nominated in PO).

- Reusable autoclavable clear polycarbonate humidifier bottle container jar (250 mL to 300 mL volume) — Autoclavable PC bottle 250—300 mL.

- Internal microporous gas diffuser bubble stem tube assembly — Microporous diffuser stem.

- Replacement silicone sealing O-ring washer kit (2 sets) — Two spare O-ring kits per unit.

## 35.4 Warranty ≥ 2 years.

# 36. VACUUM REGULATOR WITH COLLECTION JAR

Quantity: 75

## 36.1 Clinical & Functional Specifications

Wall/terminal-mounted suction regulator with collection jar for controlled patient suction (airway, wound, surgical drainage).

## 36.2 Engineering Technical Specifications

- Body: Chromed/anodized brass regulator.

- Vacuum gauge: Analog, 0 to —760 mmHg (or 0 to —100 kPa).

- Regulation accuracy: ± 5% of setting.

- Selector: Three-position mode (REG / OFF / FULL).

- Collection jar: Autoclavable polycarbonate, 1 L or 2 L, with overflow safety float and hydrophobic bacterial filter.

- Inlet: Quick-connect probe per specified national vacuum terminal standard.

- Standards: ISO 10079-3, ISO 7396-1, CE.

## 36.3 Standard Accessories

- Precision negative pressure regulator control head with large analog suction gauge scale dial — Precision regulator with ≥ 50 mm analog gauge, 0 to —760 mmHg.

- Three-mode selection control selector valve knob (REG / OFF / FULL) — Three-position selector.

- Reusable autoclavable clear polycarbonate secretion safety collection bottle (1-liter or 2-liter jar) — Autoclavable PC jar 1 or 2 L.

- Integrated mechanical overflow protection safety check float valve mechanism — Overflow float cutoff.

- Hydrophobic anti-bacterial inline protective filter cartridge element — Hydrophobic bacterial/viral filter.

## 36.4 Warranty ≥ 2 years.

# 37. NURSE-CALL SYSTEM

Quantity: 12

## 37.1 Clinical & Functional Specifications

Wired or IP-based nurse-call communication system enabling patients to signal for nursing assistance from the bedside or bathroom. Central master station displays call origin, priority, and status. Complies with standard hospital nurse-call functional and safety requirements.

## 37.2 Engineering Technical Specifications

- Architecture: Wired or IP-based, addressable, expandable per ward.

- Central station: Digital LCD/LED display of bed and room location, call type, priority, timestamp; audio intercom.

- Bedside stations: Wall-mount speaker/microphone with call and cancel buttons, reassurance LED.

- Patient pendant: Waterproof (IPX4) handheld pendant with clip, corded connection.

- Corridor lamps: Over-door tri-colour LED dome indicating call category (nurse call / assistance / emergency).

- Emergency pull-cords: Waterproof pull-cord stations for bathroom/toilet, IP-rated, height compliant with accessibility standard.

- Priority levels: Minimum three (normal call, staff assist, code/emergency).

- Reset: Only at bedside station (per DIN VDE 0834 / equivalent).

- Logging: Event log with time-stamped audit trail; report export.

- Power: Central UPS-supplied 220—240 VAC; low-voltage bedside wiring.

- Standards: DIN VDE 0834 or UL 1069 / equivalent, IEC 60601-1 (where applicable to medical use), CE.

## 37.3 Standard Accessories

- Central station master control console dashboard unit with digital room bed grid display panel — Master console with digital display, audio, event logging, printer/report export.

- Bedside wall-mount speaker microphone intercom station modules (matching full bed distribution layout) — Wall-mount speaker/microphone station per bed, with reassurance LED.

- Waterproof handheld flexible cord patient calling call button pendants with bed clips — IPX4 handheld pendants with call/cancel button and bed-clip.

- Over-door three-color LED dome corridor status alarm indicator lights lamps — Tri-colour LED corridor dome lights.

- Waterproof pull-cord emergency calling stations for bathroom/toilet wall mounting — IP-rated waterproof pull-cord stations for wet areas.

## 37.4 Installation Requirements

Structured cabling per hospital IT/BMS standard; central master station at nursing station; UPS-backed power supply; commissioning and integration test; interface with hospital paging/mobile-DECT/pager system where applicable.

## 37.5 Documentation

User manual, service manual, wiring diagrams, IP/network configuration, IT/network commissioning report, electrical safety certificate, CE certificate.

## 37.6 Training

User training for nursing and administrative staff; biomedical/IT training on configuration, addressing, event log management, and troubleshooting.

## 37.7 Warranty

Comprehensive 2-year warranty covering hardware, software updates, preventive maintenance, technical support, spare parts, and on-site service. Guaranteed spare-parts availability ≥ 10 years.