TECHNICAL SPECIFICATIONS — BEDS, FURNITURE

Technical Tender Document

Table of Contents

1. ICU BED — ELECTRIC, 5 MOTORS	7

1.1 Clinical & Functional Specifications	7

1.2 Engineering Technical Specifications	7

1.3 Standard Accessories	8

1.4 Installation Requirements	8

1.5 Documentation	9

1.6 Training	9

1.7 Warranty	9

2. ICU BED — ELECTRIC, 3 MOTORS	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	10

2.4 Installation Requirements	11

2.5 Documentation / 6. Training / 7. Warranty	11

3. HDU BED — ELECTRIC, MULTI-FUNCTION	12

3.1 Clinical & Functional Specifications	12

3.2 Engineering Technical Specifications	12

3.3 Standard Accessories	12

3.4 As per general bed clauses above.	13

4. PACU RECOVERY BED / TROLLEY	14

4.1 Clinical & Functional Specifications	14

4.2 Engineering Technical Specifications	14

4.3 Standard Accessories	14

4.4 As per general bed clauses.	15

5. WARD BED — MANUAL, 2-CRANK	16

5.1 Clinical & Functional Specifications	16

5.2 Engineering Technical Specifications	16

5.3 Standard Accessories	16

5.4 Installation: standard ward space; Documentation, Training, Warranty as general.	16

6. PAEDIATRIC WARD BED	17

6.1 Clinical & Functional Specifications	17

6.2 Engineering Technical Specifications	17

6.3 Standard Accessories	17

6.4 As per general bed clauses; paediatric-appropriate training.	17

7. CONSULTANCY / PRIVATE ROOM BED — ELECTRIC	18

7.1 Clinical & Functional Specifications	18

7.2 Engineering Technical Specifications	18

7.3 Standard Accessories	18

7.4 As per general bed clauses.	19

8. BEDSIDE LOCKER WITH LOCKABLE DRAWER	20

8.1 Clinical & Functional Specifications	20

8.2 Engineering Technical Specifications	20

8.3 Standard Accessories	20

8.4 Installation: standard ward space; Documentation, Training, Warranty as general (Warranty ≥ 2 years).	20

9. OVERBED TABLE	21

9.1 Clinical & Functional Specifications	21

9.2 Engineering Technical Specifications	21

9.3 Standard Accessories	21

9.4 Installation: standard; Warranty ≥ 2 years.	21

10. IV DRIP STAND / POLE (MOBILE)	22

10.1 Clinical & Functional Specifications	22

10.2 Engineering Technical Specifications	22

10.3 Standard Accessories	22

10.4 Warranty ≥ 2 years.	22

11. WHEELCHAIR (ADULT)	23

11.1 Clinical & Functional Specifications	23

11.2 Engineering Technical Specifications	23

11.3 Standard Accessories	23

11.4 Warranty ≥ 2 years.	23

12. WHEELCHAIR (PAEDIATRIC)	24

12.1 Clinical & Functional Specifications	24

12.2 Engineering Technical Specifications	24

12.3 Standard Accessories	24

12.4 Warranty ≥ 2 years.	24

13. PATIENT TROLLEY FOR WARD USE	25

13.1 Clinical & Functional Specifications	25

13.2 Engineering Technical Specifications	25

13.3 Standard Accessories	25

13.4 Warranty ≥ 2 years.	25

14. SCREENS / CURTAIN RAILS (BED PRIVACY)	26

14.1 Clinical & Functional Specifications	26

14.2 Engineering Technical Specifications	26

14.3 Standard Accessories	26

14.4 Installation to ceiling structure; Warranty ≥ 2 years.	26

15. MEDICINE TROLLEY (NURSE)	27

15.1 Clinical & Functional Specifications	27

15.2 Engineering Technical Specifications	27

15.3 Standard Accessories	27

15.4 Warranty ≥ 2 years.	27

16. DRESSING TROLLEY (SS, DOUBLE SHELF)	28

16.1 Clinical & Functional Specifications	28

16.2 Engineering Technical Specifications	28

16.3 Standard Accessories	28

16.4 Warranty ≥ 2 years.	28

17. NURSE STATION DESK / COUNTER (MODULAR)	29

17.1 Clinical & Functional Specifications	29

17.2 Engineering Technical Specifications	29

17.3 Standard Accessories	29

17.4 Warranty ≥ 2 years.	29

18. AIR MATTRESS WITH PUMP	30

18.1 Clinical & Functional Specifications	30

18.2 Engineering Technical Specifications	30

18.3 Standard Accessories	30

18.4 Warranty ≥ 2 years.	30

19. CATHETER STORAGE CABINET	31

19.1 Clinical & Functional Specifications	31

19.2 Engineering Technical Specifications	31

19.3 Standard Accessories	31

19.4 Warranty ≥ 2 years.	31

20. INSTRUMENT CABINET	32

20.1 Clinical & Functional Specifications	32

20.2 Engineering Technical Specifications	32

20.3 Standard Accessories	32

20.4 Warranty ≥ 2 years.	32

21. INSTRUMENT TROLLEY (SS, DOUBLE SHELF)	33

21.1 Clinical & Functional Specifications	33

21.2 Engineering Technical Specifications	33

21.3 Standard Accessories	33

21.4 Warranty ≥ 2 years.	33

22. UTILITY / DRESSING TROLLEY (SS)	34

22.1 Clinical & Functional Specifications	34

22.2 Engineering Technical Specifications	34

22.3 Standard Accessories	34

22.4 Warranty ≥ 2 years.	34

23. SS STORAGE CABINET (LOCKABLE)	35

23.1 Clinical & Functional Specifications	35

23.2 Engineering Technical Specifications	35

23.3 Standard Accessories	35

23.4 Warranty ≥ 2 years.	35

24. MAYO STAND	36

24.1 Clinical & Functional Specifications	36

24.2 Engineering Technical Specifications	36

24.3 Standard Accessories	36

24.4 Warranty ≥ 2 years.	36

25. KICK BUCKET	37

25.1 Clinical & Functional Specifications	37

25.2 Engineering Technical Specifications	37

25.3 Standard Accessories	37

25.4 Warranty ≥ 2 years.	37

26. LINEN HAMPER	38

26.1 Clinical & Functional Specifications	38

26.2 Engineering Technical Specifications	38

26.3 Standard Accessories	38

26.4 Warranty ≥ 2 years.	38

27. FOOT STEP	39

27.1 Clinical & Functional Specifications	39

27.2 Engineering Technical Specifications	39

27.3 Standard Accessories	39

27.4 Warranty ≥ 2 years.	39

28. MEDICAL / MEDICINE REFRIGERATOR (2—8 °C)	40

28.1 Clinical & Functional Specifications	40

28.2 Engineering Technical Specifications	40

28.3 Standard Accessories	40

28.4 Installation Requirements	41

28.5 Documentation, Training, Warranty as general (Warranty ≥ 2 years, comprehensive).	41

29. AUTOCLAVE (24 L)	42

29.1 Clinical & Functional Specifications	42

29.2 Engineering Technical Specifications	42

29.3 Standard Accessories	42

29.4 Installation Requirements	42

# 1. ICU BED — ELECTRIC, 5 MOTORS

Quantity: 15

## 1.1 Clinical & Functional Specifications

Fully electrically operated intensive-care bed intended for critically ill adult patients requiring continuous monitoring, ventilation, hemodynamic support, and prolonged bed rest. The bed shall permit safe patient positioning for all standard critical-care therapies including mechanical ventilation, hemodialysis, parenteral nutrition, invasive monitoring, and pressure-injury prevention. It shall support routine nursing procedures (bathing, linen change, physiotherapy, X-ray imaging), emergency interventions (CPR, defibrillation, intubation), and safe intra-hospital patient transfer without disconnecting the patient from life-support devices. The bed shall provide independently motorized adjustment of at least five functions — backrest, thigh/knee section, height, Trendelenburg, and reverse-Trendelenburg — through a minimum of four independent DC actuators, controlled through nurse control panel, integrated side-rail controls, and a handheld patient pendant.

## 1.2 Engineering Technical Specifications

- Configuration: Four-section radiolucent patient platform (head, seat, thigh, calf), motorized, with five operating modes minimum (backrest, knee/thigh, height, Trendelenburg, reverse-Trendelenburg; auto-contour, chair, examination, and vascular positions selectable).

- Frame construction: Electro-galvanized or powder-coated steel main frame with ABS or polymer end panels; corrosion-resistant, wipe-cleanable finish.

- Patient surface: Minimum four-section radiotranslucent platform suitable for on-bed C-arm and mobile X-ray imaging; integrated cassette holder or X-ray tunnel.

- Overall dimensions (approx.): Length 2100—2250 mm; Width 950—1050 mm; suitable for standard ICU cubicle.

- Sleeping surface (approx.): 2000 × 900 mm (±50 mm).

- Height adjustment range: Minimum 400—800 mm (floor to platform), electrically adjustable.

- Backrest angle: 0° to at least 70°.

- Knee/thigh angle: 0° to at least 30°.

- Trendelenburg / reverse-Trendelenburg: Minimum ±12°.

- CPR release: Mechanical, gas-spring assisted, single-action, dual-sided, ≤ 3 seconds to horizontal.

- Safe working load: Minimum 250 kg (patient + accessories).

- Side rails: Split polymer four-section bilateral rails, retractable/foldable, with integrated indicators for bed angle, CPR, brake status, and control lockout.

- Wheels: Minimum 150 mm diameter, twin-wheel, anti-static, central braking system with directional lock (fifth wheel or equivalent for steering).

- Integrated weighing system: Digital load-cell scale, capacity ≥ 250 kg, resolution ≤ 0.5 kg, accuracy ±1% of reading, zero/tare/hold functions, LCD readout.

- Controls: Nurse control panel at footboard (backlit, membrane, fluid-resistant); dual side-rail controls; handheld patient pendant with lockout; audible/visual position and battery indicators.

- Power supply: 220—240 VAC, 50/60 Hz.

- Battery backup: Sealed maintenance-free rechargeable battery; minimum 30 full motor operations or ≥ 2 hours standby on full charge.

- Motors: Silent DC actuators, IP54 or better, with overload/thermal protection and anti-crush safety.

- Ingress protection: Minimum IPX4 for control and drive units.

- Environmental: Operating 10—40 °C, 20—90% RH non-condensing.

- Applicable standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-52 (medical beds), ISO 13485, CE MDR / US FDA equivalent.

## 1.3 Standard Accessories

- High-density memory foam mattress with vapor-permeable, anti-static, waterproof cover — Cold-cured HR/viscoelastic foam, density ≥ 45 kg/m³, thickness 120—150 mm, multi-zone construction; cover bi-elastic polyurethane on polyester, waterproof, vapor-permeable (MVTR ≥ 500 g/m²/24h), antimicrobial, antistatic, flame-retardant (BS 7175/CRIB 5 or equivalent), 360° welded/zippered seam, compatible with common hospital disinfectants.

- Integrated nurse control panel embedded in footboard and dual split side rails — Backlit membrane keypad, fluid-resistant (IPX4), function lockout keys for each motor, position indicators, brake/CPR/battery LEDs, symbols per IEC 60601-1 iconography.

- Dual mechanical and gas-spring assisted CPR levers on both sides of head end — Colour-coded (red), single-action release, gas-assisted return, backrest to horizontal in ≤ 3 seconds without power.

- Digital weigh-scale system with LCD interface display screen — 4-load-cell system, capacity ≥ 250 kg, resolution ≤ 0.5 kg, accuracy ±1%, zero/tare/hold/freeze/patient-weight-lock functions.

- Emergency backup battery system for fully functional remote operations — Sealed lead-acid or Li-ion, ≥ 24 V, auto-recharge, ≥ 2 hours standby / ≥ 30 full-cycle motor operations, low-battery indicator.

- Heavy-duty central locking system with multi-axis steering caster brake pedals — Minimum 150 mm anti-static twin-wheel casters, central 3-position pedal (brake/neutral/steer), foot-operable from both sides.

- Telescopic dual-hook IV pole with storage slot attachments — 304 stainless steel, height-adjustable 1200—2000 mm, load ≥ 10 kg, insertable at all four corners.

- Wall-protection roller bumpers mounted on all four corners of the bed frame — Polymer roller bumpers on all four corners, protecting bed, wall, and equipment.

## 1.4 Installation Requirements

Standard single-phase 220—240 VAC, 50/60 Hz supply with earthed 3-pin socket per bed; UPS-compatible; equipotential bonding stud to hospital earth bus; minimum floor loading capacity 500 kg/m²; ICU cubicle clearance ≥ 3.6 × 3.0 m; ambient 18—26 °C, 30—70% RH.

## 1.5 Documentation

User manual, service manual, spare-parts catalogue, wiring/hydraulic diagrams, electrical safety test certificate (IEC 62353), factory test certificate, CE/FDA certificate of conformity, calibration certificate for weigh-scale, preventive-maintenance schedule, installation and commissioning report — all in English, one hard + one soft copy per bed.

## 1.6 Training

On-site clinical user training for nursing/ICU staff; biomedical engineering training on scheduled and breakdown maintenance; troubleshooting; battery replacement; scale calibration; minimum two sessions.

## 1.7 Warranty

Comprehensive 2-year warranty covering labour, parts, preventive maintenance visits (minimum two per year), breakdown maintenance, software updates, and technical support. Guaranteed spare-parts and after-sales support ≥ 10 years.

# 2. ICU BED — ELECTRIC, 3 MOTORS

Quantity: 12

## 2.1 Clinical & Functional Specifications

Electric intensive-care bed for adult critically ill patients requiring monitoring and moderate life-support. Provides motorized adjustment of at least three functions — backrest, thigh/knee, and height — through independent DC actuators, with mechanical Trendelenburg where applicable. Suitable for prolonged bed rest, ventilation support, and routine ICU nursing procedures.

## 2.2 Engineering Technical Specifications

- Configuration: Four-section platform, three independent motors (backrest, knee, height); mechanical Trendelenburg/reverse-Trendelenburg by lever ± 12°.

- Frame: Powder-coated or galvanized steel; ABS/HPL end panels.

- Overall dimensions: Approx. 2100 × 1000 mm.

- Sleeping surface: Approx. 2000 × 900 mm.

- Height range: Minimum 400—780 mm.

- Backrest: 0° to ≥ 70°.

- Knee angle: 0° to ≥ 30°.

- Safe working load: ≥ 230 kg.

- CPR release: Mechanical, dual-sided, single-action.

- Wheels: ≥ 125 mm anti-static twin casters, central braking.

- Controls: Wired hand pendant + integrated side-rail/footboard controls with lockout.

- Power: 220—240 VAC, 50/60 Hz; battery backup ≥ 2 hours.

- Standards: IEC 60601-1, 60601-1-2, 60601-2-52, ISO 13485, CE/FDA.

## 2.3 Standard Accessories

- Pressure-redistribution mattress with fluid-proof, antimicrobial zippered cover — HR foam ≥ 40 kg/m³, thickness 120—140 mm, multi-zone; PU cover waterproof, vapor-permeable, antimicrobial, antistatic, flame-retardant, 360° zippered.

- Wired patient hand pendant controller and backrest angle indicator gauges — Fluid-resistant (IPX4) pendant with lockable functions, cable ≥ 2 m; mechanical/electronic backrest angle indicator visible to nurse.

- Bilateral mechanical emergency CPR quick-release levers at the headboard — Red-coded, single-hand operation, ≤ 3 s to horizontal.

- Integrated battery backup pack for motorized emergency operations — Sealed rechargeable ≥ 24 V, low-battery LED, ≥ 30 motor cycles.

- Central braking steering system operating on all caster wheels — Central pedal accessible both sides, brake/neutral/steer positions.

- Height-adjustable IV pole with dual suspension hooks — SS 304, 1200—2000 mm telescopic, load ≥ 8 kg.

- Protective rubber bumper disks on the perimeter chassis corner points — Polymer bumpers on all four corners.

## 2.4 Installation Requirements

As per ICU bed 5-mode.

## 2.5 Documentation / 6. Training / 7. Warranty

As per ICU bed 5-mode.

# 3. HDU BED — ELECTRIC, MULTI-FUNCTION

Quantity: 10

## 3.1 Clinical & Functional Specifications

Electrically operated high-dependency care bed for step-down patients requiring intensive monitoring but not full ICU life-support. Provides motorized adjustment of backrest, knee, and height, with cardiac chair and examination positioning capability. Supports patient mobilization, physiotherapy, and prolonged bed rest.

## 3.2 Engineering Technical Specifications

- Configuration: Four-section platform, minimum three motors, cardiac chair and examination positions.

- Frame: Powder-coated steel; ABS end panels.

- Overall dimensions: ≈ 2100 × 970 mm.

- Height range: 400—780 mm.

- Backrest: 0° to ≥ 65°; Knee: 0° to ≥ 30°.

- Trendelenburg/reverse: Mechanical, ± 12° minimum.

- Safe working load: ≥ 220 kg.

- CPR release: Mechanical, both sides.

- Wheels: ≥ 125 mm individual or central-locking casters, anti-static.

- Controls: Handheld pendant with function lockout.

- Power: 220—240 VAC; battery backup ≥ 1 hour.

- Standards: IEC 60601-1, 60601-2-52, CE/FDA.

## 3.3 Standard Accessories

- High-density foam mattress with wipe-clean, antistatic polyurethane cover — HR foam ≥ 40 kg/m³, thickness 100—130 mm; PU cover waterproof, antimicrobial, antistatic, flame-retardant.

- Ergonomic split side rails with integrated manual drop locks — Polymer/aluminum, split 3- or 4-section bilateral, drop or retract mechanism, patient-safe locking.

- Handheld push-button control pendant for adjustment monitoring — Fluid-resistant, illuminated keys, function lockout.

- Rechargeable auxiliary battery backup power unit assembly — Sealed rechargeable, ≥ 24 V, ≥ 1 hour standby.

- Individual locking caster wheels with anti-static rubber tread wraps — ≥ 125 mm, four wheels, individual brakes, anti-static.

- Removable ABS plastic headboards and footboards panels — Impact-resistant ABS, wipe-cleanable, quick-release.

- Telescopic IV drip pole with secure frame mounting clamps — SS 304, height-adjustable, load ≥ 8 kg.

## 3.4 As per general bed clauses above.

# 4. PACU RECOVERY BED / TROLLEY

Quantity: 8

## 4.1 Clinical & Functional Specifications

Height-adjustable mobile recovery trolley for post-anesthesia care unit use, permitting patient transfer from OR to recovery to ward without transfer. Supports Trendelenburg, reverse-Trendelenburg, backrest elevation, and rapid CPR flat positioning. Hydraulically operated for reliability and mobility. Suitable for adult and adolescent patients recovering from general or regional anesthesia.

## 4.2 Engineering Technical Specifications

- Configuration: Two- or three-section radiolucent platform, hydraulic elevation and Trendelenburg.

- Frame: Epoxy-coated steel with SS or aluminum accents; radiotranslucent top.

- Overall length: 2050—2200 mm; width 700—800 mm.

- Height range: 550—900 mm (hydraulic).

- Backrest: Gas-spring assisted, 0° to ≥ 75°.

- Trendelenburg / reverse: ± 15° minimum.

- Safe working load: ≥ 200 kg.

- Wheels: ≥ 200 mm, four twin-wheel casters, central brake, fifth-wheel steering.

- Oxygen cylinder holder: Under-frame cradle for E-size cylinder.

- Standards: IEC 60601-1, 60601-2-52, CE/FDA.

## 4.3 Standard Accessories

- Premium recovery trolley mattress with non-slip, liquid-impermeable cover coating — HR foam ≥ 40 kg/m³, thickness ≥ 80 mm, non-slip base, waterproof PU cover, antimicrobial, flame-retardant.

- Collapsible full-length safety side rails with positive-lock down tabs — Aluminum/SS collapsible rails, positive locking, patient-safe.

- Dual foot-pedal hydraulic pump levers for elevation and Trendelenburg adjustments — Bilateral foot pedals, sealed hydraulic system with overload valve.

- Central braking pedal system with an integrated fifth-wheel steering guide track — Central brake, directional steering wheel for corridor maneuvering.

- Heavy-duty oxygen cylinder storage cradle built into the undercarriage chassis — Under-frame cradle for E-size (≥ 4 L) cylinder, secure retention.

- Dual-end folding push bars and perimeter wall-protection bumper molding — Foldable push handles both ends, PVC/polymer bumper strip.

- Telescopic heavy-duty double-hook IV pole — SS 304, telescopic, load ≥ 8 kg.

## 4.4 As per general bed clauses.

# 5. WARD BED — MANUAL, 2-CRANK

Quantity: 201

## 5.1 Clinical & Functional Specifications

Manually operated hospital ward bed with two independent crank mechanisms for backrest and knee/height adjustment. Intended for general ward inpatients requiring routine nursing care and comfort positioning without power assistance.

## 5.2 Engineering Technical Specifications

- Configuration: Three- or four-section platform, two crank handles (backrest, knee/height).

- Frame: Powder-coated MS or galvanized steel, epoxy finish.

- Overall dimensions: ≈ 2100 × 950 mm.

- Sleeping surface: ≈ 1900 × 850 mm.

- Fixed or two-height selectable design (fixed height ≈ 550 mm).

- Backrest: 0° to ≥ 70° (crank).

- Knee: 0° to ≥ 30° (crank).

- Safe working load: ≥ 200 kg.

- Wheels: ≥ 100 mm, individual brake casters, anti-static.

- Standards: IEC 60601-2-52 (mechanical requirements), ISO 13485, CE.

## 5.3 Standard Accessories

- Medical-grade high-density foam mattress with water-resistant cover — HR foam ≥ 35 kg/m³, thickness ≥ 100 mm, waterproof PU cover, antimicrobial, flame-retardant.

- Fold-away dual-crank mechanical handle mechanisms at the foot end — Chrome-plated foldable cranks, smooth low-effort operation.

- Full-length aluminum alloy collapsible side rails with thumb release latches — Aluminum full-length rails, easy one-hand release.

- Individual lockable swiveling casters on diagonally opposing wheel pairs — ≥ 100 mm anti-static, two diagonally braked.

- Removable lightweight composite head and foot panels — ABS or laminated composite, wipe-cleanable, quick-release.

- Universal IV pole mounting sockets located at all four corners — Four SS-lined sockets accepting standard 20—25 mm poles.

## 5.4 Installation: standard ward space; Documentation, Training, Warranty as general.

# 6. PAEDIATRIC WARD BED

Quantity: 10

## 6.1 Clinical & Functional Specifications

Manually adjustable ward bed for paediatric patients (approximately 2—12 years). Provides safe sleeping enclosure with high-profile side rails to prevent falls, adjustable backrest, and mobility for routine nursing care.

## 6.2 Engineering Technical Specifications

- Configuration: Two- or three-section platform, gas-spring assisted backrest.

- Frame: Powder-coated MS, child-safe rounded edges.

- Overall dimensions: ≈ 1600 × 800 mm.

- Sleeping surface: ≈ 1500 × 700 mm.

- Side-rail height: ≥ 500 mm above mattress surface.

- Backrest: 0° to ≥ 60°.

- Safe working load: ≥ 100 kg.

- Wheels: ≥ 100 mm, four individually braked casters, non-marking.

- Standards: IEC 60601-2-52, CE.

## 6.3 Standard Accessories

- Paediatric density-optimized foam mattress with playful, water-resistant cover fabric — HR foam ≥ 30 kg/m³, thickness ≥ 80 mm, waterproof PU cover, antimicrobial, flame-retardant, child-friendly print.

- Full-length sliding safety side bars with high-profile child-safe security locks — Height ≥ 500 mm, tamper-resistant dual-action lock, gaps compliant with child-safety spacing (≤ 60 mm).

- Manually adjustable gas-spring backrest mechanism with side locking latch levers — Gas-spring assist, positive-lock at multiple angles.

- Heavy-duty independent locking wheels with non-marring rubber tread tracks — ≥ 100 mm, all-brake, non-marking.

- Removable transparent acrylic plastic head and footboard sheets — Impact-resistant acrylic, permitting visual observation, wipe-cleanable.

- Telescopic IV drip stand attachment assembly — SS 304 telescopic pole, load ≥ 5 kg.

## 6.4 As per general bed clauses; paediatric-appropriate training.

# 7. CONSULTANCY / PRIVATE ROOM BED — ELECTRIC

Quantity: 51

## 7.1 Clinical & Functional Specifications

Fully electric private-suite hospital bed combining clinical function with a domestic aesthetic. Provides motorized backrest, knee, height, and Trendelenburg / reverse-Trendelenburg positioning through handheld remote. Intended for private wards, VIP suites, and consultancy rooms.

## 7.2 Engineering Technical Specifications

- Configuration: Four-section platform, minimum three motors, mechanical or motorized Trendelenburg ± 12°.

- Frame: Powder-coated steel with decorative wood-grain or laminated end panels.

- Overall dimensions: ≈ 2150 × 1000 mm.

- Height range: 400—800 mm.

- Backrest: 0° to ≥ 70°.

- Knee: 0° to ≥ 30°.

- Safe working load: ≥ 200 kg.

- Wheels: ≥ 125 mm anti-static, central braking.

- Controls: Handheld pendant with lockout; optional side-rail controls.

- Power: 220—240 VAC; battery backup ≥ 1 hour.

- Standards: IEC 60601-1, 60601-2-52, CE/FDA.

## 7.3 Standard Accessories

- Premium therapeutic pressure-relief mattress with breathable, aesthetic fabric cover — Multi-zone HR/viscoelastic foam ≥ 40 kg/m³, thickness ≥ 120 mm, breathable waterproof cover with pleasant textile finish, antimicrobial, flame-retardant.

- Ergonomic wooden or decorative ABS side rail structures with fold-down tracks — Aesthetic finish, fold-down mechanism, secure locking.

- Flexible patient-access hand remote control pendant line cord hook — Backlit membrane pendant, function lockout, cable ≥ 2 m with holder.

- Emergency backup battery power module for full functional safety operations — ≥ 24 V rechargeable, ≥ 30 motor cycles.

- Central locking mechanism steering casters with single-foot pedal operation — Central brake pedal both sides, steering position.

- Removable decorative wooden-grain head and foot panels — HPL/laminated with wood-grain finish, wipe-cleanable, quick-release.

- Dual IV pole entry frame sockets — SS-lined sockets at both head-end corners.

## 7.4 As per general bed clauses.

# 8. BEDSIDE LOCKER WITH LOCKABLE DRAWER

Quantity: 319

## 8.1 Clinical & Functional Specifications

Mobile bedside patient locker providing secure and accessible storage for patient personal belongings and hygiene items. Designed for hospital ward use with wipe-clean, corrosion-resistant construction.

## 8.2 Engineering Technical Specifications

- Construction: Epoxy-coated MDF, HPL, or ABS body on powder-coated steel frame; seamless wipe-cleanable surfaces.

- Dimensions (approx.): 450 × 450 × 850 mm (W × D × H).

- Compartments: Minimum one lockable drawer + one open/closed lower cabinet with adjustable shelf.

- Load capacity: ≥ 15 kg per shelf/drawer.

- Wheels: Four ≥ 50 mm swivel casters, two with brakes.

- Finish: Antimicrobial, wipe-cleanable, chemical-resistant to hospital disinfectants.

- Standards: ISO 13485, CE, IEC 60601-1 (mechanical), fire retardancy per applicable local code.

## 8.3 Standard Accessories

- Slide-out top drawer tray with Master-key cylinder lock mechanism and two keys — Smooth-glide slides, cylinder lock, two keys per unit, master-key group compatible.

- Adjustable internal shelf plate inside lower cabinet enclosure space — Height-adjustable shelf, ≥ 15 kg load, wipe-clean surface.

- Side-mounted pull-out towel rail bracket hook — Chrome/SS retractable rail on one side.

- Swiveling plastic caster wheels (two front casters equipped with brake tabs) — ≥ 50 mm, non-marking, front two with brakes.

- Wipe-clean seamless ABS top cover plate protective surface mat — Seamless ABS or HPL top, spill-retaining edge, wipe-cleanable.

## 8.4 Installation: standard ward space; Documentation, Training, Warranty as general (Warranty ≥ 2 years).

# 9. OVERBED TABLE

Quantity: 262

## 9.1 Clinical & Functional Specifications

Height-adjustable, over-bed table permitting patient use for meals, reading, and personal care while in bed. Cantilevered base slides beneath ward bed. Gas-spring or mechanical height adjustment.

## 9.2 Engineering Technical Specifications

- Construction: Chrome-plated or powder-coated steel column on cantilever base; HPL or seamless ABS top.

- Top dimensions: ≈ 400 × 800 mm.

- Height range: 750—1150 mm, gas-spring assisted.

- Safe working load: ≥ 25 kg.

- Wheels: Four low-profile ≥ 50 mm casters.

- Standards: ISO 13485, CE.

## 9.3 Standard Accessories

- Gas-spring assisted height-adjustment lift lever mechanism assembly — Sealed gas-spring, single-lever activation, smooth continuous adjustment.

- Seamless molded ABS top platform tray with spill-retaining raised perimeter lips — Seamless ABS top, raised edges, wipe-cleanable, antimicrobial.

- Twin twin-wheel low-profile rolling casters for easy clearance path tracking — Four twin-wheel low-profile casters, non-marking.

- Integrated book or magazine easel support ridge layout tray — Recessed easel ridge or tilt panel.

## 9.4 Installation: standard; Warranty ≥ 2 years.

# 10. IV DRIP STAND / POLE (MOBILE)

Quantity: 332

## 10.1 Clinical & Functional Specifications

Mobile intravenous fluid stand for hanging infusion bags/bottles at controlled height for gravity infusion. Height-adjustable with multi-hook suspension array and stable weighted base.

## 10.2 Engineering Technical Specifications

- Construction: SS 304 telescopic vertical pole; die-cast or SS weighted base.

- Height adjustment: 1300—2100 mm, telescopic with locking knob.

- Hooks: Minimum 4 heavy-duty rotating hooks at top.

- Load per hook: ≥ 2 kg; total load ≥ 8 kg.

- Base: Five-leg star base, diameter ≥ 600 mm, weighted for stability.

- Wheels: Five anti-static casters ≥ 50 mm, minimum two with locking brakes.

- Standards: ISO 13485, CE.

## 10.3 Standard Accessories

- Heavy-duty stainless steel central vertical pole with manual turn-lock screw knob — SS 304, telescopic, positive-lock knob.

- Four-hook heavy-duty top suspension tree array unit adapter — Four SS rotating hooks, ≥ 2 kg each.

- Weighted five-leg low center-of-gravity star base casting ring — Die-cast/SS, weighted, diameter ≥ 600 mm.

- Smooth-rolling anti-static caster wheels (minimum two wheels with lock locks) — ≥ 50 mm anti-static, minimum two lockable.

## 10.4 Warranty ≥ 2 years.

# 11. WHEELCHAIR (ADULT)

Quantity: 20

## 11.1 Clinical & Functional Specifications

Foldable, manually operated wheelchair for adult patient transportation within the hospital environment. Enables safe intra-hospital transfer of ambulatory-impaired patients.

## 11.2 Engineering Technical Specifications

- Frame: Chromed/powder-coated steel dual cross-brace, foldable.

- Overall dimensions (open): ≈ 1050 × 650 × 900 mm.

- Seat width: 450—500 mm.

- Rear wheels: ≥ 600 mm spoked, solid or pneumatic flat-free tires, push-rim.

- Front casters: ≥ 200 mm solid.

- Safe working load: ≥ 120 kg.

- Folded width: ≤ 350 mm.

- Standards: ISO 7176 series, ISO 13485, CE.

## 11.3 Standard Accessories

- Heavy-duty dual-steel cross-brace frame with anti-rust chrome coating — Chromed steel, foldable, corrosion-resistant.

- Durable washable nylon upholstery seating sheet and rear documentation pocket — Reinforced flame-retardant nylon, washable, rear pouch.

- Swing-away detachable footrest assemblies with adjustable composite footplates — Swing-away, detachable, height-adjustable footplates.

- Flip-back padded armrests with protective side panel shields — Flip-back padded armrests with skirt guards.

- Toggle-locking manual wheel brake levers on rear large wheels — Positive-lock lever brakes, both sides.

- Solid flat-free front caster assemblies and large spoke rear wheels — Solid PU front casters, spoked rear wheels with push-rim.

## 11.4 Warranty ≥ 2 years.

# 12. WHEELCHAIR (PAEDIATRIC)

Quantity: 2

## 12.1 Clinical & Functional Specifications

Foldable pediatric wheelchair for safe intra-hospital transport of child patients.

## 12.2 Engineering Technical Specifications

- Frame: Powder-coated steel, high-visibility color, foldable cross-brace.

- Seat width: 300—380 mm.

- Rear wheels: 400—500 mm.

- Front casters: ≥ 150 mm solid PU.

- Safe working load: ≥ 60 kg.

- Standards: ISO 7176, ISO 13485, CE.

## 12.3 Standard Accessories

- Compact child-sized cross-brace steel frame with high-visibility color powder coating — Powder-coated steel, child-friendly color.

- Washable padded fabric seat cushion liners with integrated lap safety belt straps — Removable washable padded liners with 3-point or lap belt.

- Swing-away height-adjustable footplates with rear heel-loop strap supports — Swing-away footplates, heel loops.

- Desk-length desk armrests with smooth plastic arm pad liners — Desk-length armrests, padded.

- Push-to-lock manual toggle wheel brake levers handles — Both-side toggle brakes.

- High-density polyurethane flat-free wheel tires set — Solid PU tires, puncture-free.

## 12.4 Warranty ≥ 2 years.

# 13. PATIENT TROLLEY FOR WARD USE

Quantity: 12

## 13.1 Clinical & Functional Specifications

Mobile height-adjustable patient trolley for intra-departmental patient transport, examination, and short-term observation.

## 13.2 Engineering Technical Specifications

- Configuration: Two-section radiolucent platform, gas-spring backrest.

- Frame: Epoxy-coated MS/SS, ABS end panels.

- Overall dimensions: ≈ 2050 × 700 mm.

- Height: 600—850 mm (fixed or hydraulic).

- Backrest: 0° to ≥ 70°.

- Safe working load: ≥ 180 kg.

- Wheels: ≥ 150 mm, four twin-wheel casters, individual or central brake.

- Standards: IEC 60601-1, 60601-2-52, CE.

## 13.3 Standard Accessories

- High-density transit mattress pad with waterproof, antistatic cover wrap — HR foam ≥ 35 kg/m³, ≥ 60 mm thick, waterproof PU cover, antimicrobial, antistatic, flame-retardant.

- Collapsible safety side rails with automatic structural locking mechanics — Full-length collapsible, positive lock.

- Gas-spring assisted backrest elevation lever system bracket — Gas-spring assist, angle-lock lever.

- Four individual caster lock levers with directional tracking locks — All-brake casters with direction lock.

- Integrated oxygen bottle storage tray grid underneath the platform base — Under-frame cradle for E-size O₂ cylinder.

- Dual-end fixed push handles and perimeter rubber buffer guards — Fixed handles both ends, perimeter bumpers.

## 13.4 Warranty ≥ 2 years.

# 14. SCREENS / CURTAIN RAILS (BED PRIVACY)

Quantity: 235

## 14.1 Clinical & Functional Specifications

Ceiling-mounted bed-privacy curtain system providing patient privacy in shared wards, examination areas, and recovery bays.

## 14.2 Engineering Technical Specifications

- Track: Extruded aluminum alloy, anodized, minimum 2 mm wall thickness; internal glide channel.

- Track lengths: Configurable 2—6 m per bed as per site layout.

- Curtain fabric: Antimicrobial, flame-retardant (BS 5867 Part 2/B or equivalent), polyester with mesh top ≥ 500 mm for air circulation and sprinkler compliance.

- Ceiling drop rods: ≥ 10 mm dia SS, height-adjustable.

- Standards: Local building code, fire retardancy standards.

## 14.3 Standard Accessories

- Rigid ceiling-mount aluminum alloy suspension tracks rail profiles — Anodized aluminum, internal glide channel.

- Antimicrobial, flame-retardant mesh-top fabric privacy curtains panels — Polyester with ≥ 500 mm mesh top, antimicrobial, flame-retardant, washable.

- Low-friction plastic roller wheel hook carriers glide assemblies — Nylon rollers, smooth glide, corrosion-free.

- Wall-anchor support stay bars and terminal ceiling drop rods hangers — SS drop rods and wall stays for track support.

## 14.4 Installation to ceiling structure; Warranty ≥ 2 years.

# 15. MEDICINE TROLLEY (NURSE)

Quantity: 14

## 15.1 Clinical & Functional Specifications

Mobile lockable trolley for safe distribution and administration of unit-dose medications in the ward. Enables individualized patient medication compartments with security controls.

## 15.2 Engineering Technical Specifications

- Construction: Epoxy-coated steel or ABS body on SS or steel frame.

- Dimensions: ≈ 750 × 500 × 950 mm.

- Drawers: Multi-drawer configuration (minimum 8—20 patient bins) with central locking.

- Load per drawer: ≥ 5 kg.

- Wheels: ≥ 100 mm swivel, four casters, minimum two with brakes.

- Standards: ISO 13485, CE.

## 15.3 Standard Accessories

- Key-lockable multi-drawer structural configuration with central locking frame system — Multi-drawer, cylinder central lock, master-key.

- Removable adjustable internal pharmaceutical layout trays dividers — Modular dividers for unit-dose organization.

- Slide-out side work surface writing shelf board — Retractable HPL/ABS writing surface.

- Molded ABS top work deck plate with raised safety guard perimeter walls — Seamless ABS top with raised edges.

- Four heavy-duty swivel casters wheels with tracking diagonal dual brake locks — ≥ 100 mm, four casters, two diagonal brakes.

## 15.4 Warranty ≥ 2 years.

# 16. DRESSING TROLLEY (SS, DOUBLE SHELF)

Quantity: 23

## 16.1 Clinical & Functional Specifications

Stainless-steel double-shelf trolley for transport of sterile dressing sets, instruments, and consumables to bedside for wound care and minor procedures.

## 16.2 Engineering Technical Specifications

- Construction: SS 304, seamless welded, satin/polished finish.

- Dimensions: ≈ 750 × 500 × 850 mm.

- Shelves: Two, ≥ 1 mm SS sheet, minimum 15 kg per shelf.

- Wheels: ≥ 100 mm anti-static swivel, minimum two lockable.

- Standards: ISO 13485, CE.

## 16.3 Standard Accessories

- Premium stainless steel frame tubes with two seamless flat shelf trays — SS 304 seamless.

- Three-sided stainless steel protective guard rails around both storage decks — SS guard rails on three sides of each shelf.

- Removable deep-profile stainless steel utility bowls (set of 2, with mount hooks) — Two removable SS bowls with hooks.

- Anti-static swivel caster wheels (two casters matching with locking tabs) — ≥ 100 mm, two locking.

## 16.4 Warranty ≥ 2 years.

# 17. NURSE STATION DESK / COUNTER (MODULAR)

Quantity: 14

## 17.1 Clinical & Functional Specifications

Modular ergonomic nurse station providing workspace for clinical documentation, computer workstations, monitoring, and communication.

## 17.2 Engineering Technical Specifications

- Construction: Modular HPL panels on steel frame; scratch-resistant, wipe-cleanable.

- Length: Configurable modules 1.5—3.0 m per module.

- Depth: 600—800 mm; Height 750 mm work surface + optional privacy panel.

- Load: ≥ 100 kg per module worksurface.

- Standards: ISO 13485 (where applicable), local building code, fire-retardant finishes.

## 17.3 Standard Accessories

- Modular privacy facade front panels with scratch-resistant high-pressure laminate finishes — HPL front facia, ≥ 0.8 mm thickness, scratch-resistant.

- Integrated wire-management grommets and desktop under-counter cable cable trays — Grommets and steel cable trays for power/data.

- Mobile three-drawer under-desk filing pedestal storage unit with drawer locks — Mobile 3-drawer pedestal with cylinder lock, one file-size drawer.

- Ergonomic lower keyboard tray glides track systems — Retractable keyboard tray on smooth ball-bearing glides.

## 17.4 Warranty ≥ 2 years.

# 18. AIR MATTRESS WITH PUMP

Quantity: 59

## 18.1 Clinical & Functional Specifications

Alternating-pressure air mattress system for prevention and treatment of stage I—IV pressure injuries in immobile and high-risk patients.

## 18.2 Engineering Technical Specifications

- Mattress type: Alternating pressure, tubular or bubble-cell design, minimum 17 cells, ≥ 100 mm cell height.

- Materials: PU-coated nylon, waterproof, antimicrobial, vapor-permeable top.

- Cycle time: Selectable 10—25 minutes.

- Pressure range: Adjustable 20—60 mmHg according to patient weight.

- Patient weight capacity: ≥ 135 kg (options up to 200 kg).

- Pump: Microprocessor-controlled, silent (\< 30 dB), low-air-loss capability.

- Control panel: Digital display with pressure, cycle time, alarm indicators.

- CPR release: Rapid deflation \< 20 seconds via CPR valve.

- Power: 220—240 VAC, 50/60 Hz; ≤ 20 W.

- Standards: IEC 60601-1, 60601-1-2, ISO 13485, CE/FDA.

## 18.3 Standard Accessories

- Alternating pressure micro-pore air cell mattress pad overlays — Multi-cell alternating tubes, PU-coated nylon, ≥ 100 mm cell height.

- Microprocessor-controlled pressure variable compressor pump system module — Digital pump, selectable cycle and pressure, alarms for low pressure and power failure.

- Integrated bed-rail support hanger brackets on rear pump casing — Bed-end hanger brackets for pump suspension.

- Emergency CPR rapid-deflation rip-cord valve plug attachment — Rapid CPR valve, \< 20 s full deflation.

- Waterproof, vapor-permeable top comfort stretch sheet sheet cover — Bi-elastic PU top cover, waterproof, vapor-permeable, antimicrobial.

- Emergency puncture repair patch utility kit — PU patch kit with adhesive.

## 18.4 Warranty ≥ 2 years.

# 19. CATHETER STORAGE CABINET

Quantity: 1

## 19.1 Clinical & Functional Specifications

Secure vertical storage cabinet for orderly presentation and easy retrieval of urinary, cardiac, and vascular catheters, preventing kinking and contamination.

## 19.2 Engineering Technical Specifications

- Construction: Epoxy-coated steel or SS 304 body with tempered glass doors.

- Dimensions: ≈ 900 × 400 × 1800 mm.

- Compartments: Multiple pull-out catheter hanging racks + adjustable shelves.

- Locking: Cylinder lock with master key.

- Standards: ISO 13485, CE.

## 19.3 Standard Accessories

- Heavy-duty medical cabinet body with dual glass-paneled lockable doors — Steel/SS body, tempered glass double doors, cylinder lock.

- Sliding telescopic heavy-duty pull-out catheter hook racks tracking mounts — Telescopic pull-out rails with catheter hanging hooks.

- Adjusting horizontal storage compartment shelves boxes — Adjustable-height wire or SS shelves.

- Master-key locking cylinder handle setup with twin keys — Cylinder lock, two keys.

## 19.4 Warranty ≥ 2 years.

# 20. INSTRUMENT CABINET

Quantity: 6

## 20.1 Clinical & Functional Specifications

Lockable glass-front cabinet for secure display and storage of surgical and clinical instruments.

## 20.2 Engineering Technical Specifications

- Construction: SS 304 frame with tempered safety glass doors and sides.

- Dimensions: ≈ 900 × 450 × 1800 mm.

- Shelves: Minimum 4 adjustable tempered-glass shelves.

- Load per shelf: ≥ 20 kg.

- Levelling feet: Height-adjustable.

- Standards: ISO 13485, CE.

## 20.3 Standard Accessories

- Stainless steel frame cabinet enclosure with tempered safety glass view doors — SS 304 frame, tempered glass doors.

- Adjustable heavy-duty clear tempered safety glass shelf plates (set of 4 shelves) — Four tempered glass shelves, height-adjustable, ≥ 20 kg load.

- Secure key-locking door handle lock mechanisms with duplicate master keys — Cylinder lock, two keys.

- Height-adjustable floor leveling feet studs bases — SS levelling feet.

## 20.4 Warranty ≥ 2 years.

# 21. INSTRUMENT TROLLEY (SS, DOUBLE SHELF)

Quantity: 12

## 21.1 Clinical & Functional Specifications

Mobile stainless-steel trolley with two shelves for transporting surgical and clinical instruments within OR, wards, and clinics.

## 21.2 Engineering Technical Specifications

- Construction: SS 304 frame with two seamless shelves.

- Dimensions: ≈ 750 × 500 × 850 mm.

- Load per shelf: ≥ 20 kg.

- Wheels: ≥ 100 mm anti-static, minimum two lockable.

- Standards: ISO 13485, CE.

## 21.3 Standard Accessories

- High-grade polished stainless steel structural pipes assembly setup — SS 304 polished tubes.

- Two seamless sound-dampened stainless steel storage shelves plates — Two SS shelves, ≥ 1 mm thickness.

- Four-side containment protective guard bars surrounding shelf edges — Perimeter SS guard rails.

- Silent-rolling anti-static caster wheels with protective bumper rings — ≥ 100 mm, anti-static, with bumper rings.

## 21.4 Warranty ≥ 2 years.

# 22. UTILITY / DRESSING TROLLEY (SS)

Quantity: 6

## 22.1 Clinical & Functional Specifications

Stainless-steel dressing and utility trolley with waste bin for wound care and general clinical utility.

## 22.2 Engineering Technical Specifications

- Construction: SS 304 dual-deck frame.

- Dimensions: ≈ 800 × 500 × 900 mm.

- Wheels: ≥ 100 mm swivel, all-brake.

- Standards: ISO 13485, CE.

## 22.3 Standard Accessories

- Seamless stainless steel dual-deck design trolley frame construction — Seamless SS 304 dual-deck.

- Removable plastic waste bucket receptacle bin with pivoting support frame clamp — Removable HDPE bin with SS pivot clamp.

- Ergonomic structural push handle bar loop extension — SS ergonomic push handle.

- Four multi-axis swiveling casters wheels with dual locking brake levers — Four ≥ 100 mm swivel with brakes.

## 22.4 Warranty ≥ 2 years.

# 23. SS STORAGE CABINET (LOCKABLE)

Quantity: 14

## 23.1 Clinical & Functional Specifications

Heavy-duty lockable stainless-steel storage cabinet for consumables, linen, and clinical supplies.

## 23.2 Engineering Technical Specifications

- Construction: SS 304 double-walled body.

- Dimensions: ≈ 900 × 450 × 1800 mm.

- Shelves: Minimum 4 adjustable SS shelves.

- Load per shelf: ≥ 30 kg.

- Locking: Three-point locking with cylinder lock and master key.

- Standards: ISO 13485, CE.

## 23.3 Standard Accessories

- Heavy-duty double-walled stainless steel sheets cabinet structure — Double-walled SS 304 body.

- Adjustable reinforced structural stainless steel shelves (set of 4 shelves) — Four adjustable SS shelves, ≥ 30 kg load.

- Three-point secure locking mechanism latch with master cylinder key setup — Three-point rod-lock, cylinder key, master-keyed.

- Sealed bottom base frame kick plate strip — Sealed SS kick plate at base.

## 23.4 Warranty ≥ 2 years.

# 24. MAYO STAND

Quantity: 6

## 24.1 Clinical & Functional Specifications

Height-adjustable mobile instrument stand providing sterile field working platform for surgical procedures, positioned across the patient.

## 24.2 Engineering Technical Specifications

- Construction: SS 304 frame; removable SS instrument tray.

- Tray dimensions: ≈ 500 × 350 mm.

- Height range: 850—1350 mm, foot-pedal or twist-lock.

- Load capacity: ≥ 10 kg.

- Wheels: Four ≥ 50 mm anti-static, low-clearance.

- Standards: ISO 13485, CE.

## 24.3 Standard Accessories

- Removable stainless steel surgical instrument presentation tray deck plate — SS 304 removable tray, autoclavable.

- Foot-pedal operated or manual twist mechanical lift elevator lock system — Foot pedal or twist-lock height adjustment.

- Heavy-duty low-clearance double-prong steel frame foot base — Twin-prong low-clearance SS base fitting over OR table.

- Anti-static smooth-glide caster wheels tracking mounts — Four ≥ 50 mm anti-static casters.

## 24.4 Warranty ≥ 2 years.

# 25. KICK BUCKET

Quantity: 6

## 25.1 Clinical & Functional Specifications

Mobile foot-operated stainless-steel bucket on wheeled ring frame for hygienic disposal of soiled swabs and waste in OR and treatment rooms.

## 25.2 Engineering Technical Specifications

- Bucket: SS 304, ≥ 12 L, graduated, removable.

- Frame: Powder-coated MS or SS ring frame.

- Wheels: Four ≥ 50 mm anti-static swivel casters.

- Bumper: Continuous rubber bumper on outer ring.

- Standards: ISO 13485, CE.

## 25.3 Standard Accessories

- Removable graduated stainless steel open operating room bucket bowl — SS 304, ≥ 12 L, graduated inside.

- Heavy-duty tip-resistant wheeled base frame trolley ring housing — Tip-resistant wheeled ring frame.

- Continuous perimeter protective rubber buffer wrap strip bumper ring — Continuous rubber bumper.

- Conductive swiveling anti-static casters wheels (set of 4) — Four anti-static swivel casters.

## 25.4 Warranty ≥ 2 years.

# 26. LINEN HAMPER

Quantity: 14

## 26.1 Clinical & Functional Specifications

Foot-pedal mobile linen collection hamper for hygienic collection of soiled hospital linen.

## 26.2 Engineering Technical Specifications

- Frame: SS 304 tubular frame with foot-pedal spring-lid.

- Bag capacity: ≥ 90 L.

- Wheels: Four ≥ 50 mm swivel anti-static.

- Standards: ISO 13485, CE.

## 26.3 Standard Accessories

- Stainless steel ring support frame loop holding bag mount — SS 304 ring frame with bag retention.

- Foot-pedal activated heavy-duty spring-lid mechanism panel — Foot-pedal spring-lid, hands-free.

- Heavy-duty washable synthetic canvas canvas linen collection bags (2 bags pack) — Two washable canvas bags per unit.

- Non-marring anti-static swivel caster wheels setups — Four anti-static casters, non-marking.

## 26.4 Warranty ≥ 2 years.

# 27. FOOT STEP

Quantity: 14

## 27.1 Clinical & Functional Specifications

Single- or double-step patient step for safe access to and from hospital beds and examination couches.

## 27.2 Engineering Technical Specifications

- Construction: Powder-coated steel or SS frame; slip-resistant rubber tread top.

- Dimensions: ≈ 400 × 400 × 200 mm (single step) or 500 × 400 × 400 mm (double).

- Load capacity: ≥ 150 kg.

- Feet: Non-slip rubber caps on all legs.

- Standards: ISO 13485, CE.

## 27.3 Standard Accessories

- Heavy-duty reinforced double steel step frame structural base — Powder-coated reinforced steel frame.

- Corrugated slip-resistant thick rubber top tread mats platform linings — Corrugated slip-resistant rubber top ≥ 5 mm.

- Reinforced non-skid rubber feet shoe caps tips on all support legs — Non-slip rubber tips all legs.

## 27.4 Warranty ≥ 2 years.

# 28. MEDICAL / MEDICINE REFRIGERATOR (2—8 °C)

Quantity: 15

## 28.1 Clinical & Functional Specifications

Purpose-built medical refrigerator for storage of vaccines, biologicals, and thermosensitive pharmaceuticals within a controlled 2—8 °C range. Provides continuous temperature monitoring, alarm, and data logging.

## 28.2 Engineering Technical Specifications

- Storage volume: Selectable 100—400 L configurations.

- Temperature range: 2—8 °C, uniformity ± 1 °C.

- Temperature accuracy: ± 0.5 °C.

- Ambient rating: Suitable for 10—35 °C ambient.

- Refrigerant: CFC/HCFC-free (R600a or equivalent).

- Cabinet: CFC-free insulated PU foam, powder-coated steel exterior, ABS/SS interior.

- Door: Triple-pane tempered glass with heater; lockable.

- Shelves: Minimum 3 adjustable epoxy-coated wire shelves.

- Control: Microprocessor with LED/LCD digital display of chamber temperature.

- Alarms: Visual + audible for high/low temperature, door-open, power failure, sensor failure, low battery.

- Data logging: Built-in or via USB / RS-232 / Ethernet; downloadable temperature log; battery-backup memory.

- Battery backup: ≥ 72 hours for alarms and data logger.

- Power: 220—240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2, WHO PQS E003, ISO 13485, CE/FDA.

## 28.3 Standard Accessories

- Microprocessor control digital display temperature tracking module system — Microprocessor controller with LED/LCD digital readout, resolution 0.1 °C.

- Integrated acoustic and visual temperature deviation sensor alarm alarms — Multi-alarm (temperature, door, power, sensor), battery-backed, remote-alarm dry contact.

- Triple-pane vacuum tempered safety glass door window with lock mechanism — Anti-condensation triple-pane glass door, cylinder lock, two keys.

- Adjustable epoxy-coated wire wire organizing storage shelves (set of 4 shelves) — Four epoxy-coated corrosion-resistant wire shelves, adjustable height.

- Integrated digital tracking chart recorder or data logging USB slot port — Built-in data logger with USB / RS-232 / Ethernet download, minimum 3-month internal memory, PDF/CSV export.

## 28.4 Installation Requirements

Dedicated 220—240 VAC earthed socket, UPS-compatible; minimum 100 mm clearance on sides and rear; ambient 10—32 °C; away from direct sunlight and heat sources; leveled floor.

## 28.5 Documentation, Training, Warranty as general (Warranty ≥ 2 years, comprehensive).

# 29. AUTOCLAVE (24 L)

Quantity: 1

## 29.1 Clinical & Functional Specifications

Bench-top steam sterilizer for terminal sterilization of surgical instruments, wrapped and unwrapped loads, and heat-stable materials. Provides validated saturated steam sterilization cycles for hospital CSSD and clinic use.

## 29.2 Engineering Technical Specifications

- Chamber: SS 316L, cylindrical, ≥ 24 L capacity.

- Chamber dimensions (approx.): Ø 240 × 500 mm.

- Sterilization temperature: Selectable 121 °C and 134 °C.

- Working pressure: Up to 2.1 bar (absolute).

- Cycle types: Minimum three programs (unwrapped, wrapped, liquid).

- Class (per EN 13060): Minimum Class N; Class B preferred for wrapped/hollow loads.

- Vacuum: Fractional pre-vacuum for Class B.

- Drying: Post-cycle drying phase.

- Water: Distilled water reservoir with low-level cutoff.

- Control: Microprocessor with digital display of temperature, pressure, and cycle status.

- Safety: Overpressure valve, overtemperature cutoff, door-interlock, low-water cutoff.

- Printer / data logging: Optional built-in printer / USB data output.

- Power: 220—240 VAC, 50/60 Hz, ≤ 2 kW.

- Standards: IEC 60601-1, EN 13060, ISO 13485, CE/FDA.

## 29.3 Standard Accessories

- Stainless steel sterilization chamber drum cylinder with secure screw latch wheel lid — SS 316L chamber with positive-lock lid and safety interlock.

- Perforated stainless steel internal load organizing trays (set of 2 trays) — Two perforated SS trays with tray handler.

- Dual-scale high-precision temperature and steam pressure monitoring gauge dials — Analog + digital indication of temperature and pressure.

- Integrated mechanical over-pressure exhaust safety vent safety valve module — Calibrated safety relief valve.

- Low-water element power cutoff protector switch indicator sensor — Low-water sensor with automatic heater cutoff and alarm.

## 29.4 Installation Requirements

Dedicated 220—240 VAC earthed socket (16 A); distilled water supply; adequate ventilation; heat-tolerant benchtop; ≥ 100 mm clearance rear/sides.

5—Warranty ≥ 2 years.