BLOOD BANK — TECHNICAL SPECIFICATIONS

Technical Tender Document

Table of Contents

1. Blood Grouping and Cross-Matching Workstation	9

1.1 Clinical & Functional Specifications	9

1.2 Engineering Technical Specifications	9

1.3 Standard Accessories	10

1.4 Installation Requirements	10

1.5 Documentation	10

1.6 Training	10

1.7 Warranty	10

2. Agglutination Viewer	12

2.1 Clinical & Functional Specifications	12

2.2 Engineering Technical Specifications	12

2.3 Standard Accessories	12

2.4 Installation Requirements	12

2.5 Documentation	13

2.6 Training	13

2.7 Warranty	13

3. Blood Bank Refrigerator (2—6 °C, 500-Bag Capacity)	14

3.1 Clinical & Functional Specifications	14

3.2 Engineering Technical Specifications	14

3.3 Standard Accessories	15

3.4 Installation Requirements	15

3.5 Documentation	15

3.6 Training	15

3.7 Warranty	15

4. Plasma Freezer (−30 °C / −40 °C)	17

4.1 Clinical & Functional Specifications	17

4.2 Engineering Technical Specifications	17

4.3 Standard Accessories	17

4.4 Installation Requirements	18

4.5 Documentation	18

4.6 Training	18

4.7 Warranty	18

5. Platelet Incubator with Agitator	19

5.1 Clinical & Functional Specifications	19

5.2 Engineering Technical Specifications	19

5.3 Standard Accessories	19

5.4 Installation Requirements	20

5.5 Documentation	20

5.6 Training	20

5.7 Warranty	20

6. Water Bath / Heat Block	21

6.1 Clinical & Functional Specifications	21

6.2 Engineering Technical Specifications	21

6.3 Standard Accessories	21

6.4 Installation Requirements	21

6.5 Documentation	22

6.6 Training	22

6.7 Warranty	22

7. Binocular Microscope	23

7.1 Clinical & Functional Specifications	23

7.2 Engineering Technical Specifications	23

7.3 Standard Accessories	23

7.4 Installation Requirements	24

7.5 Documentation	24

7.6 Training	24

7.7 Warranty	24

8. Donor Blood Collection Chair	25

8.1 Clinical & Functional Specifications	25

8.2 Engineering Technical Specifications	25

8.3 Standard Accessories	25

8.4 Installation Requirements	26

8.5 Documentation	26

8.6 Training	26

8.7 Warranty	26

9. Blood Collection Monitor / Mixer	27

9.1 Clinical & Functional Specifications	27

9.2 Engineering Technical Specifications	27

9.3 Standard Accessories	27

9.4 Installation Requirements	28

9.5 Documentation	28

9.6 Training	28

9.7 Warranty	28

10. Plasma Thawing Bath	29

10.1 Clinical & Functional Specifications	29

10.2 Engineering Technical Specifications	29

10.3 Standard Accessories	29

10.4 Installation Requirements	30

10.5 Documentation	30

10.6 Training	30

10.7 Warranty	30

11. Tube Sealer for Blood Bags	31

11.1 Clinical & Functional Specifications	31

11.2 Engineering Technical Specifications	31

11.3 Standard Accessories	31

11.4 Installation Requirements	31

11.5 Documentation	32

11.6 Training	32

11.7 Warranty	32

12. Validated Blood Transport Box	33

12.1 Clinical & Functional Specifications	33

12.2 Engineering Technical Specifications	33

12.3 Standard Accessories	33

12.4 Installation Requirements	34

12.5 Documentation	34

12.6 Training	34

12.7 Warranty	34

13. Bench-Top Centrifuge	35

13.1 Clinical & Functional Specifications	35

13.2 Engineering Technical Specifications	35

13.3 Standard Accessories	35

13.4 Installation Requirements	36

13.5 Documentation	36

13.6 Training	36

13.7 Warranty	36

14. Incubator	37

14.1 Clinical & Functional Specifications	37

14.2 Engineering Technical Specifications	37

14.3 Standard Accessories	37

14.4 Installation Requirements	37

14.5 Documentation	38

14.6 Training	38

14.7 Warranty	38

15. Donor Blood Collection Bed	39

15.1 Clinical & Functional Specifications	39

15.2 Engineering Technical Specifications	39

15.3 Standard Accessories	39

15.4 Installation Requirements	40

15.5 Documentation	40

15.6 Training	40

15.7 Warranty	40

16. Tube Stripper	41

16.1 Clinical & Functional Specifications	41

16.2 Engineering Technical Specifications	41

16.3 Standard Accessories	41

16.4 Installation Requirements	41

16.5 Documentation	41

16.6 Training	41

16.7 Warranty	41

17. Emergency Cabinet (Blood Bank Crash Cart)	42

17.1 Clinical & Functional Specifications	42

17.2 Engineering Technical Specifications	42

17.3 Standard Accessories	42

17.4 Installation Requirements	42

17.5 Documentation	43

17.6 Training	43

17.7 Warranty	43

18. Slide Storage Cabinet	44

18.1 Clinical & Functional Specifications	44

18.2 Engineering Technical Specifications	44

18.3 Standard Accessories	44

18.4 Installation Requirements	44

18.5 Documentation	44

18.6 Training	44

18.7 Warranty	44

19. Plasma Extractor	46

19.1 Clinical & Functional Specifications	46

19.2 Engineering Technical Specifications	46

19.3 Standard Accessories	46

19.4 Installation Requirements	46

19.5 Documentation	46

19.6 Training	46

19.7 Warranty	47

20. Weighing Machine	48

20.1 Clinical & Functional Specifications	48

20.2 Engineering Technical Specifications	48

20.3 Standard Accessories	48

20.4 Installation Requirements	48

20.5 Documentation	48

20.6 Training	48

20.7 Warranty	49

21. CLIA (Chemiluminescence Immunoassay) Machine	50

21.1 Clinical & Functional Specifications	50

21.2 Engineering Technical Specifications	50

21.3 Standard Accessories	51

21.4 Installation Requirements	51

21.5 Documentation	51

21.6 Training	51

21.7 Warranty	51

22. Refrigerated Centrifuge	52

22.1 Clinical & Functional Specifications	52

22.2 Engineering Technical Specifications	52

22.3 Standard Accessories	52

22.4 Installation Requirements	53

22.5 Documentation	53

22.6 Training	53

22.7 Warranty	53

23. Refrigerator Thermometer	54

23.1 Clinical & Functional Specifications	54

23.2 Engineering Technical Specifications	54

23.3 Standard Accessories	54

23.4 Installation Requirements	54

23.5 Documentation	55

23.6 Training	55

23.7 Warranty	55

24. Small CBC (3-Part Differential) Analyser	56

24.1 Clinical & Functional Specifications	56

24.2 Engineering Technical Specifications	56

24.3 Standard Accessories	56

24.4 Installation Requirements	57

24.5 Documentation	57

24.6 Training	57

24.7 Warranty	57

25. Waste Disposal Equipment Bundle	58

25.1 Clinical & Functional Specifications	58

25.2 Engineering Technical Specifications	58

25.3 Standard Accessories	58

25.4 Installation Requirements	58

25.5 Documentation	58

25.6 Training	58

25.7 Warranty	59

26. Phlebotomy Tray	60

26.1 Clinical & Functional Specifications	60

26.2 Engineering Technical Specifications	60

26.3 Standard Accessories	60

26.4 Installation Requirements	60

26.5 Documentation	60

26.6 Training	60

26.7 Warranty	60

27. Apheresis Machine	61

27.1 Clinical & Functional Specifications	61

27.2 Engineering Technical Specifications	61

27.3 Standard Accessories	62

27.4 Installation Requirements	62

27.5 Documentation	62

27.6 Training	62

27.7 Warranty	62

# 1. Blood Grouping and Cross-Matching Workstation

Quantity: 1

## 1.1 Clinical & Functional Specifications

The workstation shall be a complete, integrated bench-top system intended for immunohematology testing including ABO/Rh blood grouping, antibody screening, antibody identification, direct and indirect antiglobulin (Coombs') testing, and cross-matching using the column agglutination (gel card) technique or equivalent internationally accepted methodology. The system shall enable safe, standardized, and reproducible pre-transfusion testing in a hospital blood bank environment and shall be suitable for both routine and emergency workloads. It shall permit manual, semi-automated, or automated operation as configured, and shall provide clear visual reading of agglutination reactions. Workflow shall support sample identification, reagent dispensing, incubation, centrifugation, and result interpretation as an integrated process.

## 1.2 Engineering Technical Specifications

- Configuration: Integrated bench-top immunohematology workstation comprising, at minimum, a dedicated gel card centrifuge, gel card incubator, calibrated pipetting system, and reagent/card organization accessories.

- Test methodology: Column agglutination technology (gel or glass bead) or equivalent internationally accepted immunohematology technique.

- Test capacity (centrifuge): Minimum 12 gel cards per run or equivalent.

- Test capacity (incubator): Minimum 12 gel cards per cycle or equivalent.

- Centrifuge speed: Selectable/pre-programmed within the range required for gel card testing (typically up to approximately 1500 rpm), electronically controlled.

- Centrifuge timer: Digitally adjustable, minimum 0—15 minutes, with automatic cycle termination.

- Speed accuracy: Within ±2% of set value.

- Imbalance detection: Automatic electronic imbalance sensing with cycle abort.

- Lid safety: Electromagnetic lid interlock preventing operation with lid open and preventing opening during rotation.

- Incubator temperature: Factory pre-set/adjustable at 37 °C ± 1 °C, microprocessor-controlled.

- Temperature uniformity: Within ±1 °C across all card positions.

- Temperature display: Digital LED/LCD, resolution 0.1 °C.

- Alarms: Audible and visual alarms for over-temperature, under-temperature, and end-of-cycle.

- User interface: Membrane keypad or touchscreen with clearly labeled controls; multilingual or English menu.

- Pipettes: Adjustable single-channel and multi-channel micropipettes covering the working range required for immunohematology (typically 5—1000 µL) with calibration certificate.

- Housing: Corrosion-resistant, chemical-resistant, easy-to-disinfect enclosure.

- Power supply: 220—240 VAC, 50/60 Hz (or as specified by purchaser), with integrated surge protection.

- Operating environment: 15—30 °C ambient, 20—80% RH non-condensing.

- Weight / footprint: Compact tabletop configuration suitable for standard laboratory bench installation.

- Applicable standards: IEC 61010-1, IEC 61010-2-020 (centrifuge safety), IEC 61010-2-101 (IVD equipment), IEC 61326 (EMC), ISO 13485 manufacturing, CE-IVD or US-FDA cleared or equivalent.

## 1.3 Standard Accessories

- Gel Card Centrifuge Head — Dedicated rotor for gel card format, corrosion-resistant metallic construction, compatible with standard gel/column agglutination cards, electronically controlled variable speed, integrated digital timer, imbalance sensor, and safety lid interlock.

- Gel Card Incubator Block — Dry-block incubator with microprocessor-controlled heating maintained at 37 °C ± 1 °C, capacity matched to centrifuge (minimum 12 cards), digital temperature display with 0.1 °C resolution, over-temperature alarm.

- Micropipette Kit — Set comprising adjustable single-channel and multi-channel air-displacement pipettes covering 5—1000 µL range, autoclavable lower parts, calibration certificate traceable to international standards, ergonomic design.

- Reagent Vial Rack and Card Organizer — Chemical-resistant polymer construction, holders for standard reagent vials and gel/column cards, stable non-slip base, easy to clean and disinfect.

- Main Power Cable with Surge Protection — Medical-grade detachable power cord matched to local mains, integrated surge/spike protection module rated for hospital electrical environment.

## 1.4 Installation Requirements

Standard laboratory bench with stable, vibration-free surface. Single-phase 220—240 VAC supply with proper earthing, connected preferably through UPS/stabilized supply. Ambient room temperature 18—28 °C with adequate ventilation. Adequate bench space for centrifuge, incubator, and reagent storage.

## 1.5 Documentation

User Manual (English), Service Manual, Parts Catalogue, Calibration Certificate for centrifuge and pipettes, Factory Test Certificate, Electrical Safety Test Report, Preventive Maintenance Schedule, Installation and Commissioning Report, CE/FDA or equivalent regulatory certificate, ISO 13485 certificate.

## 1.6 Training

Clinical/application training for blood bank technologists, biomedical engineering training for in-house engineers, preventive maintenance and basic troubleshooting training, delivered on-site after installation.

## 1.7 Warranty

Minimum two (2) years comprehensive warranty including labour, spare parts, preventive maintenance visits, breakdown maintenance, software updates, and technical support.

# 2. Agglutination Viewer

Quantity: 2

## 2.1 Clinical & Functional Specifications

A dedicated optical viewing device intended for macroscopic reading and interpretation of agglutination reactions in blood grouping, cross-matching, and antibody screening performed in test tubes or slides. The viewer shall provide uniform, glare-free illumination and magnification suitable to reliably distinguish agglutinated from non-agglutinated reactions.

## 2.2 Engineering Technical Specifications

- Configuration: Bench-top viewer with concave reflecting mirror or equivalent optical arrangement, integrated illumination, and magnifier.

- Mirror: High-reflectivity concave mirror with position-adjustable tilt/pivot mounting.

- Illumination: Integrated LED light source, cool operation, long life (minimum 20,000 hours typical), with adjustable brightness.

- Color temperature: Neutral daylight, typical 5000—6500 K, suitable for accurate reading of hemolysis and agglutination.

- Magnifier: Optical magnifying lens, minimum 2× magnification or equivalent.

- Tube holder: Removable rack accommodating standard 10—13 mm blood bank test tubes.

- Enclosure: Chemical-resistant, easy-to-clean housing.

- Power supply: 220—240 VAC, 50/60 Hz via medical-grade adapter; low-voltage DC operation at the viewer.

- Weight: Approximately 1—3 kg.

- Applicable standards: IEC 61010-1, IEC 61326 (EMC), CE marked or equivalent.

## 2.3 Standard Accessories

- Concave Mirror Assembly — High-reflectivity front-surface or equivalent concave mirror with adjustable tilt pivot for optimal viewing angle.

- LED Illumination Module — Long-life LED array with adjustable brightness control, uniform light distribution, low heat emission.

- Magnifying Lens Assembly — Optical-grade lens providing minimum 2× magnification, anti-scratch coating.

- Removable Test Tube Rack — Chemical-resistant material, holes sized for standard blood bank tubes, easy to clean and disinfect.

- AC Power Adapter and Cord — Medical-grade external adapter compatible with local mains supply.

## 2.4 Installation Requirements

Standard laboratory bench with single-phase 220—240 VAC mains outlet. No special HVAC or plumbing.

## 2.5 Documentation

User Manual, Warranty Certificate, Electrical Safety Test Report, CE or equivalent regulatory certificate.

## 2.6 Training

Basic application training for blood bank staff on correct use and cleaning.

## 2.7 Warranty

Minimum two (2) years comprehensive warranty covering parts and labour.

# 3. Blood Bank Refrigerator (2—6 °C, 500-Bag Capacity)

Quantity: 3

## 3.1 Clinical & Functional Specifications

Purpose-built medical-grade refrigerator intended exclusively for the storage of whole blood and packed red cell units at controlled temperature within 2—6 °C, in accordance with international blood banking standards. The unit shall maintain uniform temperature throughout the storage compartment, provide continuous monitoring and recording of temperature, and provide alarms in the event of temperature excursion or power failure. It shall not be used for reagents, food, or any non-blood item.

## 3.2 Engineering Technical Specifications

- Configuration: Upright, forced-air-cooled, medical blood bank refrigerator with glass or solid door.

- Storage capacity: Suitable for minimum 500 standard 450 mL blood bags (or as configured by manufacturer for equivalent capacity).

- Temperature range: Continuously maintained at 4 °C ± 1 °C within the operational range of 2—6 °C.

- Temperature uniformity: Within ±1.5 °C across all shelves under steady-state loaded conditions.

- Cooling system: CFC-free/HFC refrigerant, hermetically sealed compressor with forced-air circulation.

- Insulation: High-density CFC-free polyurethane foam.

- Interior: Corrosion-resistant stainless steel or equivalent, with adjustable/roll-out drawers.

- Drawers: Minimum 4 drawers (or as per capacity), stainless steel, roll-out on rails, transparent front shields.

- Door: Double/triple-pane tempered safety glass (or insulated solid) with magnetic gasket seal, lockable.

- Temperature control: Microprocessor-based with digital LED/LCD display, resolution 0.1 °C.

- Temperature recording: Continuous 7-day circular chart recorder AND/OR integrated digital data logger with USB export.

- Alarms: Audible and visual alarms for high/low temperature, door open, power failure, sensor failure, and low battery.

- Battery backup for alarm: Minimum 24 hours for alarm and monitoring.

- Power failure protection: Automatic restart with parameter retention.

- Interior lighting: Automatic on door opening, LED type.

- Interior access port: Provision for external reference thermometer probe.

- Communication: USB port and/or RS-232/Ethernet for data download and remote monitoring capability.

- Power supply: 220—240 VAC, 50/60 Hz, single phase.

- Operating ambient: 10—32 °C, up to 80% RH non-condensing.

- Noise level: Typical ≤55 dB(A) at 1 m.

- Applicable standards: IEC 61010-1, IEC 60601-1 where applicable, IEC 62366, IEC 61326 (EMC), EN 60068 environmental, CE-IVD or FDA cleared, compliant with AABB / Council of Europe guidelines for blood storage.

## 3.3 Standard Accessories

- Microprocessor Control & Display Console — Digital microprocessor controller with backlit LED/LCD display of chamber temperature, set points, and alarm status; resolution 0.1 °C.

- Acoustic/Visual Alarms with Battery Backup — Integrated multi-parameter alarm with rechargeable battery providing minimum 24 hours alarm autonomy on mains failure.

- Continuous Chart Recorder — 7-day circular chart temperature recorder with quartz movement, supplied with one-year supply of paper charts and marker pens.

- Roll-out Stainless Steel Drawers — Adjustable, corrosion-resistant stainless steel drawers on smooth roller rails, with clear plastic front shields for visual inventory.

- Dual-Pane Tempered Glass Door with Lock — Double/triple glazed tempered safety glass door with heated frame (where applicable) to prevent condensation, magnetic gasket, and cylinder key lock.

- USB Data Logging Interface — Integrated USB port allowing download of digital temperature history to standard PC in CSV or equivalent open format.

## 3.4 Installation Requirements

Dedicated single-phase 220—240 VAC, 15 A minimum, connected to essential/emergency power circuit and preferably UPS-backed for alarm continuity. Level floor with minimum 100 mm clearance at rear and sides for ventilation. Ambient temperature 10—32 °C. No direct sunlight or proximity to heat sources.

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Temperature Mapping/Validation Report, Calibration Certificate for sensors traceable to national/international standards, Electrical Safety Test Report, CE/FDA regulatory certificate, ISO 13485 certificate, Installation and Commissioning Report.

## 3.6 Training

Application training for blood bank staff on operation, alarm response, temperature record review; biomedical engineering training on preventive maintenance and basic troubleshooting.

## 3.7 Warranty

Minimum two (2) years comprehensive warranty on complete unit including compressor (minimum 5 years on compressor recommended), covering parts, labour, preventive maintenance visits, and breakdown maintenance.

# 4. Plasma Freezer (−30 °C / −40 °C)

Quantity: 1

## 4.1 Clinical & Functional Specifications

Medical-grade upright freezer dedicated to the rapid freezing and long-term storage of fresh frozen plasma (FFP) and plasma-derived components at temperatures of −30 °C or below, in compliance with international blood banking guidelines. The unit shall provide uniform low-temperature storage, continuous monitoring and recording, and alarms for temperature excursion and power failure.

## 4.2 Engineering Technical Specifications

- Configuration: Upright medical plasma freezer, forced-air cooled.

- Operating temperature: Adjustable within −20 °C to −40 °C, factory set point −30 °C or −40 °C as specified.

- Temperature uniformity: Within ±3 °C at steady state.

- Storage capacity: Suitable for a clinically useful number of plasma bags (typically minimum 200 units of 250—300 mL bags, or equivalent).

- Cooling system: CFC-free hermetically sealed compressor, single- or cascade-stage as required to reach set-point.

- Insulation: High-density CFC-free polyurethane foam.

- Interior: Stainless steel or equivalent corrosion-resistant material with roll-out drawers/inventory racks with internal dividers.

- Door: Insulated solid door with heated perimeter gasket to prevent frost build-up and vacuum lock.

- Locks: Dual security cylinder locks with master-key configuration.

- Temperature control: Microprocessor-based with digital LCD display, 0.1 °C resolution.

- Temperature recording: Continuous circular chart recorder and/or integrated digital data logger with USB or Ethernet download.

- Alarms: Audible/visual for high/low temperature, door open, power failure, sensor failure, and battery low; backup battery minimum 24 hours for alarms.

- Communication: USB and/or RS-232/Ethernet interface.

- Power supply: 220—240 VAC, 50/60 Hz, single phase.

- Ambient operating range: 10—32 °C.

- Applicable standards: IEC 61010-1, IEC 61326 (EMC), CE marked or FDA cleared, ISO 13485; compliant with AABB/EU blood component storage guidelines.

## 4.3 Standard Accessories

- Digital Microprocessor Temperature Control Panel — Backlit LCD display showing chamber temperature, set point, and alarm status; adjustable set points with password protection.

- Acoustic/Visual Temperature and Power-Failure Alarms — Integrated multi-parameter alarms with backup battery providing minimum 24 hours autonomy.

- Continuous Chart or Digital Data Logger — 7-day circular chart recorder with quartz movement and/or digital data logger with minimum 1-year internal memory and USB export.

- Heavy-Duty Pull-Out Stainless Steel Inventory Drawers — Corrosion-resistant, with internal dividers for organized plasma bag storage.

- Heated Perimeter Door Gaskets — Electrically warmed gasket assembly preventing frost formation and door vacuum lock.

- Dual Cylinder Security Locks — Two independent cylinder locks with master-key configuration for controlled access.

## 4.4 Installation Requirements

Dedicated single-phase 220—240 VAC, 15 A minimum, on essential/emergency power circuit with UPS backup for alarm continuity. Level floor, minimum 150 mm rear/side clearance for ventilation. Ambient temperature 10—32 °C. Avoid direct sunlight and heat sources.

## 4.5 Documentation

User Manual, Service Manual, Parts Catalogue, Temperature Mapping/Validation Report, Calibration Certificate traceable to international standards, Electrical Safety Test Report, CE/FDA certificate, ISO 13485 certificate, Installation & Commissioning Report.

## 4.6 Training

Application training for blood bank staff; biomedical engineering training on preventive maintenance, defrost cycle, alarm handling, and troubleshooting.

## 4.7 Warranty

Minimum two (2) years comprehensive warranty on complete unit, with minimum 5-year warranty on compressor recommended.

# 5. Platelet Incubator with Agitator

Quantity: 1

## 5.1 Clinical & Functional Specifications

Dedicated platelet storage system consisting of a temperature-controlled incubator with an integrated horizontal agitator, intended for continuous gentle agitation and controlled-temperature storage of platelet concentrates at 20—24 °C in accordance with international blood banking guidelines. The system shall maintain platelet viability and function by providing uniform temperature and continuous gentle motion.

## 5.2 Engineering Technical Specifications

- Configuration: Combined temperature-controlled incubator cabinet with internal horizontal reciprocating (flat-bed) agitator.

- Temperature range: Adjustable within 20—24 °C; factory set point typically 22 °C ± 1 °C.

- Temperature uniformity: Within ±1 °C across all shelves under steady state.

- Agitation: Horizontal side-to-side reciprocating motion, typical 60—70 cycles per minute, adjustable/factory preset.

- Agitation amplitude: Approximately 35—50 mm, as per international standard for platelet agitation.

- Capacity: Minimum 3 agitator shelves accommodating a clinically useful number of standard platelet bags (typical minimum 48 single-donor bags total or equivalent).

- Interior: Corrosion-resistant material, forced-air circulation for uniform heat distribution.

- Door: Tempered safety glass viewport with continuous magnetic seal.

- Temperature control: Microprocessor-based digital control with LED/LCD display, resolution 0.1 °C.

- Recording: Integrated circular chart recorder and/or digital data logger with USB export.

- Alarms: Audible/visual for high/low temperature, agitation failure (motion stop), door open, power failure, sensor failure; battery backup ≥24 hours for alarms.

- Communication: USB and/or Ethernet for data download.

- Power supply: 220—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61326, CE-IVD or FDA cleared, ISO 13485.

## 5.3 Standard Accessories

- Microprocessor-Controlled Horizontal Agitator Rack — Reciprocating flat-bed agitator with side-to-side motion at 60—70 cpm, amplitude \~35—50 mm, quiet DC motor drive.

- Motion & Temperature Alarm System — Integrated digital alarm module detecting agitation failure and temperature deviation with audible/visual indication and backup battery.

- Circular Chart Temperature Recorder — 7-day quartz-driven recorder with paper chart pads, supplied with one-year consumable supply.

- Forced-Air Circulation Fan System — Internal fan(s) providing uniform heat distribution throughout the chamber.

- Tempered Safety Glass Viewport Door — Insulated tempered glass door with continuous magnetic gasket seal for thermal integrity.

- Removable Stainless Steel Agitator Shelves — Perforated stainless steel or wire-mesh shelves for uniform airflow around platelet bags, easily removable for cleaning.

## 5.4 Installation Requirements

Dedicated 220—240 VAC single-phase supply, essential/UPS-backed. Level floor with adequate ventilation clearance. Ambient 10—32 °C. Avoid direct sunlight.

## 5.5 Documentation

User Manual, Service Manual, Parts Catalogue, Temperature Mapping Report, Calibration Certificate, Electrical Safety Test Report, CE/FDA certificate, ISO 13485 certificate.

## 5.6 Training

Application training for staff; biomedical engineering PPM and troubleshooting training.

## 5.7 Warranty

Minimum two (2) years comprehensive warranty.

# 6. Water Bath / Heat Block

Quantity: 2

## 6.1 Clinical & Functional Specifications

Laboratory-grade temperature-controlled bath and/or dry heat block intended for controlled heating of serological reagents, samples, and small volumes of fluids for blood bank testing (e.g., incubation of tubes at 37 °C, thawing controls). The unit shall provide uniform temperature, precise digital control, and safety cut-off.

## 6.2 Engineering Technical Specifications

- Configuration: Bench-top water bath and/or dry heat block with interchangeable aluminum blocks.

- Temperature range: Ambient +5 °C to 100 °C, adjustable.

- Setting resolution: 0.1 °C.

- Uniformity: Within ±0.5 °C at 37 °C.

- Stability: Within ±0.2 °C at set point.

- Capacity: Minimum 4—8 liters (water bath) or equivalent block sizes for standard tube diameters (10, 12, 15 mm).

- Chamber: Stainless steel (SS 304 minimum).

- Cover: Removable stainless steel gable cover.

- Safety: Over-temperature cut-off, dry-boil protection.

- Display: Digital LED, high visibility.

- Power supply: 220—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61326, CE marked.

## 6.3 Standard Accessories

- Digital Microprocessor Temperature Controller — High-visibility LED display, 0.1 °C resolution, PID control.

- Removable Stainless Steel Gable Cover — SS 304, gable/pitched design to reduce evaporative loss and prevent condensate dripping onto samples.

- Perforated Bottom Support Plate — Stainless steel perforated plate for tube rack support and uniform heating.

- Interchangeable Aluminum Heating Blocks — Set of anodized aluminum blocks for common tube diameters (10, 12, 15 mm) for dry-block operation.

- Over-Temperature Safety Cut-off Module — Independent thermal cut-off with visual/audible alarm.

## 6.4 Installation Requirements

Bench-top, 220—240 VAC single-phase mains, distilled water source for water bath top-up.

## 6.5 Documentation

User Manual, Calibration Certificate, Electrical Safety Test Report, CE certificate.

## 6.6 Training

Basic operator training on operation, cleaning, calibration verification.

## 6.7 Warranty

Minimum two (2) years comprehensive warranty.

# 7. Binocular Microscope

Quantity: 4

## 7.1 Clinical & Functional Specifications

Laboratory binocular biological microscope intended for routine microscopic examination of blood smears, cross-matched samples, agglutination reactions, and general hematology and blood-bank slide examination. The microscope shall provide clear, high-resolution imaging across low, medium, and high magnifications including oil immersion, with ergonomic viewing and stable mechanical performance.

## 7.2 Engineering Technical Specifications

- Configuration: Upright binocular biological microscope with Koehler or pre-centered LED illumination.

- Optical system: Infinity-corrected or achromatic finite-conjugate optical system.

- Head: Binocular, inclined 30° (typical), 360° rotatable.

- Interpupillary adjustment: Within 48—75 mm.

- Diopter adjustment: On at least one eyepiece.

- Eyepieces: Widefield 10× with minimum field number 18, focusable.

- Objectives: Achromatic 4×, 10×, 40× (spring-loaded), 100× oil immersion (spring-loaded).

- Nosepiece: Quadruple revolving nosepiece, ball-bearing mounted.

- Stage: Mechanical X-Y stage, minimum 120 × 130 mm, travel ≥75 × 50 mm, with slide holder.

- Focus: Coaxial coarse and fine, fine graduation ≤2 µm, with tension adjustment and upper limit stopper.

- Condenser: Abbe condenser NA ≥1.25 with iris diaphragm and blue filter.

- Illumination: Built-in LED, minimum 3 W, adjustable brightness, color temperature \~5500—6500 K, service life ≥20,000 hours.

- Power supply: Universal input 100—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, CE marked, ISO 9001 manufacturing.

## 7.3 Standard Accessories

- Binocular Head with Diopter Adjustment — Inclined 30°, 360° rotatable, interpupillary 48—75 mm, at least one eyepiece with diopter compensation.

- Widefield Eyepieces (10×, 1 pair) — Focusable 10× widefield eyepieces, field number ≥18, anti-fungal coating.

- Achromatic Objective Lens Set (4×, 10×, 40×, 100× oil) — Parfocal, parcentric, DIN standard threading, 40× and 100× spring-loaded.

- Coaxial Coarse and Fine Focusing Knobs — Concentric on both sides, with tension adjustment and upper focus stopper.

- Mechanical X-Y Stage with Slide Holder — Rack-and-pinion X-Y drive, coaxial controls, integrated slide clamp.

- Abbe Condenser with Iris and Blue Filter — NA ≥1.25, height-adjustable, with iris diaphragm and swing-in blue daylight filter.

- Built-in LED Light Source with Brightness Regulator — Long-life LED with rotary brightness control.

- Immersion Oil Bottle and Dust Cover — Non-drying synthetic immersion oil, nD 1.515; vinyl dust cover.

## 7.4 Installation Requirements

Stable vibration-free bench, 100—240 VAC mains outlet, ambient temperature 10—35 °C, low-dust environment.

## 7.5 Documentation

User Manual, Optical Data Sheet, Warranty Certificate, CE certificate.

## 7.6 Training

Basic user training on operation, cleaning of optics, and objective handling.

## 7.7 Warranty

Minimum two (2) years comprehensive warranty on mechanical and optical components; minimum 5 years on LED module recommended.

# 8. Donor Blood Collection Chair

Quantity: 4

## 8.1 Clinical & Functional Specifications

Ergonomic, adjustable donor chair intended for whole-blood and apheresis donor comfort and safety during blood collection. The chair shall enable rapid conversion to the Trendelenburg (head-down) position for management of vasovagal reactions, and shall provide adjustable armrests, leg rest, and neck support to accommodate donors of varying body size.

## 8.2 Engineering Technical Specifications

- Configuration: Multi-position donor chair with motorized or hydraulic height adjustment and Trendelenburg capability.

- Frame: Heavy-gauge steel, corrosion-resistant epoxy powder-coated finish.

- Height adjustment: Motorized (electric linear actuators) or hydraulic; adjustable range within approximately 550—900 mm seat height.

- Backrest tilt: Approximately 0—70° adjustable.

- Leg rest: Adjustable, capable of full Trendelenburg positioning.

- Trendelenburg angle: Minimum 12° head-down tilt.

- Load capacity: Minimum safe working load 180 kg.

- Armrests: Dual adjustable padded armrests with multi-axis swivel and locking, height and angle adjustable.

- Upholstery: Seamless antimicrobial fluid-proof vinyl (or equivalent), latex-free, resistant to blood, disinfectants, and IPA.

- Casters: Four heavy-duty locking swivel castors, minimum 75 mm diameter.

- Control: Handheld wired remote pendant with clearly labelled functions; foot switches optional.

- Emergency: Manual emergency Trendelenburg release/override on power failure.

- Power supply (motorized): 100—240 VAC, 50/60 Hz, IEC 60601-1 medical grade.

- Applicable standards: IEC 60601-1 (medical electrical safety), IEC 60601-1-2 (EMC), IEC 60601-2-52 where applicable, CE MDR or FDA cleared, ISO 13485.

## 8.3 Standard Accessories

- Motorized/Hydraulic Height Adjustment with Trendelenburg — Electric actuator or sealed hydraulic system providing smooth height adjustment and rapid Trendelenburg tilt with manual override.

- Adjustable Padded Armrests — Dual armrests with multi-axis swivel joints and lockable positioning, upholstered with same antimicrobial vinyl as seat.

- Neck Support Pillow and Leg Rest Cushion — Removable, high-density foam, antimicrobial vinyl cover, washable.

- Handheld Wired Remote Pendant — Ergonomic pendant with clearly labelled buttons for all motorized functions; IPX rating suitable for wipe-cleaning.

- Antimicrobial Fluid-Proof Vinyl Upholstery — Seamless, latex-free, chemical- and disinfectant-resistant, tested for microbial resistance.

- Heavy-Duty Locking Caster Wheels — Four minimum 75 mm swivel casters, at least two with total-lock brakes.

## 8.4 Installation Requirements

Level floor; single-phase 220—240 VAC outlet for motorized version; adequate space for full recline and Trendelenburg.

## 8.5 Documentation

User Manual, Service Manual, Electrical Safety Test Report, CE/FDA certificate.

## 8.6 Training

Operator training on positioning, emergency Trendelenburg, and cleaning.

## 8.7 Warranty

Minimum two (2) years comprehensive warranty covering frame, upholstery, actuators, and controls.

# 9. Blood Collection Monitor / Mixer

Quantity: 4

## 9.1 Clinical & Functional Specifications

Portable electronic device intended to continuously mix (rock) blood as it is collected into the donor bag and to accurately measure and control the collected volume by weight. The device shall automatically stop or alert at a preset target volume, thereby ensuring anticoagulant-to-blood ratio and preventing over- or under-collection.

## 9.2 Engineering Technical Specifications

- Configuration: Bench-top / trolley-mounted electronic collection monitor with integrated rocking cradle and load-cell weighing system.

- Weighing range: 0—999 g (minimum), typical up to 1500 g.

- Weight resolution: 1 g.

- Weight accuracy: Within ±1% of reading or ±1 g, whichever is greater.

- Volume settings: Programmable target volume (typical 200—600 mL) with volume-to-weight conversion factor selectable for whole blood/plasma.

- Rocking motion: Continuous rocking cradle, typical 25—35 cycles/min, smooth motion.

- Automatic clamp: Integrated line clamp automatically closing at target volume (where equipped) with acoustic completion alarm.

- Display: Digital LCD/LED showing weight, volume, time, and flow rate.

- Alarms: Slow flow, no flow, target reached, low battery.

- Battery: Internal rechargeable battery, minimum 4 hours continuous operation between charges.

- Charging time: Maximum 8 hours to full charge.

- Power supply: 100—240 VAC, 50/60 Hz via medical adapter.

- Enclosure: Impact-resistant, easy to clean.

- Weight: Approximately 3—5 kg.

- Applicable standards: IEC 60601-1, IEC 60601-1-2, CE MDR or FDA cleared, ISO 13485.

## 9.3 Standard Accessories

- Motorized Rocking Cradle Platform — Smooth reciprocating rocker for continuous gentle mixing of collected blood at 25—35 cpm.

- Digital Control Console — Backlit LCD/LED display of collected volume, weight, elapsed time, and flow status.

- Automated Target-Volume Clamp — Solenoid- or motor-actuated tubing clamp with acoustic completion alarm to prevent over-collection.

- Internal Rechargeable Battery — Sealed lithium-ion or NiMH pack providing minimum 4 hours continuous operation.

- AC Power Supply Adapter and Cord — Medical-grade external adapter, universal input 100—240 VAC.

- Hard Transport Case — Rigid impact-resistant carrying case with custom-fitted foam inserts for equipment and accessories.

## 9.4 Installation Requirements

Portable device; 220—240 VAC outlet required for charging. No special HVAC.

## 9.5 Documentation

User Manual, Calibration Certificate for load cell traceable to national standards, Electrical Safety Test Report, CE/FDA certificate.

## 9.6 Training

Operator training on setup, calibration verification, alarm handling.

## 9.7 Warranty

Minimum two (2) years comprehensive warranty including battery.

# 10. Plasma Thawing Bath

Quantity: 1

## 10.1 Clinical & Functional Specifications

Dedicated dry- or wet-type thawing device intended for the controlled, rapid, and safe thawing of fresh frozen plasma (FFP), cryoprecipitate, and other frozen blood components at 37 °C in accordance with international blood banking standards, without damage to plasma proteins or contamination of the bag ports.

## 10.2 Engineering Technical Specifications

- Configuration: Water bath or dry-thaw system with microprocessor temperature control.

- Temperature: Precisely maintained at 37 °C ± 1 °C.

- Uniformity: Within ±1 °C throughout working volume.

- Capacity: Minimum 4 standard plasma bags simultaneously.

- Thawing time: Typical 15—20 minutes for standard 250—300 mL plasma bag.

- Water circulation (wet type): Integrated pump for uniform heat distribution.

- Basket/holder: Removable stainless steel dividers/holders for plasma bags, keeping bags separated and ports above water level.

- Timer: Digital cycle timer with acoustic completion alarm.

- Safety: Over-temperature cut-off, low-water level protection.

- Drainage: Water drain valve with hose connection.

- Housing: Stainless steel chamber and outer housing.

- Power supply: 220—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61326, CE MDR or FDA cleared, ISO 13485.

## 10.3 Standard Accessories

- Microprocessor Temperature Controller — Calibrated to 37 °C ± 1 °C with digital display and adjustable set point (limited range) with password protection.

- Removable Stainless Steel Basket Dividers — Corrosion-resistant dividers for organized bag placement, easy to remove for cleaning and disinfection.

- Water Circulation Pump Module — Integrated pump for uniform temperature distribution throughout the water volume.

- Digital Cycle Timer with Acoustic Alarm — Backlit digital display with programmable cycle time and end-of-cycle audible alarm.

- Drainage Outlet Valve with Hose — Bottom drain valve with flexible hose for periodic water change and cleaning.

## 10.4 Installation Requirements

Bench-top; 220—240 VAC supply; access to distilled water for filling; nearby drain or waste vessel for emptying.

## 10.5 Documentation

User Manual, Service Manual, Calibration Certificate, Electrical Safety Test Report, CE/FDA certificate.

## 10.6 Training

Operator training on operation, water management, disinfection, and calibration verification.

## 10.7 Warranty

Minimum two (2) years comprehensive warranty.

# 11. Tube Sealer for Blood Bags

Quantity: 2

## 11.1 Clinical & Functional Specifications

Electronic device intended for the hermetic sealing of PVC blood bag tubing during collection, component preparation, and transfusion, providing a sterile leak-proof seal without cutting the tubing or generating significant heat that could damage cellular components.

## 11.2 Engineering Technical Specifications

- Configuration: Handheld or bench-top RF (radio-frequency) or equivalent electronic tube sealer.

- Sealing technology: RF dielectric sealing or equivalent internationally accepted method.

- Tube compatibility: Standard PVC blood bag tubing (typical 3—5 mm OD).

- Sealing time: Typical 1—3 seconds per seal.

- Seal integrity: Hermetic, leak-proof under standard blood bank pressure testing.

- Cycles between charges (portable version): Minimum 300 seals per full charge.

- Battery: Internal rechargeable lithium-ion for cordless operation.

- Charging time: Maximum 4 hours to full charge.

- Safety: Sealing head trigger sensor to prevent accidental activation; safety splash shield over jaws.

- Housing: Impact- and chemical-resistant, ergonomically designed.

- Power supply: 100—240 VAC, 50/60 Hz via charging dock.

- Weight (handheld): Approximately 0.5—1.5 kg.

- Applicable standards: IEC 60601-1, IEC 60601-1-2, IEC 61010 as applicable, CE MDR or FDA cleared, ISO 13485.

## 11.3 Standard Accessories

- RF Sealing Head with Automated Trigger — Radio-frequency electrode head with automatic sensor-triggered sealing cycle.

- Handheld Ergonomic Sealing Wand — Lightweight sealing wand with flexible coiled cable (corded models) or fully cordless design.

- Clear Plastic Safety Splash Shield — Transparent shield over sealing jaws for operator protection.

- Internal Rechargeable Lithium-Ion Battery — Sealed rechargeable pack providing minimum 300 seals per charge.

- Desktop Charging Dock — Wall-mountable or bench-top charging cradle with universal 100—240 VAC input.

- Rigid Transit Storage Container — Impact-resistant carrying case with foam insert.

## 11.4 Installation Requirements

Bench-top charging; 220—240 VAC outlet. No special HVAC.

## 11.5 Documentation

User Manual, Service Manual, Electrical Safety Test Report, CE/FDA certificate.

## 11.6 Training

Operator training on operation, safety, and battery management.

## 11.7 Warranty

Minimum two (2) years comprehensive warranty including battery.

# 12. Validated Blood Transport Box

Quantity: 4

## 12.1 Clinical & Functional Specifications

Passively cooled, temperature-validated transport container intended for the safe transfer of blood and blood components between blood bank, wards, operating theatres, and outreach centres, maintaining product temperature within required range (2—10 °C for red cells; other ranges as applicable) for a validated period.

## 12.2 Engineering Technical Specifications

- Configuration: Rigid insulated transport box with locking lid and integrated temperature display.

- Outer shell: High-density polyurethane insulated container with impact-resistant outer skin.

- Capacity: Suitable for a clinically useful number of blood bags (typical minimum 6—12 units of 450 mL).

- Temperature maintenance: Validated to maintain internal temperature within 2—10 °C for minimum 24 hours at ambient 25 °C with proper pre-conditioned gel packs.

- Insulation performance: Validation report to be supplied.

- Temperature indicator: Integrated digital LCD temperature display on exterior lid, resolution 0.1 °C, min/max memory.

- Sensor: NTC or equivalent probe with independent battery, minimum 12-month operation on single battery.

- Latches: Secure snap or lever latches, tamper-evident where possible.

- Handles/straps: Integrated side handles and adjustable shoulder carrying strap.

- Interior dividers: Plastic dividers/partitions to separate bags from ice packs and each other.

- Cleaning: Fully wipe-cleanable with hospital disinfectants.

- Applicable standards: ISTA 7D or equivalent thermal validation, CE marked where applicable, compliant with WHO/AABB guidelines for blood transport.

## 12.3 Standard Accessories

- High-Density Polyurethane Insulated Shell — Rigid impact-resistant outer skin with secure latches and validated thermal insulation.

- Validated Reusable Phase-Change Gel Packs — Complete set of pre-conditioned gel packs sized for the container, reusable, with pre-conditioning instructions.

- Integrated Digital Temperature Display — Exterior-mounted LCD indicator with min/max memory, resolution 0.1 °C.

- Internal Plastic Dividers — Removable partitions separating bags from direct ice contact and preventing bag movement.

- Shoulder Carrying Strap and Side Handles — Heavy-duty synthetic fabric strap with adjustable length; integrated ergonomic side handles.

## 12.4 Installation Requirements

Storage in dry, clean area between uses. Access to freezer for pre-conditioning gel packs.

## 12.5 Documentation

User Manual, Thermal Validation Report, Cleaning/Disinfection Instructions, Warranty Certificate.

## 12.6 Training

Operator training on pre-conditioning, packing, temperature monitoring, and cleaning.

## 12.7 Warranty

Minimum two (2) years comprehensive warranty on container, latches, and electronic components.

# 13. Bench-Top Centrifuge

Quantity: 2

## 13.1 Clinical & Functional Specifications

General-purpose laboratory centrifuge intended for the separation of blood components in tubes, including serum/plasma separation, sample preparation for immunohematology, and routine pre-analytical processing.

## 13.2 Engineering Technical Specifications

- Configuration: Bench-top centrifuge with interchangeable swing-out and/or fixed-angle rotor.

- Maximum speed: Minimum 4000 rpm (swing-out) / minimum 5000 rpm (fixed-angle).

- Maximum RCF: Minimum 3000 × g.

- Speed setting: Digitally adjustable in 10 rpm/1 × g steps.

- Speed accuracy: Within ±1% of set value.

- Timer: 0—99 minutes digital, plus continuous mode, with automatic braking.

- Rotor: Swing-out rotor with minimum 4 × 50 mL capacity (or matched buckets for blood bank tubes); adapters for 10, 12, 15 mL tubes.

- Acceleration/braking: Multiple selectable ramp profiles.

- Imbalance detection: Automatic electronic sensing with cycle abort.

- Lid safety: Electromagnetic interlock, no lid opening during rotation, no start with lid open.

- Emergency lid release: Manual override on power failure.

- Display: Digital LCD/LED showing speed/RCF, time, and program.

- Noise: ≤60 dB(A) at maximum speed.

- Power supply: 220—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61010-2-020, IEC 61326, CE marked.

## 13.3 Standard Accessories

- Interchangeable Swing-Out or Fixed-Angle Rotor Head — Corrosion-resistant metal rotor for blood tubes, autoclavable buckets where applicable.

- Tube Buckets/Adapters — Complete set of buckets/adapters for standard blood banking tubes (10, 12, 15 mL and blood-bag pilot tubes).

- Digital Microprocessor Control Interface — RPM/RCF and run-time programmable with backlit digital display.

- Automated Electronic Lid-Lock and Imbalance Sensor — Motorized/electromagnetic lid lock with independent imbalance detection.

- Main Power Cable with Fuse — Detachable IEC power cord with integrated protective fuse.

## 13.4 Installation Requirements

Sturdy vibration-free bench, 220—240 VAC, 10 A minimum outlet, adequate ventilation clearance around unit.

## 13.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate (speed and timer), Electrical Safety Test Report, CE certificate.

## 13.6 Training

Operator training on rotor selection, balancing, program setup, and safety.

## 13.7 Warranty

Minimum two (2) years comprehensive warranty; motor warranty minimum 5 years recommended.

# 14. Incubator

Quantity: 1

## 14.1 Clinical & Functional Specifications

General-purpose laboratory incubator intended for controlled-temperature incubation of serological reactions, samples, and reagents in blood bank workflows, at temperatures typically between ambient and 60 °C.

## 14.2 Engineering Technical Specifications

- Configuration: Bench-top or floor-standing forced-air or gravity-convection incubator.

- Chamber capacity: Minimum 50 L (or as specified).

- Temperature range: Ambient +5 °C to 60 °C, adjustable.

- Setting resolution: 0.1 °C.

- Uniformity: Within ±1 °C at 37 °C.

- Stability: Within ±0.5 °C at set point.

- Interior: Stainless steel (SS 304) with rounded corners for easy cleaning.

- Insulation: CFC-free high-density insulation.

- Shelves: Minimum 3 adjustable perforated stainless steel shelves.

- Door: Outer insulated door and inner tempered glass viewing door with magnetic seal.

- Control: Microprocessor PID controller with digital LED display.

- Safety: Independent over-temperature cut-off with audible/visual alarm.

- Power supply: 220—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61326, CE marked.

## 14.3 Standard Accessories

- Digital Microprocessor Temperature Controller — PID controller with LED display, 0.1 °C resolution, programmable set point.

- Adjustable Perforated Stainless Steel Shelves — Set of 3 SS 304 perforated shelves for uniform air circulation.

- Tempered Glass Inner Viewing Door — Full-perimeter magnetic seal for thermal integrity while permitting sample observation without opening the outer door.

- Over-Temperature Alarm Module — Independent thermostat with visual and audible alarm.

- Medical-Grade Power Cord — Detachable IEC cord matched to local mains.

## 14.4 Installation Requirements

Level bench or floor, 220—240 VAC single-phase supply, ambient 10—32 °C, adequate ventilation.

## 14.5 Documentation

User Manual, Service Manual, Calibration Certificate, Electrical Safety Test Report, CE certificate.

## 14.6 Training

Operator training on operation, cleaning, and calibration verification.

## 14.7 Warranty

Minimum two (2) years comprehensive warranty.

# 15. Donor Blood Collection Bed

Quantity: 8

## 15.1 Clinical & Functional Specifications

Multi-section adjustable bed intended for donor comfort during whole-blood donation, permitting supine and semi-recumbent positioning, rapid conversion to Trendelenburg for management of vasovagal reactions, and easy access by phlebotomy staff.

## 15.2 Engineering Technical Specifications

- Configuration: Two- or three-section donor couch/bed with manually adjustable or gas-spring assisted backrest and leg section.

- Frame: Heavy-gauge steel with anti-rust epoxy powder-coated finish.

- Backrest adjustment: 0—70° minimum.

- Leg section adjustment: Independent, permitting Trendelenburg positioning.

- Trendelenburg angle: Minimum 12° head-down.

- Dimensions: Approximately 1900 × 650 mm sleeping surface (adult standard).

- Height: Fixed or manually adjustable within approximately 550—800 mm.

- Load capacity: Minimum safe working load 180 kg.

- Armboards: Dual wide upholstered armboards with locking knobs for phlebotomy.

- Mattress: High-density seamless vinyl-covered fluid-proof antimicrobial pad, minimum 75 mm thickness.

- Casters: Four locking swivel casters, minimum 100 mm diameter.

- IV pole socket: Universal receptacle at head end.

- Applicable standards: IEC 60601-2-52 principles where applicable, CE MDR or equivalent, ISO 13485.

## 15.3 Standard Accessories

- Powder-Coated Steel Frame — Multi-section frame with corrosion-resistant epoxy powder coating.

- Backrest and Leg Section Adjustment Levers — Manual or gas-spring assisted, smooth operation with reliable locking.

- Upholstered Wide Armboards — Dual armboards with padded upholstery and locking knobs for secure positioning.

- Seamless Antimicrobial Vinyl Mattress Pad — Fluid-proof high-density foam mattress with seamless antimicrobial cover.

- Locking Swivel Caster Wheels — Four heavy-duty casters, minimum two with total-lock brakes.

- Universal IV Pole Sockets — Receptacles at the head section accepting standard IV drip stands.

## 15.4 Installation Requirements

Level floor; no electrical connection for manual model.

## 15.5 Documentation

User Manual, Cleaning/Disinfection Instructions, Warranty Certificate, CE certificate.

## 15.6 Training

Operator training on adjustment, Trendelenburg positioning, and cleaning.

## 15.7 Warranty

Minimum two (2) years comprehensive warranty on frame, upholstery, and mechanical components.

# 16. Tube Stripper

Quantity: 1

## 16.1 Clinical & Functional Specifications

Handheld mechanical device intended for expressing (stripping) residual blood or plasma from the donor tubing segment into the primary blood bag prior to sealing, thereby mixing anticoagulant with the tubing content for later cross-matching.

## 16.2 Engineering Technical Specifications

- Configuration: Handheld mechanical roller-jaw device.

- Rollers: Spring-loaded self-adjusting rollers accommodating standard PVC blood bag tubing (3—5 mm OD).

- Handle: Ergonomic anti-slip texturized grip.

- Construction: Corrosion-resistant metal frame with polymer/rubber roller surfaces.

- Weight: Approximately 150—300 g.

- Cleaning: Wipe-cleanable with hospital disinfectants.

- Applicable standards: CE marked, ISO 13485 manufacturing.

## 16.3 Standard Accessories

- Mechanical Roller Jaw Assembly — Handheld frame with paired rollers for tube stripping.

- Spring-Loaded Self-Adjusting Rollers — Rollers with spring tension automatically accommodating standard tubing diameters.

- Anti-Slip Textured Grip Handles — Ergonomic handle surface for secure operation.

- Textile Storage Pouch — Protective slip case for storage and transport.

## 16.4 Installation Requirements

None.

## 16.5 Documentation

User Manual, Cleaning Instructions, CE certificate.

## 16.6 Training

Operator demonstration on correct use.

## 16.7 Warranty

Minimum one (1) to two (2) years comprehensive warranty.

# 17. Emergency Cabinet (Blood Bank Crash Cart)

Quantity: 1

## 17.1 Clinical & Functional Specifications

Mobile emergency trolley intended to hold resuscitation equipment, emergency drugs, and consumables required for immediate response to donor adverse reactions and other emergencies within the blood bank.

## 17.2 Engineering Technical Specifications

- Configuration: Mobile crash cart with multiple drawers, defibrillator shelf, IV pole, CPR board, sharps holder, and waste bins.

- Body: High-impact ABS and/or stainless steel construction with corrosion-resistant finish.

- Drawers: Minimum 4 drawers with central locking system, smooth ball-bearing slides, modular divider set.

- CPR board: Rigid resuscitation backboard mounted on rear.

- IV pole: Height-adjustable telescopic pole, stainless steel.

- Defibrillator shelf: Reinforced shelf with security strap.

- Sharps container bracket: Side-mounted holder.

- Waste bins: Dual color-coded medical waste bins.

- Casters: Four heavy-duty swivel casters, minimum two lockable, minimum 125 mm diameter.

- Load capacity: Minimum 100 kg total.

- Applicable standards: CE marked, ISO 13485 manufacturing.

## 17.3 Standard Accessories

- ABS/Stainless Steel Mobile Trolley Body — Impact- and chemical-resistant body with central locking mechanism.

- Rear-Mounted CPR Backboard — Rigid resuscitation board for chest compressions.

- Telescopic IV Pole — Height-adjustable stainless steel IV pole with hooks.

- Defibrillator Shelf with Security Strap — Reinforced shelf with retention strap.

- Sharps Container Bracket and Dual Waste Bins — Side-mounted brackets and color-coded medical waste receptacles.

- Modular Drawer Divider Set — Configurable dividers for organized medication and consumable storage.

## 17.4 Installation Requirements

Level floor; suitable storage location near donor area.

## 17.5 Documentation

User Manual, Warranty Certificate, CE certificate.

## 17.6 Training

Operator familiarization with layout and locking system.

## 17.7 Warranty

Minimum two (2) years comprehensive warranty.

# 18. Slide Storage Cabinet

Quantity: 1

## 18.1 Clinical & Functional Specifications

Dedicated storage cabinet intended for the systematic long-term archival of pathology and blood-bank glass slides, providing organized retrieval, protection from breakage and contamination, and secure locking access.

## 18.2 Engineering Technical Specifications

- Configuration: Multi-drawer sheet-steel cabinet.

- Capacity: Minimum 5,000 standard 25 × 75 mm slides (or as specified).

- Construction: Heavy-duty sheet steel with anti-scratch epoxy powder-coated finish.

- Drawers: Smooth-gliding pull-out drawers on ball-bearing slides, with front index card holders.

- Internal dividers: Vertical grids sized for standard pathology slides.

- Locking: Central cylinder key lock system with two duplicate keys.

- Base: Levelling feet.

- Applicable standards: ISO 9001 manufacturing.

## 18.3 Standard Accessories

- Sheet-Steel Cabinet Housing — Heavy-duty body with anti-scratch epoxy powder finish.

- Pull-Out Drawers with Index Card Holders — Smooth-gliding drawers with front card slots for identification.

- Vertical Divider Grids — Internal partitions sized for standard microscope slides.

- Central Cylinder Locking System — Single-key locking of all drawers, with duplicate keys supplied.

- Base Levelling Feet — Adjustable feet for stable placement on uneven floors.

## 18.4 Installation Requirements

Level floor; no electrical connection.

## 18.5 Documentation

User Manual, Warranty Certificate.

## 18.6 Training

Basic user familiarization.

## 18.7 Warranty

Minimum two (2) years warranty on structure and mechanical components.

# 19. Plasma Extractor

Quantity: 1

## 19.1 Clinical & Functional Specifications

Manual mechanical device intended for the separation and transfer of plasma or other supernatant components from centrifuged blood bags into satellite bags by controlled compression of the primary bag against a pressure plate.

## 19.2 Engineering Technical Specifications

- Configuration: Manual spring-loaded plate-type plasma extractor.

- Pressure plate: Transparent acrylic panel permitting visual monitoring during extraction.

- Base plate: Heavy-duty stainless steel with non-slip rubber feet.

- Compression mechanism: Spring-loaded with manual toggle lever lock for controlled and consistent pressure.

- Bag hanging pins: Integrated pins on support frame for satellite bag suspension.

- Compatibility: Standard 450 mL / 350 mL blood bag configurations.

- Cleaning: Fully wipe-cleanable with hospital disinfectants.

- Applicable standards: CE marked, ISO 13485 manufacturing.

## 19.3 Standard Accessories

- Spring-Loaded Transparent Acrylic Pressure Plate — Clear panel permitting visual monitoring of extraction.

- Stainless Steel Base Plate with Non-Slip Feet — Heavy-duty base with rubber feet for stability.

- Manual Toggle Lever Lock — Handle for engagement and locking of the compression plate.

- Integrated Blood Bag Hanging Pins — Support pins for satellite bag suspension during extraction.

## 19.4 Installation Requirements

Bench-top; no electrical connection.

## 19.5 Documentation

User Manual, Cleaning Instructions, CE certificate.

## 19.6 Training

Operator training on correct extraction technique.

## 19.7 Warranty

Minimum two (2) years comprehensive warranty.

# 20. Weighing Machine

Quantity: 1

## 20.1 Clinical & Functional Specifications

Electronic laboratory weighing scale intended for accurate weighing of blood bags, reagents, and samples in the blood bank, with tare and calibration functions and rechargeable operation.

## 20.2 Engineering Technical Specifications

- Configuration: Bench-top electronic platform scale with remote indicator.

- Capacity: Minimum 5 kg (typical 5—15 kg).

- Resolution: 1 g or better.

- Accuracy: Within ±0.1% of full scale or ±1 g, whichever is greater.

- Platform: Stainless steel deck with non-slip surface, minimum 200 × 200 mm.

- Indicator: Remote wall-mountable LCD display with flexible cord.

- Tare: Electronic tare across full range.

- Calibration: Automatic electronic calibration function with software support.

- Power supply: 100—240 VAC adapter with internal rechargeable backup battery, minimum 20 hours operation.

- Applicable standards: OIML R 76 or NTEP class III equivalent, IEC 61010-1, CE marked.

## 20.3 Standard Accessories

- Stainless Steel Platform Deck — Corrosion-resistant SS platform with non-slip coating.

- Remote Wall-Mounted Digital Indicator — LCD display on flexible cord for convenient positioning.

- Automated Electronic Tare/Calibration Function — Software-supported tare and calibration.

- AC Adapter and Rechargeable Backup Battery — Universal 100—240 VAC adapter and internal battery for uninterrupted operation.

## 20.4 Installation Requirements

Level bench; 220—240 VAC outlet.

## 20.5 Documentation

User Manual, Calibration Certificate traceable to national standards, Electrical Safety Test Report, CE certificate.

## 20.6 Training

Operator training on operation, tare, calibration verification.

## 20.7 Warranty

Minimum two (2) years comprehensive warranty including battery and load cell.

# 21. CLIA (Chemiluminescence Immunoassay) Machine

Quantity: 1

## 21.1 Clinical & Functional Specifications

Fully automated chemiluminescence immunoassay analyser intended for the qualitative and quantitative detection of transfusion-transmissible infection (TTI) markers and other immunoassay parameters relevant to blood bank screening, including but not limited to HIV Ag/Ab, HBsAg, anti-HCV, syphilis, and other markers as required. The system shall provide random-access, high-throughput, walk-away operation with primary-tube sampling, positive sample identification, integrated quality control, and LIS connectivity.

## 21.2 Engineering Technical Specifications

- Configuration: Fully automated bench-top or floor-standing chemiluminescence immunoassay analyser.

- Technology: Chemiluminescence immunoassay (CLIA) or equivalent internationally accepted technique.

- Throughput: Minimum 60 tests/hour typical (target ≥100 tests/hour for TTI screening applications).

- Loading: Continuous or batch loading; random-access architecture.

- Sample identification: Integrated barcode reader for positive sample and reagent identification.

- Sample type: Serum, plasma; primary-tube sampling with clot and level detection.

- Sample volume: Minimum sample volume typical 10—200 µL as per assay.

- Reagent storage: Onboard refrigerated reagent compartment maintained at 2—10 °C with independent temperature monitoring.

- Reagent capacity: Minimum 10 assay channels or equivalent onboard.

- Reaction chamber: Temperature-controlled at 37 °C ± 0.5 °C.

- Detector: Photomultiplier tube (PMT) or equivalent high-sensitivity photodetector.

- Calibration: Automated calibration with lot-specific parameters via barcode.

- QC: Integrated multi-level QC with Levy-Jennings and Westgard rule support.

- Water requirement: Deionized water supply with in-line or reservoir configuration.

- Waste: Segregated liquid and solid waste with automatic full-level detection.

- User interface: Integrated PC workstation with color medical display, minimum 19-inch diagonal.

- Software: Dedicated CLIA data analysis software with user access levels, audit trail, and QC management.

- Connectivity: Bi-directional LIS/HIS via HL7 or ASTM; Ethernet; USB.

- Cybersecurity: Compliant with IEC 80001 principles; user authentication; audit log.

- UPS: Dedicated UPS supplied for controlled shutdown.

- Power supply: 220—240 VAC, 50/60 Hz, single phase, dedicated circuit.

- Applicable standards: IEC 61010-1, IEC 61010-2-101 (IVD), IEC 61326, IEC 62304 (software), IEC 62366 (usability), ISO 13485, ISO 14971, CE-IVD or FDA cleared.

## 21.3 Standard Accessories

- Automated Sample Pipetting Arm — Robotic pipettor with liquid-level detection, clot detection, and crash sensor.

- Refrigerated Reagent Storage Compartment — Independent cooling zone maintained at 2—10 °C with continuous temperature monitoring.

- Integrated Washer Manifold Module — Microplate/tube washer with fluid lines and level sensors.

- Processing Computer Workstation and Medical Display — Dedicated PC with minimum 19-inch medical-grade color display, keyboard, and mouse.

- Dedicated CLIA Data Analysis Software License — Full analysis package with QC, calibration, patient results, and audit trail; perpetual or as specified license.

- Laser Report Printer and System UPS — Networked laser printer for reports and dedicated online UPS with runtime sufficient for controlled shutdown (minimum 20 minutes at full load).

## 21.4 Installation Requirements

Dedicated 220—240 VAC circuit with proper earthing, connected to online UPS; deionized water supply and waste drainage; ambient temperature 18—28 °C, 30—80% RH non-condensing; adequate bench/floor space with service clearance; network point for LIS connectivity.

## 21.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Installation Qualification (IQ), Operational Qualification (OQ), Calibration Certificate, Electrical Safety Test Report, CE-IVD/FDA certificate, ISO 13485 certificate, software validation summary.

## 21.6 Training

Comprehensive clinical/application training for laboratory technologists, biomedical engineering training, preventive maintenance and troubleshooting training, LIS interface training.

## 21.7 Warranty

Minimum two (2) years comprehensive warranty including labour, spare parts, preventive maintenance visits, breakdown maintenance, software updates, and technical support.

# 22. Refrigerated Centrifuge

Quantity: 1

## 22.1 Clinical & Functional Specifications

High-capacity refrigerated centrifuge intended for the temperature-controlled separation of whole blood into components (red cells, plasma, platelets, cryoprecipitate) directly from blood bags, and for temperature-sensitive sample processing.

## 22.2 Engineering Technical Specifications

- Configuration: Floor-standing high-capacity refrigerated centrifuge with swing-out rotor for blood bags.

- Maximum speed: Minimum 4000 rpm with blood-bag rotor.

- Maximum RCF: Minimum 5000 × g.

- Speed setting: Digitally adjustable in 10 rpm/1 × g steps, accuracy within ±1%.

- Temperature range: −10 °C to +40 °C (minimum), maintainable during high-speed operation.

- Temperature stability: Within ±1 °C of set point at steady state.

- Timer: 0—99 hours 59 minutes digital, plus continuous mode; automatic braking.

- Rotor: Swing-out rotor for minimum 4 × standard 450 mL blood bags (quad bag configuration) with matched buckets/cups (minimum 6 supplied).

- Acceleration/braking profiles: Multiple selectable ramps for delicate separation.

- Imbalance detection: Automatic with cycle abort.

- Lid safety: Motorized/electromagnetic interlock with manual emergency release on power failure.

- Refrigeration: CFC-free hermetically sealed compressor.

- User interface: Color touchscreen or backlit LCD/LED display.

- Programs: Minimum 50 user programs, password-protected.

- Communication: USB and/or Ethernet for program export and service.

- Power supply: 220—240 VAC, 50/60 Hz; or 380—415 VAC three-phase as specified; dedicated circuit rated for full load.

- Applicable standards: IEC 61010-1, IEC 61010-2-020, IEC 61326, CE marked, ISO 13485.

## 22.3 Standard Accessories

- High-Capacity Swing-Out Rotor Head — Corrosion-resistant swing-out rotor for quad blood-bag configuration.

- Heavy-Duty Blood Bag Buckets/Cups (Set of 6) — Matched buckets sized for standard 450 mL blood bags, autoclavable where applicable.

- CFC-Free Refrigeration System — Hermetically sealed compressor maintaining chamber temperature down to −10 °C or lower.

- Digital Microprocessor Control Console — Backlit color display of RPM, RCF, time, temperature, and program.

- Motorized Safety Lid Lock and Imbalance Sensor — Electromagnetic interlock with electronic imbalance detection.

- Main Power Distribution Cable — Heavy-duty cable matched to unit's electrical rating and local mains configuration.

## 22.4 Installation Requirements

Dedicated single- or three-phase supply as per unit rating, on essential/UPS circuit for control (motor typically on raw mains). Level, vibration-tolerant reinforced floor. Adequate ventilation clearance around unit. Ambient 15—30 °C.

## 22.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate (speed, timer, temperature), Electrical Safety Test Report, CE certificate, ISO 13485 certificate.

## 22.6 Training

Application training on component preparation programs; biomedical engineering PPM and troubleshooting training.

## 22.7 Warranty

Minimum two (2) years comprehensive warranty; compressor minimum 5 years recommended.

# 23. Refrigerator Thermometer

Quantity: 2

## 23.1 Clinical & Functional Specifications

Digital dual-channel thermometer intended for continuous independent monitoring and verification of temperature inside blood bank refrigerators and freezers, with high/low alarm thresholds and traceable calibration.

## 23.2 Engineering Technical Specifications

- Configuration: Digital dual-channel thermometer with external probes.

- Channels: Minimum 2 (internal chamber + ambient).

- Measurement range: −50 °C to +70 °C (minimum).

- Resolution: 0.1 °C.

- Accuracy: Within ±0.5 °C over 0—40 °C, within ±1 °C over full range.

- Probes: Glycol-encapsulated (buffered) sensor for chamber, with thin extended leads suitable for closed-door installation.

- Display: Backlit LCD showing both channels with min/max memory.

- Alarms: Programmable high/low audible and visual alarms.

- Mounting: Magnetic wall clips and integrated kickstand.

- Power: Long-life coin-cell or standard AAA batteries, minimum 12 months operation.

- Calibration: Certificate traceable to national/international standards.

- Applicable standards: IEC 61010-1, EN 12830 principles, CE marked.

## 23.3 Standard Accessories

- Dual-Channel Digital Display Monitor — LCD showing internal and ambient temperatures with min/max memory and alarm indicators.

- Glycol-Encapsulated Sensor Probes — Buffered probes for stable chamber-temperature measurement, with thin flat leads suitable for door pass-through.

- High/Low Threshold Alert Alarms — User-programmable audible and visual alarms.

- Magnetic Wall-Mount Clips and Kickstand — Rear magnetic clips and integrated kickstand for flexible mounting.

- Battery Pack — Long-life coin-cell or AAA batteries providing minimum 12 months operation.

## 23.4 Installation Requirements

Mounting on/near monitored refrigerator; batteries supplied.

## 23.5 Documentation

User Manual, Calibration Certificate traceable to national standards, CE certificate.

## 23.6 Training

Basic user familiarization with alarms and calibration.

## 23.7 Warranty

Minimum two (2) years comprehensive warranty.

# 24. Small CBC (3-Part Differential) Analyser

Quantity: 1

## 24.1 Clinical & Functional Specifications

Compact automated hematology analyser providing complete blood count (CBC) with three-part white blood cell differential, intended for donor screening (hemoglobin, hematocrit) and basic hematology support in the blood bank.

## 24.2 Engineering Technical Specifications

- Configuration: Bench-top automated 3-part differential hematology analyser.

- Parameters: Minimum CBC with 3-part differential — WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT, LYM%, MID%, GRA%, LYM#, MID#, GRA#, RDW, MPV — with histograms.

- Throughput: Minimum 30 samples/hour.

- Sample volume: Minimum ≤20 µL whole blood (venous or capillary).

- Sample modes: Whole blood, pre-dilute; open-tube and closed-tube (where equipped).

- Measurement principles: Impedance for cell counting; cyanide-free or SLS hemoglobin methodology.

- Linearity (WBC): 0—100 × 10⁹/L or better.

- Linearity (RBC): 0—7 × 10¹²/L or better.

- Linearity (PLT): 0—1000 × 10⁹/L or better.

- Linearity (HGB): 0—25 g/dL or better.

- Precision (CV): WBC ≤3%, RBC ≤2%, HGB ≤2%, PLT ≤5% at normal levels.

- User interface: Color LCD touchscreen, minimum 7-inch diagonal.

- Data storage: Minimum 10,000 patient results with histograms.

- Printer: Integrated thermal graphics printer with paper.

- Connectivity: USB, Ethernet, LIS via HL7 or ASTM.

- Reagents: Cyanide-free where applicable; supplied as system-open or system-matched.

- Power supply: 100—240 VAC, 50/60 Hz.

- Applicable standards: IEC 61010-1, IEC 61010-2-101, IEC 61326, CE-IVD or FDA cleared, ISO 13485.

## 24.3 Standard Accessories

- Internal Sample Aspiration Probe with Auto-Wash — Automated aspiration probe with wash block cleaning cycle.

- High-Resolution Color LCD Touchscreen — Minimum 7-inch color touchscreen user interface.

- Integrated Thermal Graphics Printer — Built-in printer with two rolls of paper supplied.

- External Reagent Tubing Kit — Reagent lines, bottle caps, and level-sensor wires for diluent/lyse/rinse reagents.

- Electronic Validation/Calibration Check Strip — Verification tool for routine electronic system check.

- Medical-Grade AC/DC Power Supply — Universal 100—240 VAC adapter and detachable cord.

## 24.4 Installation Requirements

Bench-top; 220—240 VAC on UPS-backed circuit; ambient 15—30 °C, 30—80% RH; adequate space for reagents and waste containers.

## 24.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Calibration Certificate, Electrical Safety Test Report, CE-IVD/FDA certificate, ISO 13485 certificate.

## 24.6 Training

Application training for technologists; biomedical engineering PPM and troubleshooting training.

## 24.7 Warranty

Minimum two (2) years comprehensive warranty including labour, spare parts, PPM, breakdown maintenance, and software updates.

# 25. Waste Disposal Equipment Bundle

Quantity: 1

## 25.1 Clinical & Functional Specifications

Integrated set of waste management equipment intended for the safe collection, segregation, and temporary storage of infectious and sharps waste generated in the blood bank, in accordance with international biohazard waste management guidelines.

## 25.2 Engineering Technical Specifications

- Configuration: Bundle of foot-pedal-operated bins, sharps container brackets, biohazard bag holder trolley, and personal protective equipment.

- Infectious waste bins: Minimum 3 bins, foot-pedal operated, color-coded (typical yellow and red per local regulations), capacity minimum 20 L each, HDPE construction, autoclavable lids.

- Sharps container brackets: Heavy-duty wall-mounting brackets for standard sharps containers.

- Bag holder trolley: Stainless steel ring-type biohazard bag holder on castors.

- PPE: Heavy-duty industrial rubber gloves (multiple pairs) and safety eye goggles (multiple pairs).

- Compliance: WHO / local biohazard waste management guidelines.

- Applicable standards: CE marked where applicable.

## 25.3 Standard Accessories

- Foot-Pedal Infectious Waste Bins (Set of 3) — HDPE bins with pedal-operated lids, color-coded yellow/red per local biohazard code, minimum 20 L each.

- Wall-Mounting Sharps Container Brackets — Heavy-duty brackets accepting standard sharps containers.

- Stainless Steel Biohazard Bag Trolley — Ring-type bag holder on locking casters for mobile collection.

- Industrial Rubber Gloves and Safety Goggles — Heavy-duty chemical-resistant gloves and impact-resistant goggles in a pack.

## 25.4 Installation Requirements

Suitable locations for bins, brackets, and trolley; no electrical connection.

## 25.5 Documentation

User Manual/Guide, Compliance Statement.

## 25.6 Training

Staff training on biohazard waste segregation and PPE use.

## 25.7 Warranty

Minimum one (1) to two (2) years warranty on bins, brackets, and trolley.

# 26. Phlebotomy Tray

Quantity: 1

## 26.1 Clinical & Functional Specifications

Portable tray intended for organized carriage of phlebotomy consumables and equipment to donor stations, including tubes, tourniquets, antiseptic wipes, adhesive bandages, and sharps container.

## 26.2 Engineering Technical Specifications

- Configuration: Molded rigid plastic carrying basket with ergonomic handle.

- Material: High-impact ABS or equivalent, chemical- and disinfectant-resistant.

- Dimensions: Approximately 400 × 300 × 100 mm typical.

- Internal organization: Removable tube rack inserts with multi-hole sizes, integrated sharps container compartment, side slots for wipes/tourniquets/bandages.

- Cleaning: Fully wipe-cleanable with hospital disinfectants.

- Applicable standards: CE marked where applicable.

## 26.3 Standard Accessories

- Rigid Plastic Carrying Basket — High-impact molded plastic tray with ergonomic handle.

- Removable Multi-Hole Tube Rack Inserts — Interchangeable inserts sized for standard blood collection tubes.

- Integrated Sharps Container Holder — Dedicated compartment for a small sharps container.

- Side Slots for Antiseptic Wipes, Tourniquets, and Bandages — Molded compartments for phlebotomy consumables.

## 26.4 Installation Requirements

None.

## 26.5 Documentation

User Guide.

## 26.6 Training

None required.

## 26.7 Warranty

Minimum one (1) year warranty against manufacturing defects.

# 27. Apheresis Machine

Quantity: 1

## 27.1 Clinical & Functional Specifications

Fully automated apheresis system intended for the safe collection of specific blood components (platelets, plasma, red cells, or combinations) from healthy donors, and/or therapeutic apheresis procedures where clinically indicated. The system shall use closed, single-use disposable kits, provide continuous or intermittent flow separation, and shall monitor donor safety parameters throughout the procedure.

## 27.2 Engineering Technical Specifications

- Configuration: Floor-standing/trolley-mounted automated apheresis system with integrated centrifuge, pump system, optical sensors, touchscreen interface, IV pole, and backup battery.

- Procedures: Configurable for at minimum platelet apheresis (plateletpheresis); additional plasma and/or multi-component protocols as specified.

- Separation principle: Continuous- or intermittent-flow centrifugation with optical interface detection.

- Pump system: Multi-pump peristaltic drive with independent programmable flow control for donor, anticoagulant, and collection lines.

- Flow rate — donor draw: Adjustable within approximately 30—140 mL/min.

- Flow rate — anticoagulant: Adjustable per protocol.

- Pressure monitoring: Continuous monitoring of draw (venous) and return pressures with programmable alarm thresholds.

- Optical sensors: Line sensors for red-cell/plasma interface detection and air-in-line detection.

- Air detection: Ultrasonic air-in-line detectors on all donor lines with automatic clamp actuation.

- Occlusion detection: Automated line occlusion clamps with pressure-based detection.

- Disposable set: Fully closed sterile single-use disposable kit; kit auto-identification recommended.

- User interface: Color touchscreen, minimum 10-inch diagonal, high resolution.

- Donor monitoring: Continuous display of collected volume, ACD used, elapsed time, procedure progress, and alarms.

- Data storage: Onboard storage of procedure records; export via USB and Ethernet.

- Connectivity: USB, Ethernet, and bi-directional connectivity to donor management/LIS system.

- Battery backup: Internal rechargeable battery supporting safe procedure completion or safe termination on mains failure, minimum 15 minutes.

- Trolley: Heavy-duty base frame with anti-static locking casters and integrated IV pole with pressure gauge.

- Power supply: 100—240 VAC, 50/60 Hz.

- Cybersecurity: Compliant with IEC 80001 principles; user authentication; audit trail.

- Applicable standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-1-6 (usability), IEC 60601-2-24 principles (infusion pump safety) where applicable, IEC 62304 (software), IEC 62366, ISO 13485, ISO 14971, CE MDR or FDA cleared.

## 27.3 Standard Accessories

- Multi-Pump Automated Blood Separation Centrifuge Chamber — Integrated centrifuge drive with matched disposable separation chamber interface.

- Optical Line Sensors — Sensors tracking red-cell/plasma interface and detecting air bubbles on all critical lines.

- High-Resolution Color Touchscreen Interface — Adjustable-angle color touchscreen, minimum 10-inch diagonal, for procedure setup and monitoring.

- Dual-Height IV Pole with Donor Pressure Gauge — Adjustable IV pole with integrated donor-line pressure gauge/display.

- Pressure Occlusion Clamps and Line Valves — Automated barometric clamps and line occlusion valves for safety interlocks.

- Internal Emergency Backup Battery — High-capacity rechargeable battery module for safe procedure continuation/termination on mains failure.

- Heavy-Duty Trolley Frame with Anti-Static Casters — Mobile base frame with locking anti-static caster wheels.

## 27.4 Installation Requirements

Dedicated 220—240 VAC outlet on essential/UPS-backed circuit; ambient temperature 18—28 °C, 30—80% RH non-condensing; adequate space beside donor bed/chair with clearance for movement; network point for donor management system if required.

## 27.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Installation Qualification (IQ), Operational Qualification (OQ), Calibration Certificate, Electrical Safety Test Report, CE MDR/FDA certificate, ISO 13485 certificate, software validation summary.

## 27.6 Training

Comprehensive clinical/application training for apheresis operators and blood bank medical officers; biomedical engineering training on preventive maintenance and troubleshooting; refresher training as required.

## 27.7 Warranty

Minimum two (2) years comprehensive warranty including labour, spare parts, preventive maintenance visits, breakdown maintenance, software updates, and technical support.