DERMATOLOGY DEPARTMENT — TECHNICAL SPECIFICATIONS

Technical Tender Document

Table of Contents

1. Handheld Dermatoscope	7

1.1 Clinical & Functional Specifications	7

1.2 Engineering Technical Specifications	7

1.3 Standard Accessories	8

1.4 Installation Requirements	8

1.5 Documentation	8

1.6 Training	8

1.7 Warranty	8

2. Wood's Lamp	9

2.1 Clinical & Functional Specifications	9

2.2 Engineering Technical Specifications	9

2.3 Standard Accessories	9

2.4 Installation Requirements	10

2.5 Documentation	10

2.6 Training	10

2.7 Warranty	10

3. LED Examination / Procedure Light	11

3.1 Clinical & Functional Specifications	11

3.2 Engineering Technical Specifications	11

3.3 Standard Accessories	12

3.4 Installation Requirements	12

3.5 Documentation	12

3.6 Training	12

3.7 Warranty	12

4. Magnifying LED Examination Lamp	13

4.1 Clinical & Functional Specifications	13

4.2 Engineering Technical Specifications	13

4.3 Standard Accessories	13

4.4 Installation Requirements	14

4.5 Documentation	14

4.6 Training	14

4.7 Warranty	14

5. Digital Clinical Photography Camera	15

5.1 Clinical & Functional Specifications	15

5.2 Engineering Technical Specifications	15

5.3 Standard Accessories	15

5.4 Installation Requirements	16

5.5 Documentation	16

5.6 Training	16

5.7 Warranty	16

6. Electrocautery / Radiofrequency Cautery Unit	17

6.1 Clinical & Functional Specifications	17

6.2 Engineering Technical Specifications	17

6.3 Standard Accessories	17

6.4 Installation Requirements	18

6.5 Documentation	18

6.6 Training	18

6.7 Warranty	18

7. Cryotherapy Unit (Liquid Nitrogen Cryogun)	19

7.1 Clinical & Functional Specifications	19

7.2 Engineering Technical Specifications	19

7.3 Standard Accessories	19

7.4 Installation Requirements	20

7.5 Documentation	20

7.6 Training	20

7.7 Warranty	20

8. Skin Biopsy Punch Set	21

8.1 Clinical & Functional Specifications	21

8.2 Engineering Technical Specifications	21

8.3 Standard Accessories	21

8.4 Installation Requirements	21

8.5 Documentation	21

8.6 Training	21

8.7 Warranty	21

9. Minor Surgical Instrument Set (Dermatology)	23

9.1 Clinical & Functional Specifications	23

9.2 Engineering Technical Specifications	23

9.3 Standard Accessories	23

9.4 Installation Requirements	23

9.5 Documentation	23

9.6 Training	23

9.7 Warranty	24

10. Curette Set	25

10.1 Clinical & Functional Specifications	25

10.2 Engineering Technical Specifications	25

10.3 Standard Accessories	25

10.4 Installation Requirements	25

10.5 Documentation	25

10.6 Training	25

10.7 Warranty	25

11. Comedone Extractor Set	26

11.1 Clinical & Functional Specifications	26

11.2 Engineering Technical Specifications	26

11.3 Standard Accessories	26

11.4 Installation Requirements	26

11.5 Documentation	26

11.6 Training	26

11.7 Warranty	26

12. Tabletop Autoclave (20—25 L)	27

12.1 Clinical & Functional Specifications	27

12.2 Engineering Technical Specifications	27

12.3 Standard Accessories	27

12.4 Installation Requirements	28

12.5 Documentation	28

12.6 Training	28

12.7 Warranty	28

13. Stainless Steel Instrument Tray	29

13.1 Clinical & Functional Specifications	29

13.2 Engineering Technical Specifications	29

13.3 Standard Accessories	29

13.4 Installation Requirements	29

13.5 Documentation	29

13.6 Training	29

13.7 Warranty	29

14. Kidney Tray / Bowl	30

14.1 Clinical & Functional Specifications	30

14.2 Engineering Technical Specifications	30

14.3 Standard Accessories	30

14.4 Installation Requirements	30

14.5 Documentation	30

14.6 Training	30

14.7 Warranty	30

15. Full Body Narrowband UVB Phototherapy Cabinet	31

15.1 Clinical & Functional Specifications	31

15.2 Engineering Technical Specifications	31

15.3 Standard Accessories	32

15.4 Installation Requirements	32

15.5 Documentation	32

15.6 Training	32

15.7 Warranty	32

16. Hand & Foot Narrowband UVB Phototherapy Unit	33

16.1 Clinical & Functional Specifications	33

16.2 Engineering Technical Specifications	33

16.3 Standard Accessories	33

16.4 Installation Requirements	34

16.5 Documentation	34

16.6 Training	34

16.7 Warranty	34

17. Phototherapy Protective Goggles	35

17.1 Clinical & Functional Specifications	35

17.2 Engineering Technical Specifications	35

17.3 Standard Accessories	35

17.4 Installation Requirements	35

17.5 Documentation	35

17.6 Training	35

17.7 Warranty	36

# 1. Handheld Dermatoscope

Quantity: 2

## 1.1 Clinical & Functional Specifications

Handheld portable dermatoscopic examination device intended for non-invasive in-vivo examination of pigmented and non-pigmented skin lesions, moles, nevi, melanoma screening, and general dermatological surface diagnostics. The device shall enable magnified visualization of subsurface skin structures, pigment networks, vascular patterns, and epidermal-dermal junction features not visible to the naked eye. It shall support both contact (immersion/gel-based) and non-contact examination techniques and shall provide user-selectable polarized and non-polarized illumination modes to differentiate superficial from deep structures. The unit shall be lightweight, ergonomic, battery-operated, and suitable for outpatient dermatology clinics, teledermatology, screening camps, and bedside examination.

## 1.2 Engineering Technical Specifications

- Optical Magnification: Minimum 10× achromatic magnification or equivalent; higher configurable magnification acceptable

- Illumination Source: High-intensity white LED array, minimum 6 LEDs or equivalent, with color temperature approximately 5000—6500 K and Colour Rendering Index (CRI) ≥ 90

- LED Service Life: Minimum 10,000 operating hours

- Illumination Modes: User-selectable/toggle-switchable between polarized (cross-polarized) and non-polarized modes, and contact / non-contact examination modes

- Field of View: Minimum 25 mm diameter or better

- Focus: Fine adjustable focus ring or fixed-focus with optical clarity across the full field of view

- Optical System: Achromatic multi-lens system with anti-reflective coating; distortion-free edge-to-edge imaging

- Power Source: Rechargeable lithium-ion battery integrated in handle; Ni-MH acceptable as equivalent

- Battery Runtime: Minimum 60 minutes of continuous illumination on full charge

- Charging Time: Maximum 3—4 hours for full charge

- Battery Indicator: Low-battery visual indicator

- Power Adapter Input: 100—240 VAC, 50/60 Hz, universal

- Housing: Impact-resistant medical-grade polymer / anodized aluminum construction

- Weight (head + handle): Approximately 150—350 g

- Dimensions: Compact handheld ergonomic form factor

- Digital Interface (where applicable): Adapter compatibility to interface with smartphone/DSLR/mirrorless cameras for image capture (optional/configurable)

- Environmental Operating Conditions: 10—40 °C, 15—85% RH non-condensing

- Applicable Standards: IEC 60601-1 (electrical safety), IEC 60601-1-2 (EMC), ISO 13485, CE MDR / FDA clearance, RoHS

## 1.3 Standard Accessories

- Contact glass plate with 10 mm metric scale crosshair grid — Optical-grade scratch-resistant glass, calibrated 10 mm reticle for lesion measurement, autoclavable or chemically disinfectable, snap-on/threaded compatibility with dermatoscope head.

- Non-contact glass plate / faceplate adapter assembly — Optical-grade non-contact faceplate with anti-reflective coating, interchangeable with contact plate, suitable for immersion-fluid-free examination.

- Polarized / non-polarized toggle illumination control module — Integrated in the head assembly with tactile selector switch, permitting instantaneous switching between cross-polarized and non-polarized modes without loss of focus.

- Rechargeable lithium-ion battery handle with knurled grip surface — Sealed lithium-ion cell, minimum 60 minutes runtime, textured non-slip ergonomic grip, integrated charge-status indicator.

- Desktop battery charging dock cradle with mains adapter cord — Freestanding charging cradle, universal input 100—240 VAC, LED charge indicator, overcharge/short-circuit protection.

- Rigid form-fitted protective storage case — Impact-resistant polymer carry case with custom-cut foam inserts for head, handle, plates, and charger.

## 1.4 Installation Requirements

Bench/tabletop use only. Standard 220—240 VAC / 50 Hz mains socket for the charging dock. No special HVAC, gas, water, or shielding requirements. Minimum bench footprint approximately 300 × 300 mm for cradle placement.

## 1.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, electrical safety test report, CE/FDA certificates, warranty certificate, and declaration of conformity.

## 1.6 Training

On-site clinical application training for dermatologists and paramedical staff, biomedical training for cleaning, disinfection, battery handling, and basic troubleshooting.

## 1.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, battery, LEDs, charger, software updates (if applicable), preventive maintenance, and technical support.

# 2. Wood's Lamp

Quantity: 1

## 2.1 Clinical & Functional Specifications

Long-wave ultraviolet examination lamp intended for diagnostic evaluation of fungal skin/scalp infections, pigmentary disorders (vitiligo, melasma), porphyrias, erythrasma, tinea, Pseudomonas infections, and general fluorescence-based dermatological diagnostics. The device shall emit filtered long-wave UV-A illumination (approximately 320—400 nm, peak near 365 nm) that induces characteristic fluorescence in specific skin conditions when used in a darkened environment. The device shall incorporate a magnifying viewing lens for enhanced visualization and shall be suitable for tabletop or wall-mounted use in outpatient dermatology clinics.

## 2.2 Engineering Technical Specifications

- UV Emission Wavelength: Long-wave UV-A, peak approximately 365 nm (within 320—400 nm range)

- UV Source: Filtered fluorescent black-light tubes or equivalent LED-based UV-A emitter; minimum 2 tubes (typically 4 pre-installed)

- Tube Rated Life: Minimum 5,000 operating hours

- Filter: Wood's glass (nickel-oxide-doped) or equivalent optical UV-pass filter blocking visible light

- Magnification: Integrated viewing lens, minimum 3× (or 3—4 dioptre equivalent)

- Lens Diameter: Minimum 100 mm

- Illumination Field: Uniform UV output across the viewing aperture

- Warm-up Time: Maximum 1 minute to full output

- Housing: Rigid impact-resistant polymer / powder-coated metal, UV-stable

- Mounting: Convertible tabletop / wall-mount / desktop cradle

- Power Supply: 100—240 VAC, 50/60 Hz

- Power Consumption: Approximately 15—40 W depending on configuration

- On/Off Switch: Integrated on housing with pilot indicator lamp

- Cable: Minimum 2 m mains cord with strain relief and surge suppression

- Weight: Approximately 1—3 kg

- Environmental Operating Conditions: 10—40 °C, 15—85% RH non-condensing

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 62471 (photobiological safety of lamps), CE / FDA, RoHS

## 2.3 Standard Accessories

- Long-wave UV-A fluorescent tubes (set of 4 pre-installed) — Filtered Wood's glass tubes, peak emission \~365 nm, minimum 5,000-hour service life, user-replaceable.

- Integrated magnifying viewing lens (3×/4× dioptre) — Optical-grade distortion-free lens, minimum 100 mm diameter, scratch-resistant coating, mounted centrally within illumination field.

- Removable opaque blackout drape hood — Lightweight opaque fabric hood for localized darkening during examination, machine-washable, elasticated for lamp-head fit.

- Wall-mount bracket / desktop cradle — Powder-coated metal bracket with mounting hardware; freestanding cradle for tabletop use.

- Mains power cord with surge protection — Minimum 2 m, medical-grade, strain relief, integrated surge filter.

## 2.4 Installation Requirements

Standard 220—240 VAC / 50 Hz outlet. Darkened examination room recommended. Wall mounting requires standard structural drywall/concrete anchorage. No HVAC, gas, or water requirements.

## 2.5 Documentation

User manual, service manual, parts catalogue, electrical safety test report, factory test certificate, CE/FDA declarations, warranty certificate.

## 2.6 Training

Clinical application training, biomedical training on tube replacement, cleaning, and electrical safety checks.

## 2.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts (excluding consumable tubes after first year), preventive maintenance, and technical support.

# 3. LED Examination / Procedure Light

Quantity: 2

## 3.1 Clinical & Functional Specifications

Mobile high-intensity shadowless LED examination and minor-procedure light intended for general clinical examination, dermatological procedures, minor surgical interventions, biopsies, and cryotherapy. The unit shall provide cool, shadow-free, color-accurate illumination of the treatment field with minimal heat generation at the patient site. It shall be mounted on a mobile pedestal with articulated arm allowing full multi-axis positioning, and shall provide adjustable light intensity for varying clinical requirements.

## 3.2 Engineering Technical Specifications

- Light Source: High-efficiency medical-grade white LED array

- Central Illuminance (at 1 m): Minimum 40,000 lux; adjustable up to 60,000 lux or higher

- Colour Temperature: Approximately 4000—5000 K (daylight neutral)

- Colour Rendering Index (CRI / Ra): Minimum 90; R9 ≥ 80

- Light Field Diameter (at 1 m): Adjustable/focusable, approximately 150—250 mm

- Working Distance: 700—1400 mm

- LED Service Life: Minimum 30,000—50,000 hours

- Dimming Range: Electronically adjustable, minimum 20%—100% output, in defined steps or continuously variable

- Heat at Operator/Patient Head: Minimal; radiant temperature rise not to exceed clinically safe limits per IEC 60601-2-41 where applicable

- Arm: Spring-balanced articulated multi-joint arm, minimum 3 pivot points, full 360° head rotation

- Mobile Base: Five-caster antistatic base, minimum two casters with locking brakes

- Base Diameter: Approximately 550—700 mm for stability

- Sterile Handle: Detachable, autoclavable at 134 °C

- Backup Battery: Integrated rechargeable Li-ion; minimum 60 minutes runtime at full output

- Battery Charging Time: Maximum 6—8 hours

- Control: Integrated electronic touch or membrane dimmer on lamp head

- Power Supply: 100—240 VAC, 50/60 Hz

- Power Consumption: Approximately 30—80 W

- Housing / Head: Aluminum / medical-grade polymer, smooth-cleanable, IPX1 or better

- Weight (total): Approximately 15—30 kg

- Environmental Operating Conditions: 10—40 °C, 15—85% RH

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-41 (surgical/exam luminaires), ISO 13485, CE MDR / FDA

## 3.3 Standard Accessories

- Five-caster mobile pedestal stand with antistatic locking brakes — Heavy-duty powder-coated steel base, minimum 550 mm diameter, two lockable antistatic casters.

- Spring-balanced articulated extension arm with multi-axis pivot — Counterbalanced multi-joint arm, drift-free positioning, minimum 3 pivots, 360° head rotation.

- Autoclavable detachable sterile handles (2 pieces) — Steam-sterilizable at 134 °C, ergonomic grip, quick-release lock.

- Integrated electronic dimmer control on lamp head — Membrane/touch keypad, intensity indication, memory of last setting.

- Built-in rechargeable backup battery pack — Li-ion, minimum 60 min runtime, auto-switchover on mains failure, deep-discharge protection.

## 3.4 Installation Requirements

Standard 220—240 VAC / 50 Hz socket. Recommended UPS-backed circuit. No fixed mounting; requires minimum floor clearance approximately 1 m² and door clearance for maneuverability. No HVAC, gas, or water requirements.

## 3.5 Documentation

User manual, service manual, parts catalogue, electrical safety test report, factory test certificate, photometric report (illuminance/CRI/CT), CE/FDA declarations, PM schedule, warranty certificate.

## 3.6 Training

Clinical training on positioning and use; biomedical training on battery care, LED module replacement, PM, and basic troubleshooting.

## 3.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, LED module, battery, preventive maintenance, and technical support.

# 4. Magnifying LED Examination Lamp

Quantity: 2

## 4.1 Clinical & Functional Specifications

Magnifying illuminated examination lamp intended for close-up dermatological assessment, minor procedures, comedone extraction, skin surface inspection, and cosmetic dermatology. The device shall combine a shadowless circular LED ring illumination with a high-clarity magnifying lens on an articulated arm, allowing hands-free magnified visualization of the treatment area.

## 4.2 Engineering Technical Specifications

- Lens Diameter: Minimum 125 mm (5-inch)

- Lens Material: Optical-grade glass (preferred) or high-clarity crown glass; distortion-free

- Magnification / Dioptre: 5 dioptre (equivalent to \~2.25× magnification); higher configurable optional

- Illumination Source: Circular LED ring

- LED Quantity: Minimum 60 LEDs or equivalent shadowless configuration

- Colour Temperature: Approximately 5000—6500 K

- CRI: Minimum 80

- Illuminance at Working Distance: Minimum 8,000 lux at 200 mm

- LED Service Life: Minimum 30,000 hours

- Dimming: Configurable/adjustable intensity (stepped or continuous)

- Arm: Spring-loaded pantograph, minimum 800 mm reach, with anti-pinch finger guards

- Lens Cover: Flip-up protective cover to shield lens when not in use

- Mounting Options: Table G-clamp and mobile rolling floor stand with five-caster base

- Power Supply: 100—240 VAC, 50/60 Hz

- Power Consumption: Approximately 10—25 W

- Weight: Approximately 5—12 kg (with floor stand)

- Environmental Operating Conditions: 10—40 °C, 15—85% RH

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, CE MDR / FDA, RoHS

## 4.3 Standard Accessories

- High-transparency optical glass lens (5-inch, 5-dioptre) — Distortion-free, scratch-resistant, evenly ground and polished.

- Shadowless circular LED ring module with protective cover — Uniform ring illumination, flip-up dust/scratch cover.

- Spring-loaded pantograph arm with anti-pinch guards — Drift-free counterbalanced arm, minimum 800 mm reach.

- Table G-clamp and mobile caster floor stand — Table clamp fits work-surface up to 60 mm thick; mobile stand with five casters, two lockable.

- Mains power cord with cable management clips — Minimum 2 m medical-grade cord with integrated routing clips.

## 4.4 Installation Requirements

Standard 220—240 VAC / 50 Hz outlet. No structural, HVAC, gas, or water requirements. Floor stand version requires minimum 500 × 500 mm base area.

## 4.5 Documentation

User manual, service manual, parts catalogue, electrical safety test report, factory test certificate, CE/FDA declarations, warranty certificate.

## 4.6 Training

Clinical use training and biomedical training on cleaning, lens care, LED replacement, and PM.

## 4.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, LED module, and technical support.

# 5. Digital Clinical Photography Camera

Quantity: 1

## 5.1 Clinical & Functional Specifications

High-resolution digital clinical photography system intended for standardized documentation of dermatological lesions, cosmetic procedures, before/after comparisons, teledermatology consultations, and clinical research. The system shall enable high-fidelity color-accurate macro imaging of skin lesions with reproducible lighting via ring or twin-flash illumination and shall support high-capacity storage, tethered transfer, and integration with hospital PACS/EMR systems where applicable.

## 5.2 Engineering Technical Specifications

- Camera Type: Digital mirrorless or DSLR interchangeable-lens system

- Sensor Type: CMOS, APS-C or full-frame equivalent

- Effective Resolution: Minimum 24 megapixels

- Colour Depth: Minimum 14-bit RAW output

- ISO Range: Minimum ISO 100—25,600, expandable

- Shutter Speed Range: Minimum 1/4000 s to 30 s, plus Bulb

- Autofocus System: Multi-point phase/contrast detection AF, minimum 100 focus points

- Macro Lens: Fixed focal length dedicated macro lens, focal length approximately 90—105 mm, minimum 1:1 reproduction ratio

- Flash Sync Speed: Minimum 1/200 s

- Viewfinder: Optical or high-resolution EVF (minimum 2.3 MP)

- LCD Display: Minimum 3-inch, minimum 1,000,000 dots, articulating or tilting

- Storage: Dual or single SD/SDHC/SDXC slot; UHS-I or UHS-II compatible

- Video Recording (optional): Minimum Full HD 1080p

- Connectivity: USB (Type-C preferred), HDMI, Wi-Fi (configurable)

- Battery: Rechargeable lithium-ion; minimum 400 exposures per charge

- Power Supply for Charger: 100—240 VAC, 50/60 Hz

- Body Construction: Magnesium alloy or high-grade polymer with dust/moisture resistance

- Weight (body + macro lens): Approximately 900 g—1.5 kg

- Environmental Operating Conditions: 0—40 °C, 15—85% RH

- Applicable Standards: IEC 60950-1 / IEC 62368-1, CE, FCC, RoHS; DICOM/HL7 export compatibility via workstation software where applicable

## 5.3 Standard Accessories

- Digital mirrorless/DSLR body with macro lens — Interchangeable-lens body, minimum 24 MP, dedicated 1:1 macro lens (\~90—105 mm).

- Ring-flash or twin-flash lighting attachment — Dedicated macro flash with adjustable output, TTL support, guide number appropriate for close-up work, shadowless illumination for dermatology imaging.

- High-speed SD memory card (minimum 64 GB) — UHS-I Class 10 or better, minimum 64 GB, minimum 90 MB/s read.

- Rechargeable lithium-ion battery packs (2 units) — Manufacturer-matched Li-ion batteries, minimum 400 exposures per charge.

- External desktop battery charger with mains cord — Universal 100—240 VAC input, LED charge status, dual-bay preferred.

- Rigid padded transport case with custom foam — Weather-resistant impact case, custom-cut foam slots for body, lens, flash, batteries, and cards.

## 5.4 Installation Requirements

Standard mains outlet for charger. Optional workstation with dedicated clinical photography software / EMR integration. No HVAC, gas, water, or shielding requirements. Recommended standardized photography corner with neutral background and controlled lighting.

## 5.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, CE/FCC declarations, warranty certificate.

## 5.6 Training

Clinical training on standardized dermatological photography protocols, use of flash and macro lens; biomedical training on cleaning, sensor care, battery management, and software.

## 5.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, batteries, and technical support.

# 6. Electrocautery / Radiofrequency Cautery Unit

Quantity: 1

## 6.1 Clinical & Functional Specifications

High-frequency radiofrequency electrosurgical unit intended for cutting, coagulation, fulguration, and desiccation of skin lesions and soft tissue during dermatological, cosmetic, and minor surgical procedures. The device shall support both monopolar and bipolar modes with independently adjustable power output and shall provide precise, controlled tissue effect with minimal thermal spread. The unit shall include patient-return-electrode monitoring, foot-switch and hand-switch activation, and audible/visual safety indicators.

## 6.2 Engineering Technical Specifications

- Operating Frequency: Radiofrequency, minimum 1.5 MHz (typical 1.5—4 MHz)

- Output Modes: Cut, Blend/Cut-Coag, Coagulation, Fulguration, Bipolar (monopolar + bipolar mandatory)

- Maximum Power Output: Minimum 50 W (typical 50—120 W configurable)

- Power Adjustment: Digitally selectable in defined steps

- Display: Digital LED/LCD showing mode, power level, and active output

- Activation: Both foot-switch and hand-switch (finger-switch pencil) supported

- Patient Return Electrode Monitoring: Continuous contact-quality monitoring with audible/visual alarm on disconnect

- Bipolar Mode: Independent power adjustment, auto-start optional

- Safety Features: Audible activation tone, isolated output, self-test on power-up, over-temperature protection

- Housing: Impact-resistant medical-grade enclosure; smooth cleanable surface

- Power Supply: 100—240 VAC, 50/60 Hz

- Power Consumption: Approximately 150—300 VA

- Weight: Approximately 3—8 kg

- Dimensions: Compact tabletop configuration

- Environmental Operating Conditions: 10—40 °C, 15—85% RH

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-2 (HF surgical equipment), ISO 13485, CE MDR / FDA

## 6.3 Standard Accessories

- High-frequency RF generator console with digital display — Compact console, digital indication of mode and wattage, membrane keypad, self-test on start-up.

- Reusable monopolar autoclavable electrode pen with hand-switch — Autoclavable at 134 °C, dual-button hand-switch (cut/coag), universal collet for electrode tips.

- Reusable cautery tip electrode assortment — Set including fine needle, ball, loop, and blade tips; autoclavable stainless steel/tungsten; universal fit.

- Reusable bipolar coagulation forceps with cable — Fine-tip bipolar forceps, autoclavable, insulated shafts, reusable connecting cable.

- Reusable flexible patient return plate with cable — Silicone/stainless flexible plate with strap, connecting cable compatible with generator's return-electrode monitoring.

- Waterproof double-pedal footswitch — IPX8 dust/water-tight sealed footswitch, dual pedal (cut/coag), heavy-duty cable.

- Mobile utility cart with drawers and casters — Powder-coated steel/medical-grade polymer cart, minimum 2 drawers, four casters (two lockable).

## 6.4 Installation Requirements

Standard 220—240 VAC / 50 Hz mains outlet with proper earthing (equipotential bonding recommended). UPS-backed circuit preferred. No HVAC, gas, water, or shielding requirements beyond normal minor-procedure room. Minimum bench footprint approximately 500 × 400 mm.

## 6.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, electrical safety test report (per IEC 62353), calibration certificate, CE/FDA declarations, PM schedule.

## 6.6 Training

Clinical training on modes and safe use; biomedical training on PM, electrical safety testing, calibration verification, and troubleshooting.

## 6.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, footswitch, cables, preventive maintenance, and technical support.

# 7. Cryotherapy Unit (Liquid Nitrogen Cryogun)

Quantity: 1

## 7.1 Clinical & Functional Specifications

Handheld liquid-nitrogen cryosurgical device intended for controlled cryogenic destruction of benign and pre-malignant skin lesions such as warts, actinic keratoses, seborrheic keratoses, molluscum contagiosum, and other dermatological indications. The unit shall enable precise trigger-controlled liquid-nitrogen spray delivery with interchangeable aperture tips for varying lesion sizes and depths, and shall provide safe cryogen containment with minimal evaporative loss.

## 7.2 Engineering Technical Specifications

- Cryogen: Liquid nitrogen (LN₂), approximately —196 °C

- Reservoir Capacity: Configurable, approximately 300—500 mL

- Reservoir Construction: Vacuum-insulated double-walled stainless steel (medical grade)

- Static Hold Time: Minimum 6 hours without refill under normal conditions

- Working Pressure: Within manufacturer-defined safe operating limits with integrated pressure-relief valve

- Spray Delivery: Trigger-actuated with fine control for intermittent or continuous spray

- Aperture Tips: Interchangeable, minimum 4 sizes (A, B, C, D — fine to coarse)

- Extension Nozzles: Straight and curved, quick-thread attachment

- Safety Features: Integrated pressure-relief valve, automatic venting, thermally insulated grip

- Filling Method: Top-fill via siphon pump attachment

- Housing: Medical-grade stainless steel with insulating handle

- Weight (empty): Approximately 0.5—1.0 kg

- Dimensions: Handheld ergonomic configuration

- Environmental Operating Conditions: 10—40 °C, 15—85% RH

- Applicable Standards: ISO 13485, CE MDR / FDA, applicable pressure-vessel and cryogenic-equipment standards

## 7.3 Standard Accessories

- Vacuum-insulated double-walled stainless steel flask (300—500 mL) — Medical-grade 304/316 stainless steel, vacuum-jacketed, minimum 6-hour hold time.

- Interchangeable cryo-spray aperture tips (A, B, C, D) — Precision-machined stainless steel tips of graduated aperture sizes, thread-compatible.

- Straight and curved extension nozzles with quick-thread rings — Stainless steel extension probes for hard-to-reach lesions.

- Fingertip trigger release lever assembly — Ergonomic trigger, tactile feedback, controlled continuous/intermittent delivery.

- Heavy-duty desktop storage stand cradle — Stable powder-coated metal cradle for upright resting of the cryogun.

- LN₂ extraction siphon pump with flexible withdrawal tube — Compatible with standard LN₂ dewar outlets, insulated withdrawal tube.

## 7.4 Installation Requirements

No electrical supply required. Requires access to a separately procured LN₂ storage dewar and well-ventilated procedure room with adequate oxygen-deficiency monitoring where applicable. No HVAC/gas/water piping required beyond general room ventilation.

## 7.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, pressure-vessel test certificate, CE/FDA declarations, warranty certificate.

## 7.6 Training

Clinical training on cryosurgical techniques; biomedical/safety training on LN₂ handling, PPE, PM, and pressure-valve inspection.

## 7.7 Warranty

Minimum comprehensive two (2) years warranty covering labour and spare parts.

# 8. Skin Biopsy Punch Set

Quantity: 2

## 8.1 Clinical & Functional Specifications

Precision reusable skin biopsy punch instrument set intended for full-thickness cylindrical skin biopsy sampling for histopathological diagnosis of dermatological conditions. The set shall provide multiple graduated punch diameters and shall enable clean, atraumatic tissue-core removal.

## 8.2 Engineering Technical Specifications

- Punch Handle: Reusable, medical-grade stainless steel (304/316), ergonomic knurled grip

- Blade Retention: Rapid-lock collet mechanism for interchangeable blades

- Blade Material: Surgical-grade stainless steel, precision-ground cutting edge

- Blade Sizes: Minimum 2 mm, 3 mm, 4 mm, and 5 mm inner diameter

- Overall Length: Approximately 100—130 mm

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 13485, ISO 7153-1 (surgical instrument materials), CE MDR / FDA

## 8.3 Standard Accessories

- Reusable stainless steel punch handles with rapid-lock collets — Medical-grade stainless steel, autoclavable, quick-lock blade retention.

- Interchangeable circular punch biopsy blades (2, 3, 4, 5 mm) — Precision-ground stainless steel cutting cylinders, autoclavable.

- Fine specimen extraction needle stylet — Stainless steel fine stylet for atraumatic core removal.

- Perforated stainless steel storage box with silicone holders — Autoclavable perforated tray with silicone finger rows for organized instrument holding.

## 8.4 Installation Requirements

No installation required. Requires access to standard steam autoclave for sterilization.

## 8.5 Documentation

Material certificates, sterilization compatibility statement, CE/FDA declarations, warranty certificate.

## 8.6 Training

Clinical familiarization for procedural technique; CSSD training for cleaning and sterilization.

## 8.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 9. Minor Surgical Instrument Set (Dermatology)

Quantity: 3

## 9.1 Clinical & Functional Specifications

Reusable minor surgical instrument set intended for dermatological excisions, biopsies, suturing, wound closure, and minor surgical procedures. All instruments shall be manufactured from medical-grade stainless steel and shall be fully autoclavable.

## 9.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel, martensitic/austenitic per ISO 7153-1

- Finish: Satin/matte non-reflective; corrosion-resistant

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 13485, ISO 7153-1, CE MDR / FDA

## 9.3 Standard Accessories

- Iris fine-tissue dissection scissors (straight and curved, sharp tips) — Approximately 110 mm length, precision-ground fine tips, autoclavable.

- Adson-Brown delicate tissue forceps — Approximately 120 mm, fine multi-teeth (7×7 configuration), autoclavable.

- Jeweler-type ultra-fine suture tying forceps (straight and angled tips) — Approximately 110 mm, ultra-fine tapered tips, non-magnetic option preferred.

- Webster needle holder (smooth jaw) — Approximately 125 mm, smooth-jaw fine microsurgical design, tungsten-carbide insert optional.

- Scalpel handle No. 3 with metric measurement scale — Standard No. 3 blade compatibility, integrated engraved metric ruler.

- Backhaus towel clamps and Halsted mosquito hemostatic forceps (curved, 2 pieces) — Standard sizes, autoclavable, ratchet-lock.

- Heavy-duty perforated stainless steel sterilization tray with locking lid — Sized to hold complete set, autoclavable perforated tray with lockable lid.

## 9.4 Installation Requirements

No installation. Requires standard steam autoclave and CSSD workflow.

## 9.5 Documentation

Material certificates, ISO 7153-1 conformity declaration, CE/FDA declarations.

## 9.6 Training

Clinical familiarization and CSSD reprocessing training.

## 9.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 10. Curette Set

Quantity: 2

## 10.1 Clinical & Functional Specifications

Reusable dermatological curette instrument set intended for scraping and removal of superficial skin lesions such as warts, seborrheic keratoses, molluscum contagiosum, and basal cell carcinomas in combination with electrosurgery.

## 10.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel per ISO 7153-1

- Handle: Ergonomic hexagonal or textured grip

- Loop Sizes: Assorted 1 mm, 2 mm, 3 mm, 4 mm

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 13485, ISO 7153-1, CE MDR / FDA

## 10.3 Standard Accessories

- Fox dermal curettes (oval-loop, assorted 1—4 mm) — Stainless steel, precision-ground oval loops, autoclavable.

- Piffard dermal curettes (pear-shaped loop, assorted sizes) — Stainless steel, pear-shaped cutting loops, autoclavable.

- Ergonomic hexagonal handles with slip-resistant grip — Textured non-slip finish for positive grip.

- Perforated stainless steel sterilization container — Autoclavable, sized to hold complete set.

## 10.4 Installation Requirements

No installation; requires CSSD sterilization capability.

## 10.5 Documentation

Material certificates, CE/FDA declarations.

## 10.6 Training

Clinical familiarization and CSSD training.

## 10.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 11. Comedone Extractor Set

Quantity: 2

## 11.1 Clinical & Functional Specifications

Reusable comedone extractor instrument set intended for the mechanical extraction of open and closed comedones, milia, and superficial acne lesions in dermatological and cosmetic dermatology practice.

## 11.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel per ISO 7153-1

- Finish: Polished, corrosion-resistant, autoclavable

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 13485, ISO 7153-1, CE MDR / FDA

## 11.3 Standard Accessories

- Unna-style comedone extractor (flat and wire loop ends) — Dual-ended stainless steel, autoclavable.

- Schamberg-style comedone extractor (crimped square loops) — Dual-ended stainless steel, autoclavable.

- Saalfeld-style comedone extractor (lance point and loop) — Dual-ended stainless steel with fine lance tip, autoclavable.

- Rigid velvet-lined instrument organizing wallet pouch — Structured wallet with individual instrument slots.

## 11.4 Installation Requirements

No installation; standard sterilization required.

## 11.5 Documentation

Material certificates, CE/FDA declarations.

## 11.6 Training

Clinical familiarization and CSSD training.

## 11.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 12. Tabletop Autoclave (20—25 L)

Quantity: 1

## 12.1 Clinical & Functional Specifications

Automatic microprocessor-controlled tabletop steam sterilizer intended for terminal sterilization of solid, hollow, packaged, and unpackaged dermatological surgical instruments, curettes, punches, and reusable accessories. The device shall operate on saturated pressurized steam principles and shall provide multiple pre-programmed sterilization cycles with automatic drying, digital monitoring, integrated safety systems, and cycle documentation output.

## 12.2 Engineering Technical Specifications

- Chamber Capacity: 20—25 L

- Chamber Material: Medical-grade stainless steel (AISI 304/316L)

- Sterilization Temperature Range: Selectable, typically 121 °C and 134 °C

- Operating Pressure: Approximately 1.1—2.1 bar (working pressure)

- Cycle Types: Minimum programmable cycles for unwrapped instruments, wrapped instruments, hollow/porous loads, and liquids (as applicable); user-configurable

- Class: Class B / Class N / Class S per EN 13060 (Class B preferred for wrapped/hollow loads)

- Drying: Automatic vacuum-assisted or heated drying phase

- Temperature Accuracy: Within ±1 °C of set point

- Pressure Accuracy: Within ±0.1 bar

- Display: Digital LCD/LED showing temperature, pressure, cycle phase, and time

- Cycle Documentation: Integrated printer or USB/RS-232 data port for cycle log export

- Water Reservoir: Integrated with low-water cutoff

- Safety Features: Door interlock, over-pressure relief valve, over-temperature cutoff, low-water cutoff, mains fuse

- Power Supply: 220—240 VAC, 50/60 Hz

- Power Consumption: Approximately 1.5—2.5 kW

- External Dimensions: Compact tabletop configuration

- Weight: Approximately 40—70 kg

- Environmental Operating Conditions: 10—40 °C, 15—85% RH

- Applicable Standards: IEC 61010-1, IEC 61010-2-040, EN 13060, IEC 60601-1-2 (EMC), ISO 13485, CE MDR / FDA, PED (Pressure Equipment Directive) where applicable

## 12.3 Standard Accessories

- Stainless steel chamber with mechanical door lock latch — Medical-grade 304/316L stainless steel, mirror-finish, safety-interlocked door.

- Perforated internal load trays (set of 3) — Stainless steel perforated trays sized to chamber.

- Tray extraction handle tool — Insulated ergonomic handle for hot-tray extraction.

- Digital temperature and pressure gauge display — Dual-parameter display with real-time cycle monitoring.

- Mechanical over-pressure exhaust safety valve — Certified pressure-relief valve per applicable pressure-vessel standards.

- Water reservoir with low-water cutoff switch — Integrated tank with automatic power cutoff on low water.

## 12.4 Installation Requirements

Dedicated 220—240 VAC / 50 Hz, minimum 15 A grounded circuit; UPS not required but recommended power stabilization. Distilled/demineralized water supply (manually filled). Well-ventilated CSSD/procedure room. No gas or fixed drainage requirement, but bench discharge tray recommended. Minimum bench area approximately 700 × 600 mm with rear clearance for ventilation.

## 12.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, EN 13060 type-test certificate, pressure-vessel certificate, electrical safety test report (IEC 62353), CE/FDA declarations, PM schedule, calibration certificate.

## 12.6 Training

Clinical/CSSD training on cycle selection, loading, and biological/chemical indicator use; biomedical training on PM, calibration, and troubleshooting.

## 12.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, gaskets, valves, preventive maintenance, and technical support.

# 13. Stainless Steel Instrument Tray

Quantity: 10

## 13.1 Clinical & Functional Specifications

Reusable medical-grade stainless steel instrument tray intended for holding, transporting, and sterilizing dermatological instruments and consumables during clinical procedures.

## 13.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel AISI 304 (18/8)

- Finish: Polished, mirror or satin, corrosion-resistant

- Construction: Seamless with radiused internal corners for cleanability

- Standard Size Range: Configurable; typical inner dimensions approximately 300 × 200 × 30 mm (or as specified)

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 7153-1, ISO 13485, CE

## 13.3 Standard Accessories

- Heavy-duty polished 304 stainless steel tray frame — Medical-grade 304, mirror/satin finish, seamless construction.

- Seamless radiused internal corners — Fully rounded internal corners eliminating crevices for effective cleaning.

- Form-fitted stainless steel drop-on cover with pull handle — Matching lid with integrated ergonomic pull handle.

## 13.4 Installation Requirements

None. Standard CSSD reprocessing compatibility.

## 13.5 Documentation

Material certificate, CE declaration.

## 13.6 Training

Not required beyond standard CSSD handling.

## 13.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 14. Kidney Tray / Bowl

Quantity: 10

## 14.1 Clinical & Functional Specifications

Reusable kidney-shaped stainless steel receiver basin intended for collection of surgical waste, fluids, swabs, and instruments during clinical procedures.

## 14.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel AISI 304 (18/8)

- Finish: Polished, corrosion-resistant, seamless

- Standard Sizes: 8-inch and/or 10-inch (approximately 200—250 mm length)

- Rim: Flanged stabilization lip for hand-held control

- Sterilization: Steam-autoclavable at 134 °C

- Applicable Standards: ISO 13485, CE

## 14.3 Standard Accessories

- Polished stainless steel kidney basin (8-inch or 10-inch) — Medical-grade 304 stainless steel, seamless kidney shape.

- Flanged perimeter stabilization lip — Reinforced flanged rim for grip control.

- Matching graduated stainless steel solution bowl — Small graduated stainless steel bowl for solutions/dressings.

## 14.4 Installation Requirements

None.

## 14.5 Documentation

Material certificate, CE declaration.

## 14.6 Training

Not required.

## 14.7 Warranty

Minimum two (2) years warranty against manufacturing defects.

# 15. Full Body Narrowband UVB Phototherapy Cabinet

Quantity: 1

## 15.1 Clinical & Functional Specifications

Full-body vertical narrowband UVB (NB-UVB) phototherapy cabinet intended for the treatment of psoriasis, vitiligo, atopic dermatitis, mycosis fungoides, and other UV-responsive dermatological conditions. The device shall emit therapeutic narrowband UVB radiation at approximately 311 nm and shall enable microprocessor-controlled dosimetry, precise timed exposure, integrated safety features, and standardized reproducible treatment delivery for standing adult patients.

## 15.2 Engineering Technical Specifications

- UV Emission Wavelength: Narrowband UVB, peak 311 ± 2 nm

- Lamp Type: Narrowband UVB fluorescent lamps (or equivalent qualified UVB source)

- Number of Lamps: Minimum 24 lamps; up to 48 or equivalent configuration

- Lamp Service Life: Minimum 500—1000 rated hours per lamp

- Irradiance at Patient Position: Minimum 8—12 mW/cm² typical, or equivalent, ensuring clinically acceptable treatment times

- Dosimetry: Integrated internal UV sensor with automatic dose compensation for lamp aging

- Treatment Programs: Microprocessor-controlled, multiple selectable protocols (psoriasis, vitiligo, custom)

- Timer Range: Selectable from seconds to minutes with 1-second resolution or better

- Cabinet Type: Vertical stand-in enclosure with patient handrails

- Interior Diameter: Sufficient for adult patient with handrails, typical approximately 800—1000 mm

- Height: Full-body coverage, approximately 2000—2200 mm

- Cooling: Forced-air ventilation with open-top exhaust

- Emergency Stop: Patient pull-cord emergency shut-off linked to master circuit breaker

- Operator Console: External remote workstation with dual digital timers, dose display, protocol selection

- Lamp Protection: Internal acrylic UV-transparent protective grilles/shields

- Power Supply: 380—415 VAC three-phase, 50/60 Hz (or 220—240 VAC single-phase per configuration)

- Power Consumption: Approximately 4—10 kVA depending on lamp count

- Weight: Approximately 200—400 kg

- Environmental Operating Conditions: 15—30 °C, 15—75% RH

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-83 (UV therapy equipment), IEC 62471 (photobiological safety), ISO 13485, CE MDR / FDA

## 15.3 Standard Accessories

- NB-UVB fluorescent lamps (24—48 lamps, 311 nm) — Certified narrowband UVB lamps, peak 311 nm, minimum 500—1000 rated hours, user-replaceable.

- Microprocessor dosimetry system with integrated UV sensor — Real-time UV output monitoring, automatic dose compensation, calibrated sensor.

- Internal handrails and open-top ventilation fan array — Patient stabilization handrails; forced-air cooling.

- Emergency patient pull-cord shut-off — Wired to master circuit breaker, immediate cabinet de-energization on activation.

- External operator console with dual digital timers — Remote workstation with dose display, protocol memory, patient database (as applicable).

- Acrylic UV-transparent protective lamp shields — Interior protective grilles/panels shielding lamps from contact.

## 15.4 Installation Requirements

Dedicated power supply per manufacturer's specification (single or three-phase). Proper equipotential earthing. Dedicated treatment room with minimum floor area approximately 3 × 3 m, ceiling height minimum 2.5 m. HVAC with adequate ventilation to dissipate heat load. UV-safe interlocked door recommended. UPS not typically required. No gas/water requirements. Structural floor loading capacity to support unit weight.

## 15.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, UV output calibration certificate, electrical safety test report (IEC 62353), photobiological safety report (IEC 62471), CE/FDA declarations, PM schedule.

## 15.6 Training

Clinical training on NB-UVB protocols, dosimetry, patient safety, and eye protection; biomedical training on lamp replacement, dosimetry calibration, PM, and electrical safety.

## 15.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, lamps (first set), dosimetry sensor, preventive maintenance, and technical support.

# 16. Hand & Foot Narrowband UVB Phototherapy Unit

Quantity: 1

## 16.1 Clinical & Functional Specifications

Compact narrowband UVB phototherapy device intended for localized treatment of hand and foot dermatoses including palmoplantar psoriasis, palmoplantar eczema, hand/foot vitiligo, and other UV-responsive localized conditions. The device shall provide simultaneous or selectable hand and foot exposure with microprocessor-controlled dosimetry and safety interlocks.

## 16.2 Engineering Technical Specifications

- UV Emission Wavelength: Narrowband UVB, peak 311 ± 2 nm

- Lamp Configuration: NB-UVB fluorescent lamps configured for simultaneous hand and/or foot exposure; minimum 4 lamps

- Lamp Service Life: Minimum 500—1000 rated hours

- Irradiance: Sufficient for clinically effective localized treatment (typical minimum 8 mW/cm² at treatment surface)

- Timer: Digital microprocessor-controlled, selectable in 1-second increments

- Auto Shut-off: Automatic cycle termination and safety interlock

- Reflectors: High-reflectivity interior aluminum reflectors behind lamp grids

- Patient Interface: Removable transparent acrylic protective plates over treatment area

- Housing: Tabletop or pedestal-mounted chamber

- Power Supply: 220—240 VAC, 50/60 Hz

- Power Consumption: Approximately 400—800 W

- Weight: Approximately 15—30 kg

- Environmental Operating Conditions: 15—30 °C, 15—75% RH

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-83, IEC 62471, ISO 13485, CE MDR / FDA

## 16.3 Standard Accessories

- Compact tabletop/pedestal dual-exposure treatment chamber — Enclosed chamber sized for simultaneous hand/foot exposure.

- NB-UVB 311 nm fluorescent lamps — Certified narrowband UVB, minimum 500—1000 hours rated life.

- Digital timer microprocessor console with auto shut-off — Membrane keypad, digital display, safety interlock.

- Interior aluminum reflector mirrors — High-reflectivity reflectors behind lamps for uniform irradiance.

- Removable transparent acrylic patient plate guards — UV-transparent acrylic protective overlay sheets, replaceable.

## 16.4 Installation Requirements

Standard 220—240 VAC / 50 Hz mains outlet with proper earthing. Well-ventilated treatment area. Bench footprint approximately 600 × 500 mm. No HVAC, gas, or water requirements beyond standard.

## 16.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, UV output calibration certificate, electrical safety test report, CE/FDA declarations, PM schedule.

## 16.6 Training

Clinical training on localized NB-UVB protocols; biomedical training on lamp replacement, dosimetry, PM, and safety testing.

## 16.7 Warranty

Minimum comprehensive two (2) years warranty covering labour, spare parts, lamps (first set), and technical support.

# 17. Phototherapy Protective Goggles

Quantity: 10

## 17.1 Clinical & Functional Specifications

UV-protective eyewear intended for patient and staff protection during UVA/UVB phototherapy sessions. The goggles shall provide total blocking of UV radiation in the therapeutic wavelength range and shall be comfortable, adjustable, and suitable for repeated clinical use.

## 17.2 Engineering Technical Specifications

- UV Attenuation: Minimum 99.9% blocking of UVA (315—400 nm) and UVB (280—315 nm)

- Optical Density: Minimum OD 5+ across 200—400 nm

- Frame Material: High-density opaque medical-grade polymer

- Fit: Full eye-cup coverage with light-tight side shielding

- Strap: Adjustable elastic silicone head band with toggle buckle

- Cleaning Compatibility: Compatible with standard hospital disinfectants (alcohol/quaternary ammonium)

- Weight: Approximately 30—80 g per pair

- Applicable Standards: EN 170 (UV filters), EN 166 (personal eye protection), ISO 13485, CE

## 17.3 Standard Accessories

- Total UV-blocking opaque polymer eye-cups (10 sets) — Rigid opaque polymer frames, full-cup design, minimum OD 5+ across UV range.

- Adjustable silicone head band straps with toggle buckle — Elasticated silicone straps with quick-adjust locking toggles.

- Neonate-safe disposable soft fabric eye-shield masks (starter pack of 5) — Soft opaque fabric wraps for neonatal/pediatric use.

- Microfiber cleaning cloths and individual storage boxes — Optical microfiber cloths and rigid polymer storage cases per pair.

## 17.4 Installation Requirements

None.

## 17.5 Documentation

UV attenuation test certificate, EN 170 / EN 166 conformity, CE declaration.

## 17.6 Training

Familiarization on correct fitting and cleaning.

## 17.7 Warranty

Minimum one (1) year warranty against manufacturing defects.