HOSPITAL CARDIOLOGY EQUIPMENT — TECHNICAL SPECIFICATIONS

Technical Tender Document

Table of Contents

1. ECG MACHINE — 6-CHANNEL WITH LCD	4

1.1 Clinical & Functional Specifications	4

1.2 Engineering Technical Specifications	4

1.3 Standard Accessories	7

1.4 Installation Requirements	8

1.5 Documentation	8

1.6 Training	8

1.7 Warranty	9

2. EXERCISE TOLERANCE TEST (ETT) / TREADMILL STRESS-TEST SYSTEM	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	12

2.4 Installation Requirements	14

2.5 Documentation	14

2.6 Training	14

2.7 Warranty	14

3. HOLTER ECG MONITOR / RECORDER (24-HOUR AMBULATORY)	16

3.1 Clinical & Functional Specifications	16

3.2 Engineering Technical Specifications	16

3.3 Standard Accessories	17

3.4 Installation Requirements	18

3.5 Documentation	18

3.6 Training	19

3.7 Warranty	19

4. HOLTER ANALYSIS WORKSTATION	20

4.1 Clinical & Functional Specifications	20

4.2 Engineering Technical Specifications	20

4.3 Standard Accessories	22

4.4 Installation Requirements	23

4.5 Documentation	23

4.6 Training	23

4.7 Warranty	23

5. HOLTER BP / AMBULATORY BP MONITOR (ABPM)	24

5.1 Clinical & Functional Specifications	24

5.2 Engineering Technical Specifications	24

5.3 Standard Accessories	26

5.4 Installation Requirements	27

5.5 Documentation	27

5.6 Training	27

5.7 Warranty	28

# 1. ECG MACHINE — 6-CHANNEL WITH LCD

Quantity: 18

## 1.1 Clinical & Functional Specifications

The equipment shall be a diagnostic-grade 12-lead resting electrocardiograph designed for acquisition, display, analysis, storage, and printing of the standard 12-lead ECG in adult and paediatric patients. It shall be suitable for use in outpatient clinics, cardiology departments, emergency rooms, ICUs, wards, and pre-operative assessment units.

The device shall simultaneously acquire all 12 leads and display any user-selectable combination of 3, 6, or 12 leads on the built-in display. It shall provide automatic, manual, rhythm, and copy modes of operation. It shall incorporate an automatic interpretation algorithm producing measurement values (heart rate, PR, QRS, QT, QTc, P/QRS/T axes) and interpretive statements for adult and paediatric patients.

Functional capabilities shall include:

- Real-time on-screen preview of all 12 leads with adjustable gain, speed, and filter settings

- Automatic detection of lead-off, poor electrode contact, and pacemaker pulses

- Automatic and manual acquisition modes with configurable pre-acquisition memory of at least 10 seconds

- Long rhythm strip recording of any user-selected lead

- Digital filters for baseline drift, muscle tremor (EMG), AC line interference (50/60 Hz), and low-pass filtering

- Storage of acquired ECGs in internal memory with patient demographics

- Export of ECG records in standard interchange formats (PDF, XML, DICOM, HL7, SCP-ECG or equivalent)

- Networking capability to hospital information system (HIS) / electronic medical record (EMR) / cardiology information system

- Password-protected user login and audit trail

- Multi-language user interface (English mandatory)

## 1.2 Engineering Technical Specifications

Configuration

- Portable/trolley-mountable 12-lead simultaneous acquisition electrocardiograph

- Minimum 6-channel simultaneous printing capability

Display

- Colour TFT LCD or equivalent display technology

- Minimum 7-inch diagonal viewing area

- Minimum resolution 800 × 480 pixels or better

- Touchscreen operation (capacitive or resistive) preferred; alternatively membrane keypad with dedicated function keys

- Simultaneous display of at least 3, 6, or 12 leads, user selectable

Signal Acquisition

- Number of leads: Standard 12-lead simultaneous acquisition

- Simultaneously acquired channels: Minimum 12

- Sampling rate: Minimum 1000 samples/second/channel or higher

- A/D conversion: Minimum 16-bit resolution

- Input impedance: Minimum 50 MΩ

- Common Mode Rejection Ratio (CMRR): Minimum 100 dB (with AC filter on)

- Frequency response (bandwidth): Minimum 0.05 Hz to 150 Hz

- Dynamic range: Minimum ±5 mV

- Sensitivity/Gain: Selectable 2.5, 5, 10, 20 mm/mV (auto-gain option)

- Paper/sweep speed: Selectable 5, 10, 25, 50 mm/s

- Time constant: Minimum 3.2 seconds

- Pacemaker pulse detection: Automatic, with pulse marker on printout

- Noise level: ≤ 15 μV peak-to-peak (referred to input)

Filters

- AC (mains) filter: 50 Hz / 60 Hz selectable

- Muscle/EMG filter: Selectable (e.g., 25 / 35 / 45 Hz or equivalent)

- Baseline drift/wander filter: Selectable

- Low-pass filter: Selectable up to at least 150 Hz

Operating Modes

- Automatic (Auto), Manual, Rhythm, Copy, and Emergency (one-touch) modes

- Configurable auto-print formats: 3×4, 3×4+1R, 3×4+3R, 6×2, 12×1, rhythm

Interpretation & Measurements

- Built-in automatic measurement of HR, PR, QRS duration, QT, QTc, P/QRS/T axes

- Automatic interpretation algorithm for adult and paediatric patients

Processor & Memory

- Embedded medical-grade processor of adequate performance

- Internal storage capacity: Minimum 200 ECG records in full-disclosure format

- Removable media support: USB / SD card

- Retrieval, review, and re-print of stored ECGs

Printer

- Integrated high-resolution thermal printer

- Minimum print resolution: 200 dots/mm vertical × 40 dots/mm horizontal (or equivalent)

- Paper width: Minimum 210 mm (A4) or equivalent to accommodate 6-channel printout

- Paper type: Thermal Z-fold or roll

- Configurable report formats with patient demographics, measurements, and interpretation

Connectivity / Interfaces

- USB ports (minimum 2)

- Ethernet (RJ-45) 10/100 Mbps or better

- Wi-Fi (IEEE 802.11 b/g/n) — optional/configurable

- Support for PDF, XML, DICOM, HL7, SCP-ECG or equivalent export formats

- Compatibility with hospital EMR / cardiology information systems

Power Supply

- Input: 100—240 VAC, 50/60 Hz, single phase

- Internal rechargeable lithium-ion battery

- Battery backup: Minimum 2 hours of continuous operation or minimum 50 ECG recordings on full charge

- Charging time: Maximum 4—6 hours

- Automatic switch-over between mains and battery

Physical

- Weight: Approximately 3—6 kg (with battery)

- Dimensions: Compact tabletop / trolley-mountable configuration suitable for hospital use

- Integrated carry handle

Environmental

- Operating temperature: 10—40 °C

- Storage temperature: —20 to +60 °C

- Relative humidity: 25—95% non-condensing

- Ingress protection: Suitable for hospital environment

Electrical Safety Classification

- Class I, Type CF defibrillation-proof applied part

- Defibrillator protection on all patient inputs

Applicable Standards

- IEC 60601-1 (General requirements for basic safety and essential performance)

- IEC 60601-1-2 (EMC)

- IEC 60601-2-25 (Particular requirements for electrocardiographs)

- IEC 60601-2-51 (Recording and analysing electrocardiographs)

- IEC 62304 (Medical device software)

- IEC 62366 (Usability engineering)

- ISO 13485 (Manufacturer QMS)

- ISO 14971 (Risk management)

- CE MDR / US FDA clearance or equivalent regulatory approval

- HL7 / DICOM / SCP-ECG compatibility

## 1.3 Standard Accessories

10-Lead Patient ECG Trunk Cable with Banana Plug Connectors — 1 set

- Medical-grade 10-lead patient cable, defibrillation-proof

- Shielded, low-noise construction with anti-tangle design

- Distal termination: Banana (4 mm) plug connectors compatible with reusable limb and chest electrodes

- Cable length: Approximately 3—4 metres

- IEC/AHA colour coding, clearly labelled

- Reusable and cleanable/disinfectable with hospital-grade agents

- Latex-free, biocompatible per ISO 10993

Reusable Adult Chest Suction Electrodes — Set of 6

- Reusable Welsh-type suction bulb chest electrodes for precordial leads V1—V6

- Metal electrode (silver/silver-chloride or nickel-plated brass) with rubber suction bulb

- Compatible with 4 mm banana socket

- Latex-free suction bulb

- Reusable, cleanable, and disinfectable

- Set to include 6 pieces, colour-coded per IEC/AHA standard

Reusable Adult Limb Extremity Clamp Electrodes — Set of 4

- Reusable limb clamp electrodes for RA, LA, RL, LL

- Corrosion-resistant metal electrode with spring-loaded clamp mechanism

- Compatible with 4 mm banana socket

- Latex-free, biocompatible

- Reusable, cleanable, and disinfectable

- Set of 4, colour-coded per IEC/AHA standard

Integrated High-Resolution Thermal Matrix Printer

- Built-in high-resolution thermal array printer

- Minimum print resolution: 200 dots/mm vertical × 40 dots/mm horizontal (or equivalent)

- Print width: Minimum 210 mm accommodating 6-channel simultaneous printout

- Supports A4 / letter and Z-fold paper

Thermal Z-Fold or Roll Chart Paper — 2 rolls/packs

- Medical-grade thermal recording paper, standard ECG grid

- Width compatible with the printer (minimum 210 mm)

- Z-fold or roll form as compatible with the machine

- Archival quality; grid clearly visible after storage

- Quantity: Minimum 2 rolls/packs supplied with each unit

Built-in Rechargeable Lithium-Ion Battery Pack

- Medical-grade rechargeable Li-ion battery

- Capacity sufficient for minimum 2 hours of continuous operation or 50 ECG records

- Charging time: Maximum 4—6 hours

- Integrated battery management with overcharge/over-discharge protection

- User-replaceable or serviceable

Main Power Supply Cord with Medical-Grade Plug

- Detachable hospital-grade AC power cord

- Length approximately 2.5—3 metres

- Region-appropriate plug (as specified by purchaser)

- Rated for 100—240 VAC, 50/60 Hz

## 1.4 Installation Requirements

- Electrical supply: 100—240 VAC, 50/60 Hz, single phase, dedicated earthed socket

- Recommended UPS backup at ward/clinic level

- Ambient temperature 10—40 °C, humidity 25—95% non-condensing

- Network point (RJ-45) recommended for EMR/HIS connectivity

- Space: Compact tabletop or mobile trolley

## 1.5 Documentation

- User Manual (English)

- Service Manual with schematics and troubleshooting

- Illustrated Parts Catalogue

- Factory Test / Quality Certificate

- Electrical Safety Test Report (IEC 62353)

- Calibration Certificate

- Preventive Maintenance Schedule

- Installation & Commissioning Report

- Certificates of compliance (CE / FDA / ISO 13485)

## 1.6 Training

- On-site clinical/application training for cardiology and nursing staff

- Biomedical engineering training on operation, first-line troubleshooting, and preventive maintenance

- Training on data management, networking, and report export

## 1.7 Warranty

- Comprehensive warranty: Minimum 2 years from date of commissioning

- Coverage: All parts, labour, preventive maintenance visits, breakdown maintenance, software updates, technical support

- Guaranteed spare parts and after-sales support for minimum 10 years

# 2. EXERCISE TOLERANCE TEST (ETT) / TREADMILL STRESS-TEST SYSTEM

Quantity: 1

## 2.1 Clinical & Functional Specifications

The system shall be a complete diagnostic stress-testing solution used to evaluate cardiovascular response to controlled physical exercise. It shall enable acquisition, display, real-time analysis, arrhythmia detection, ST-segment analysis, blood pressure monitoring, and reporting of the 12-lead ECG during rest, exercise, and recovery phases.

The system shall be suitable for cardiology departments, cardiac clinics, and pre-operative cardiac evaluation units.

Functional capabilities shall include:

- Simultaneous 12-lead ECG acquisition during rest, exercise, and recovery

- Real-time ST-segment level and slope measurement for all 12 leads

- Real-time arrhythmia detection and classification (PVC counting, ventricular tachycardia, bigeminy, trigeminy, etc.)

- Automatic and manual NIBP measurement during test (integrated or interfaced module)

- Predefined and user-configurable exercise protocols (Bruce, Modified Bruce, Naughton, Balke, Cornell, manual, etc.)

- Real-time treadmill control from operator console (speed and grade)

- Continuous heart rate calculation and target heart rate display

- Real-time full-disclosure ECG display

- Comprehensive final report with median complexes, trends, tabular summary, and interpretation

- Emergency stop with automatic treadmill deceleration

- Storage, review, and retrieval of complete test sessions

- Networking and export to HIS/EMR/cardiology information system

## 2.2 Engineering Technical Specifications

System Configuration

- Complete integrated stress-test system comprising: motorized medical treadmill, operator console/tower, medical-grade display, ECG acquisition module, telemetry/wired patient cable, stress-test software, and laser report printer

Treadmill

- Heavy-duty motorized medical-grade treadmill

- Motor rating: Minimum 2.0 HP continuous duty (AC or DC medical-grade)

- Speed range: Minimum 0.0—12.0 km/h (or 0.0—8.0 mph) continuously variable

- Elevation/grade: Minimum 0—25% adjustable

- Running surface: Minimum 1500 × 500 mm (approximate)

- Maximum patient weight capacity: Minimum 160 kg

- Safety features: Emergency stop button, side handrails, front handrail, safety key/lanyard, non-slip belt, side platforms

- Low-noise operation

- Automatic control from operator console; manual override

Operator Console / Central Workstation

- Medical-grade computer with multi-core processor (minimum quad-core, current generation)

- Minimum RAM: 8 GB (16 GB preferred)

- Storage: Minimum 500 GB SSD or equivalent

- Operating system: Current licensed medical-grade OS

- Optical drive or USB-based software media

- Minimum 4 USB ports, Ethernet (RJ-45)

Display

- Minimum 22-inch high-resolution medical-grade colour monitor

- Minimum resolution 1920 × 1080 pixels (Full HD)

- Anti-glare, wide viewing angle

ECG Acquisition

- Number of leads: 12 simultaneous

- Sampling rate: Minimum 1000 samples/second/channel

- A/D resolution: Minimum 16-bit

- Frequency response: Minimum 0.05—150 Hz

- CMRR: Minimum 100 dB

- Input impedance: Minimum 50 MΩ

- Digital filters: AC (50/60 Hz), EMG, baseline drift, motion artefact

- Real-time ST measurement resolution: 0.1 mm or better

Stress-Test Software

- Predefined protocols: Bruce, Modified Bruce, Naughton, Balke, Cornell, Ellestad, Astrand, plus user-definable

- Real-time ST-segment analysis (level and slope) for all 12 leads

- Real-time arrhythmia detection and classification

- Median complex generation and trend display

- Configurable report templates

- Full disclosure of complete test

- Storage of minimum 500 complete stress tests

Blood Pressure Monitoring

- Integrated or interfaced NIBP module (auscultatory and/or oscillometric)

- Automatic and manual measurement modes

- Configurable measurement intervals during test

Printer

- Laser printer with paper tray, minimum A4, minimum 20 ppm

UPS

- Online double-conversion UPS

- Capacity sufficient to support complete system (console + printer + monitor) for minimum 20—30 minutes

- Input: 100—240 VAC, 50/60 Hz

Power Supply (Treadmill)

- 220—240 VAC, 50/60 Hz, single phase (or as specified by purchaser)

- Dedicated earthed circuit

Environmental

- Operating temperature: 10—35 °C

- Relative humidity: 30—75% non-condensing

Electrical Safety Classification

- ECG applied part: Type CF, defibrillation-proof

- Class I equipment

Applicable Standards

- IEC 60601-1, IEC 60601-1-2 (EMC)

- IEC 60601-2-25 / 60601-2-51 (electrocardiographs)

- IEC 60601-1-6 (usability)

- IEC 62304 (software)

- ISO 13485, ISO 14971

- CE MDR / US FDA clearance or equivalent

- HL7 / DICOM compatibility

## 2.3 Standard Accessories

Heavy-Duty Motorized Medical Treadmill with Emergency Stop Switch and Safety Side Rails

- As per treadmill engineering specifications above

- Emergency stop button within immediate reach of patient and operator

- Safety side rails and front handrail, ergonomic grip

- Non-slip running belt

- Automatic deceleration on emergency stop

- Safety lanyard/key

Central Operator Console Tower with High-Performance CPU and Medical-Grade Monitor

- Medical-grade PC with specifications as listed above

- Minimum 22-inch Full HD medical-grade display

- Alphanumeric keyboard and optical mouse

- Ergonomic mobile trolley or fixed console mounting

10-Lead Stress-Test Patient ECG Cable with Active Telemetry Transmitter (or Wired equivalent)

- Medical-grade 10-lead patient cable optimized for motion-artefact rejection

- Active/buffered patient module with lightweight design

- Telemetry (wireless) or shielded wired configuration

- Defibrillation-proof, Type CF applied part

- Snap or clip connectors compatible with disposable electrodes

- IEC/AHA colour coding

- Cable length appropriate for stress testing

- Reusable, cleanable/disinfectable, latex-free

Patient Waist Belt and Cable Organizer Clip System

- Adjustable waist/chest belt for patient module attachment

- Adjustable to fit adult patient sizes

- Cable strain-relief clips/organizer to minimize motion artefact

- Reusable, washable fabric, latex-free

Dedicated Stress-Test Analysis and Reporting Software License

- Perpetual software license with all functional capabilities described above

- Free software updates during warranty period

- User management with role-based access

- Configurable report templates

- Data export in PDF/XML/DICOM/HL7 or equivalent

High-Resolution Laser Printer with Paper Tray

- Monochrome or colour laser printer

- Minimum print resolution 600 × 600 dpi

- Minimum print speed 20 ppm

- Minimum A4 paper tray with 250-sheet capacity

- USB and network connectivity

Uninterruptible Power Supply (UPS) for Full System Protection

- Online double-conversion UPS

- Capacity sized to support console, monitor, and printer for minimum 20—30 minutes

- Automatic voltage regulation, surge protection

- Sealed maintenance-free batteries

- Audible and visual status indicators

Skin Preparation Kit and Abrasive Pads

- Skin preparation kit including abrasive pads for reduction of skin impedance prior to electrode application

- Single-use, disposable

- Latex-free, biocompatible

- Adequate quantity supplied for initial commissioning use

## 2.4 Installation Requirements

- Dedicated room minimum 4 × 5 m with adequate ventilation

- Electrical supply: 220—240 VAC, 50/60 Hz, dedicated earthed circuit for treadmill (minimum 15 A)

- Separate earthed circuit for console/UPS

- Ambient temperature 10—35 °C, humidity 30—75%

- Emergency crash cart and defibrillator available on-site

- Network point (RJ-45) for EMR/HIS connectivity

- Floor: Level, non-slip, capable of supporting treadmill weight

## 2.5 Documentation

- User Manual, Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Calibration Certificate for treadmill speed/grade and NIBP

- Preventive Maintenance Schedule

- Installation and Commissioning Reports

- Compliance certificates (CE / FDA / ISO 13485)

## 2.6 Training

- Clinical/application training for cardiologists, technicians, and nurses

- Biomedical engineering training on treadmill mechanics, calibration, PM, and troubleshooting

- Software administration training

- Emergency safety and resuscitation protocol orientation

## 2.7 Warranty

- Comprehensive warranty: Minimum 2 years from commissioning

- All parts, labour, PM visits, breakdown maintenance, software updates

- Guaranteed after-sales support and spares for minimum 10 years

# 3. HOLTER ECG MONITOR / RECORDER (24-HOUR AMBULATORY)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The equipment shall be a lightweight, patient-worn digital ambulatory ECG recorder for continuous 24-hour (or extended) recording and offline analysis of the patient's ECG during routine daily activities. It shall be used for detection and evaluation of intermittent arrhythmias, ST changes, symptom—rhythm correlation, and pacemaker function assessment.

Functional capabilities shall include:

- Continuous full-disclosure recording of minimum 3-channel or 12-channel ECG for at least 24 hours (extendable to 48—72 hours where configuration allows)

- Patient event marker button for symptom logging

- Automatic lead-off detection with visual/audible indication

- High signal quality with motion-artefact rejection

- Solid-state digital recording (SD/flash memory) with retrievable data

- Compatibility with dedicated Holter analysis workstation for full-disclosure review, template matching, arrhythmia detection, ST analysis, and heart rate variability

- Pacemaker pulse detection

## 3.2 Engineering Technical Specifications

Configuration

- Compact, lightweight, patient-worn digital solid-state Holter recorder

Channels/Leads

- Minimum 3-channel (5-lead) recording; 12-channel (7-lead reconstructed 12-lead) configuration acceptable/preferred

- User-selectable configuration

Recording Duration

- Continuous recording: Minimum 24 hours; extendable to 48 or 72 hours preferred

Sampling Rate

- Minimum 1000 samples/second/channel (or equivalent)

- Optional higher rates for pacemaker detection (up to 4000 samples/second)

A/D Resolution

- Minimum 12-bit; 16-bit preferred

Frequency Response

- Minimum 0.05—100 Hz (extendable to 150 Hz where supported)

Dynamic Range

- Minimum ±5 mV

Pacemaker Detection

- Automatic pacemaker pulse detection and marking

Memory / Storage

- Removable SD card or built-in flash memory

- Capacity sufficient for minimum 72 hours of full-disclosure recording

Data Transfer

- USB or SD-card reader

- Compatible with Holter analysis workstation

Display / Indicators (on recorder)

- Small integrated LCD or LED indicators for battery status, recording status, lead-off, and time

- Patient event marker button

Weight (recorder)

- Approximately 100—300 g including battery

- Compact ergonomic form factor

Power Supply

- Standard AAA alkaline or rechargeable batteries (minimum 2 sets supplied), or integrated rechargeable battery

- Battery life: Minimum 24 hours continuous recording (48—72 hours preferred)

Environmental

- Operating temperature: 5—40 °C

- Humidity: 20—90% non-condensing

- Ingress protection: Minimum IP22 (splash resistant preferred)

Electrical Safety Classification

- Type CF applied part

- Internally powered device

Applicable Standards

- IEC 60601-1, IEC 60601-1-2 (EMC)

- IEC 60601-2-47 (ambulatory electrocardiographic systems)

- IEC 62304, ISO 13485, ISO 14971

- CE MDR / US FDA clearance or equivalent

## 3.3 Standard Accessories

Compact Digital 3-Channel or 12-Channel Holter Recorder Unit

- As per engineering specifications above

- Lightweight ergonomic housing

- Integrated event marker button, LCD/LED status indicators

- Solid-state digital recording

Reusable 5-Lead or 7-Lead Patient Cable with Snap Connectors

- Medical-grade patient cable, 5-lead (for 3-channel) or 7-lead (for 12-channel reconstruction) as compatible with the recorder

- Snap-type connectors compatible with standard disposable Ag/AgCl ECG electrodes

- Shielded, low-noise, flexible, motion-artefact tolerant

- Anti-tangle design

- Latex-free, biocompatible per ISO 10993

- IEC/AHA colour coding

- Reusable, cleanable, and disinfectable

Heavy-Duty Protective Pouch with Adjustable Waist Belt and Shoulder Strap

- Padded protective pouch to house the recorder during ambulatory recording

- Adjustable waist belt and detachable shoulder strap

- Reusable, washable fabric construction

- Cable exit port with strain relief

- Latex-free

USB Data Transfer Cable or High-Speed SD Card Reader

- USB data transfer cable (compatible connector) or dedicated high-speed SD card reader

- USB 2.0 or higher

- Compatible with Holter analysis workstation

Rechargeable or Long-Life AAA Batteries (2 sets)

- Two sets of AAA cells, alkaline long-life or rechargeable (NiMH minimum 800 mAh)

- Each set sufficient for minimum one full 24-hour recording session

- If rechargeable: include compatible charger

## 3.4 Installation Requirements

- Minimal installation; hand-over training only

- Analysis workstation to be linked with this recorder (see separate item)

- Ambient storage: 5—40 °C, humidity 20—90% non-condensing

## 3.5 Documentation

- User Manual, Service Manual, Parts Catalogue

- Factory Test / Quality Certificate

- Electrical Safety Test Report (IEC 62353)

- Preventive Maintenance Schedule

- Compliance certificates (CE / FDA / ISO 13485)

## 3.6 Training

- Clinical application training (electrode placement, patient briefing, hook-up procedure)

- Biomedical engineering training (battery management, cleaning, troubleshooting, PM)

- Data download and workstation interfacing training

## 3.7 Warranty

- Comprehensive warranty: Minimum 2 years from commissioning

- All parts, labour, PM, breakdown maintenance, software updates

- Spare parts guaranteed for minimum 7—10 years

# 4. HOLTER ANALYSIS WORKSTATION

Quantity: 1

## 4.1 Clinical & Functional Specifications

The workstation shall be a dedicated, high-performance computer-based analysis system for offline processing, review, editing, and reporting of ambulatory (Holter) ECG recordings acquired by the associated Holter recorder(s). It shall provide full-disclosure ECG review, automatic beat detection, template matching, arrhythmia classification, ST-segment analysis, heart rate variability, pacemaker analysis, and comprehensive report generation.

Functional capabilities shall include:

- Import of recordings from removable media (SD card / USB)

- Automatic beat detection and classification (normal, supraventricular, ventricular, paced, artefact)

- Template-based matching with user editing

- Full-disclosure display with configurable time compression

- Arrhythmia detection (VT, VF, SVT, AF, bradycardia, tachycardia, pauses, bigeminy, trigeminy, R-on-T, etc.)

- ST-segment level and slope analysis for all channels

- Heart rate variability (time-domain, frequency-domain, non-linear)

- Pacemaker analysis (capture, sensing, pacing counters) where applicable

- Patient diary and event correlation

- Configurable, editable summary and full report templates

- Export in PDF, XML, DICOM, HL7 or equivalent

- Networking to HIS/EMR

- Multi-user login with audit trail

## 4.2 Engineering Technical Specifications

Configuration

- Complete Holter analysis workstation comprising: computer tower, medical-grade display, analysis software, data-reader, keyboard/mouse, laser printer, UPS

Computer Tower

- Multi-core processor: Minimum current-generation quad-core (hexa/octa-core preferred)

- RAM: Minimum 16 GB DDR4 or better

- Storage: Minimum 500 GB SSD + 1 TB HDD (or equivalent single 1 TB SSD)

- Graphics: Dedicated or integrated GPU with dual-monitor support

- Optical drive: DVD writer (optional)

- USB ports: Minimum 4 (with at least 2 × USB 3.0)

- Ethernet (RJ-45) 1 Gbps

- Operating system: Current licensed medical/enterprise-grade OS

Display

- Minimum 24-inch high-resolution medical-grade diagnostic monitor

- Minimum resolution 1920 × 1080 pixels (Full HD); higher resolution preferred

- Anti-glare, wide viewing angle

- LED backlight

Analysis Software

- Perpetual, comprehensive software license with all functional capabilities described above

- Free software updates during warranty period

- Multi-user role-based access with audit trail

- Configurable report templates

- Compatible with associated Holter recorder(s)

- Data export in PDF/XML/DICOM/HL7 or equivalent

Data Reader

- High-speed USB optical / SD-card reader compatible with the Holter recorder storage media

Input Devices

- Alphanumeric keyboard and optical/laser mouse

Printer

- Laser printer for diagnostic reports

- Minimum resolution 600 × 600 dpi

- Minimum print speed 20 ppm

- Minimum A4 paper tray with 250-sheet capacity

- USB and network connectivity

UPS

- Online double-conversion UPS with surge protection

- Capacity sufficient to support complete workstation (tower + monitor + printer) for minimum 20—30 minutes

- Automatic voltage regulation

- Sealed maintenance-free batteries

Power Supply

- 100—240 VAC, 50/60 Hz, single phase, earthed

Environmental

- Operating temperature: 10—35 °C

- Relative humidity: 30—75% non-condensing

Applicable Standards

- IEC 60601-1 / IEC 60950 / IEC 62368 (as applicable to computer components)

- IEC 62304 (medical software)

- IEC 62366 (usability)

- IEC 60601-2-47 (compatibility with ambulatory ECG systems)

- IEC 80001 (IT network risk management)

- ISO 13485, ISO 14971

- CE MDR / FDA clearance or equivalent

- HL7 / DICOM compatibility

## 4.3 Standard Accessories

High-Performance Computer Tower Workstation with Multi-Core Processor

- As per engineering specifications above (minimum quad-core, ≥16 GB RAM, ≥500 GB SSD, licensed OS, adequate connectivity)

High-Resolution 24-Inch Medical-Grade Diagnostic Display Monitor

- Minimum 24-inch diagonal, minimum 1920 × 1080 pixels, medical-grade with anti-glare coating

- Wide viewing angle (minimum 170° H/V)

- Adjustable tilt/height/swivel stand

Advanced Holter Analysis, Template Matching, and Arrhythmia Reporting Software License

- Comprehensive perpetual license with all functional and analytical capabilities described above

- Free software updates during warranty period

- Compatible with the associated Holter recorder(s)

- Configurable report templates and export formats

High-Speed USB Optical Card Reader for Recorder Data Download

- USB 2.0/3.0 optical / SD-card reader

- Compatible with recorder storage medium

- Plug-and-play

Alphanumeric Keyboard and Optical Mouse

- Standard 104-key USB keyboard, spill-resistant, cleanable

- Optical/laser USB mouse with scroll wheel

- Both cleanable/disinfectable with hospital-grade agents

Laser Printer for Diagnostic Report Generation

- As per printer specifications above (≥600 dpi, ≥20 ppm, A4, USB + network)

Uninterruptible Power Supply (UPS) with Surge Protection

- Online double-conversion UPS

- Capacity sized for complete workstation, minimum 20—30 minutes backup

- Surge protection, AVR, sealed maintenance-free batteries

## 4.4 Installation Requirements

- Dedicated workstation area with ergonomic desk and chair

- Electrical supply: 100—240 VAC, 50/60 Hz, dedicated earthed socket

- Network point (RJ-45) for EMR/HIS connectivity

- Ambient temperature 10—35 °C, humidity 30—75%

- Adequate lighting for diagnostic display viewing

## 4.5 Documentation

- User Manual, Service Manual, Software Manual

- Software license certificates

- Factory Test Certificate

- Electrical Safety Test Report

- Installation and Commissioning Report

- Compliance certificates (CE / FDA / ISO 13485)

## 4.6 Training

- Clinical/application training for cardiologists and technicians

- Software administration and user management training

- Biomedical engineering training for hardware maintenance and troubleshooting

- Data management, backup, and networking training

## 4.7 Warranty

- Comprehensive warranty: Minimum 2 years from commissioning

- All parts, labour, PM, breakdown maintenance, software updates, technical support

- Spare parts guaranteed for minimum 7—10 years

# 5. HOLTER BP / AMBULATORY BP MONITOR (ABPM)

Quantity: 1

## 5.1 Clinical & Functional Specifications

The equipment shall be a lightweight, patient-worn oscillometric ambulatory blood-pressure monitor for automatic, programmable 24-hour non-invasive blood-pressure measurement during patient's daily activities and sleep. It shall be used for diagnosis and management of hypertension, evaluation of white-coat and masked hypertension, nocturnal BP profile, and antihypertensive therapy monitoring.

Functional capabilities shall include:

- Automatic programmable NIBP measurement at user-configurable intervals during daytime and night-time

- Oscillometric measurement algorithm validated for ambulatory use

- Storage of systolic, diastolic, mean arterial pressure, heart rate, and time-stamp

- Patient event/manual measurement button

- Motion artefact rejection with automatic retry

- Audible alarm silencing during night-time

- Data download via USB or dedicated interface to reporting software

- Comprehensive report generation including 24-hour trends, day/night averages, dipping status, BP load, variability, and interpretation

- Compatibility with reporting software installed on standard PC

## 5.2 Engineering Technical Specifications

Configuration

- Compact, lightweight, patient-worn digital oscillometric ABPM recorder with dedicated reporting software

Measurement Technique

- Oscillometric method, clinically validated (per ESH / BHS / AAMI / ISO 81060-2 validation protocols)

Measurement Ranges

- Systolic pressure: 60—260 mmHg (minimum)

- Diastolic pressure: 30—200 mmHg (minimum)

- Mean arterial pressure: 40—220 mmHg (minimum)

- Pulse/heart rate: 30—200 bpm (minimum)

Accuracy

- BP accuracy: Within ±3 mmHg (or ±2% of reading, whichever is greater)

- Heart rate accuracy: Within ±5% or ±3 bpm

Cuff Pressure Range

- 0—300 mmHg

- Maximum cuff pressure with automatic over-pressure protection

Measurement Intervals

- User programmable, typically 5—120 minutes

- Separate daytime and night-time interval settings

Recording Duration

- Minimum 24 hours; extendable to 48 hours preferred

- Storage capacity: Minimum 250 measurements

Motion Artefact Handling

- Automatic detection and rejection of motion-affected readings with automatic retry after configurable delay

Display / Indicators

- Integrated LCD or LED display showing time, measurement status, last reading, battery status

- Manual measurement button, event marker button

- Audible signal with configurable night-time silence

Data Transfer

- USB interface cable or wireless (as configured)

- Compatible with reporting software on standard PC (Windows or equivalent)

Reporting Software

- Perpetual license

- 24-hour trends (BP and HR)

- Day/night averages, dipping status calculation

- BP load, variability indices

- Configurable report templates

- Export in PDF/XML/HL7 or equivalent

Power Supply

- 2 × AA alkaline or rechargeable NiMH batteries (minimum one set supplied)

- Battery life: Minimum one complete 24-hour recording session per set

Weight (recorder)

- Approximately 200—350 g including batteries and cuff-tube connector

Environmental

- Operating temperature: 5—40 °C

- Humidity: 15—90% non-condensing

- Ingress protection: Minimum IP22 (splash resistant preferred)

Electrical Safety Classification

- Type BF applied part

- Internally powered

Applicable Standards

- IEC 60601-1, IEC 60601-1-2 (EMC)

- IEC 60601-2-30 (automated non-invasive sphygmomanometers)

- ISO 81060-2 (clinical validation of NIBP)

- IEC 62304, ISO 13485, ISO 14971

- CE MDR / US FDA clearance or equivalent

- Validation per ESH / BHS / AAMI protocols

## 5.3 Standard Accessories

Compact Lightweight Digital ABPM Recorder Unit

- As per engineering specifications above

- Ergonomic lightweight housing suitable for waist/belt wear

- Integrated LCD/LED indicators, manual measurement and event buttons

- Oscillometric measurement, digital solid-state storage

Reusable Adult NIBP Cuff (Medium Size)

- Reusable adult medium-size NIBP cuff (approximate arm circumference range 22—32 cm)

- Two-tube or single-tube configuration compatible with the recorder

- Durable nylon/polyester construction with hook-and-loop closure

- Artery indicator marking and size range clearly labelled

- Latex-free, biocompatible per ISO 10993

- Cleanable and disinfectable with hospital-grade agents

- Meets accuracy requirements per ISO 81060-2

Reusable Adult NIBP Cuff (Large Size)

- Reusable adult large-size NIBP cuff (approximate arm circumference range 32—42 cm)

- Same construction, material, compatibility, and compliance as medium cuff

- Latex-free, cleanable, disinfectable

Flexible Air Connecting Tube with Quick-Connect Couplings

- Medical-grade PVC or equivalent flexible air tubing

- Kink-resistant construction

- Quick-connect couplings compatible with recorder and cuff

- Length appropriate for ambulatory use (approximately 0.8—1.2 m)

- Latex-free

Heavy-Duty Fabric Carrying Pouch with Adjustable Belt Straps

- Padded protective pouch for the recorder

- Adjustable waist belt and optional shoulder strap

- Reusable, washable fabric

- Cable/tube exit port with strain relief

- Latex-free

ABPM Reporting and Analysis Software License with USB Interface Cable

- Perpetual software license with full functional capabilities described above

- Compatible with standard PC (Windows or equivalent)

- Free updates during warranty period

- USB interface cable for data transfer, compatible connector

- Data export in PDF/XML/HL7 or equivalent

AA Alkaline or Rechargeable Battery Set

- Minimum one complete set (typically 2 × AA) alkaline long-life or rechargeable NiMH (≥2000 mAh)

- Sufficient for minimum one full 24-hour recording session

- If rechargeable: compatible charger included

## 5.4 Installation Requirements

- Minimal installation

- Reporting software to be installed on a designated PC (may share Holter analysis workstation or a separate PC as specified by purchaser)

- Ambient storage: 5—40 °C, humidity 15—90% non-condensing

## 5.5 Documentation

- User Manual, Service Manual, Parts Catalogue

- Clinical Validation Certificate (per ESH / BHS / AAMI / ISO 81060-2)

- Factory Test / Quality Certificate

- Electrical Safety Test Report (IEC 62353)

- Calibration Certificate

- Preventive Maintenance Schedule

- Compliance certificates (CE / FDA / ISO 13485)

## 5.6 Training

- Clinical application training (cuff selection, hook-up, patient briefing, interval programming)

- Biomedical engineering training (battery management, calibration verification, PM, troubleshooting)

- Reporting software training (data download, report generation, export)

## 5.7 Warranty

- Comprehensive warranty: Minimum 2 years from commissioning

- All parts, labour, PM, breakdown maintenance, software updates, technical support

- Guaranteed spare parts for minimum 7—10 years