HOSPITAL PROCUREMENT — TECHNICAL SPECIFICATIONS ENT

ENT / Audiology / Diagnostic Department

Table of Contents

1. Diagnostic Two-Channel Audiometer	7

1.1 Clinical & Functional Specifications	7

1.2 Engineering Technical Specifications	7

1.3 Standard Accessories	8

1.4 Installation Requirements	9

1.5 Documentation	9

1.6 Training	9

1.7 Warranty	9

2. Impedance Audiometer / Tympanometer	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	10

2.4 Installation Requirements	11

2.5 Documentation	11

2.6 Training	11

2.7 Warranty	11

3. BERA Machine (Auditory Brainstem Response System)	12

3.1 Clinical & Functional Specifications	12

3.2 Engineering Technical Specifications	12

3.3 Standard Accessories	13

3.4 Installation Requirements	13

3.5 Documentation	13

3.6 Training	13

3.7 Warranty	14

4. Rigid ENT Endoscope Set with Camera, Light Source and HD Monitor	15

4.1 Clinical & Functional Specifications	15

4.2 Engineering Technical Specifications	15

4.3 Standard Accessories	15

4.4 Installation Requirements	16

4.5 Documentation	16

4.6 Training	16

4.7 Warranty	16

5. Flexible Nasopharyngoscope	17

5.1 Clinical & Functional Specifications	17

5.2 Engineering Technical Specifications	17

5.3 Standard Accessories	17

5.4 Installation Requirements	17

5.5 Documentation	18

5.6 Training	18

5.7 Warranty	18

6. FESS Instrument Set	19

6.1 Clinical & Functional Specifications	19

6.2 Engineering Technical Specifications	19

6.3 Standard Accessories	19

6.4 Installation Requirements	19

6.5 Documentation	19

6.6 Training	19

6.7 Warranty	20

7. Mastoidectomy Instrument Set	21

7.1 Clinical & Functional Specifications	21

7.2 Engineering Technical Specifications	21

7.3 Standard Accessories	21

7.4 Installation Requirements	21

7.5 Documentation	21

7.6 Training	21

7.7 Warranty	21

8. Tympanoplasty Instrument Set	22

8.1 Clinical & Functional Specifications	22

8.2 Engineering Technical Specifications	22

8.3 Standard Accessories	22

8.4 Installation Requirements	22

8.5 Documentation	22

8.6 Training	22

8.7 Warranty	22

9. Stapedectomy / Stapedotomy Instrument Set	23

9.1 Clinical & Functional Specifications	23

9.2 Engineering Technical Specifications	23

9.3 Standard Accessories	23

9.4 Installation Requirements	23

9.5 Documentation	23

9.6 Training	23

9.7 Warranty	23

10. ENT Operating Microscope	24

10.1 Clinical & Functional Specifications	24

10.2 Engineering Technical Specifications	24

10.3 Standard Accessories	24

10.4 Installation Requirements	25

10.5 Documentation	25

10.6 Training	25

10.7 Warranty	25

11. Battery-Operated ENT Micro-Drill	26

11.1 Clinical & Functional Specifications	26

11.2 Engineering Technical Specifications	26

11.3 Standard Accessories	26

11.4 Installation Requirements	26

11.5 Documentation	26

11.6 Training	26

11.7 Warranty	27

12. ENT Examination Unit	28

12.1 Clinical & Functional Specifications	28

12.2 Engineering Technical Specifications	28

12.3 Standard Accessories	28

12.4 Installation Requirements	28

12.5 Documentation	29

12.6 Training	29

12.7 Warranty	29

13. ENT Diagnostic and Minor-Procedure Set	30

13.1 Clinical & Functional Specifications	30

13.2 Engineering Technical Specifications	30

13.3 Standard Accessories	30

13.4 Installation Requirements	30

13.5 Documentation	30

13.6 Training	30

13.7 Warranty	30

14. ENT Microscope for OPD	31

14.1 Clinical & Functional Specifications	31

14.2 Engineering Technical Specifications	31

14.3 Standard Accessories	31

14.4 Installation Requirements	31

14.5 Documentation	31

14.6 Training	31

14.7 Warranty	31

15. Microlaryngeal Instruments Set	33

15.1 Clinical & Functional Specifications	33

15.2 Engineering Technical Specifications	33

15.3 Standard Accessories	33

15.4 Installation Requirements	33

15.5 Documentation	33

15.6 Training	33

15.7 Warranty	33

16. Telescope for Microlaryngoscopy	34

16.1 Clinical & Functional Specifications	34

16.2 Engineering Technical Specifications	34

16.3 Standard Accessories	34

16.4 Installation Requirements	34

16.5 Documentation	34

16.6 Training	34

16.7 Warranty	34

17. Rigid Oesophagoscope Set	35

17.1 Clinical & Functional Specifications	35

17.2 Engineering Technical Specifications	35

17.3 Standard Accessories	35

17.4 Installation Requirements	35

17.5 Documentation	35

17.6 Training	35

17.7 Warranty	35

18. Tonsillectomy and Adenoidectomy Instrument Set	36

18.1 Clinical & Functional Specifications	36

18.2 Engineering Technical Specifications	36

18.3 Standard Accessories	36

18.4 Installation Requirements	36

18.5 Documentation	36

18.6 Training	36

18.7 Warranty	36

19. Septoplasty Instrument Set	37

19.1 Clinical & Functional Specifications	37

19.2 Engineering Technical Specifications	37

19.3 Standard Accessories	37

19.4 Installation Requirements	37

19.5 Documentation	37

19.6 Training	37

19.7 Warranty	37

20. Rhinoplasty Instrument Set	38

20.1 Clinical & Functional Specifications	38

20.2 Engineering Technical Specifications	38

20.3 Standard Accessories	38

20.4 Installation Requirements	38

20.5 Documentation	38

20.6 Training	38

20.7 Warranty	38

21. Rigid Bronchoscope Set	39

21.1 Clinical & Functional Specifications	39

21.2 Engineering Technical Specifications	39

21.3 Standard Accessories	39

21.4 Installation Requirements	39

21.5 Documentation	39

21.6 Training	39

21.7 Warranty	39

# 1. Diagnostic Two-Channel Audiometer

Quantity: 1

## 1.1 Clinical & Functional Specifications

The equipment shall be a clinical-grade, microprocessor-controlled, two-channel diagnostic audiometer intended for comprehensive pure-tone, speech, and masking audiometry in adult and pediatric patients within an ENT / Audiology department. It shall support air-conduction, bone-conduction, and sound-field testing with independent channel control for stimulus and masking. It shall permit determination of hearing thresholds, speech reception thresholds (SRT), speech discrimination scores, SISI, tone decay, ABLB, Stenger, Békésy, and other advanced tests as clinically required. It shall include talk-forward, talk-back, and monitor functions for patient-operator communication and support both manual and automated testing workflows with real-time audiogram plotting.

## 1.2 Engineering Technical Specifications

- Type: Clinical diagnostic audiometer, Type 1 or Type 2 per IEC 60645-1, or equivalent.

- Channels: Two independent channels, fully configurable for stimulus/masking on either channel.

- Display: Integrated color TFT/LCD display, minimum 5-inch diagonal, or equivalent PC-based graphical user interface with minimum resolution 800 × 480 pixels or better.

- Controls: Membrane keypad, rotary encoders, or touchscreen; ergonomically arranged for single-hand operation.

- Frequency Range: Minimum 125 Hz to 8000 Hz for air conduction; minimum 250 Hz to 8000 Hz for bone conduction; extended high-frequency capability up to 16 kHz configurable/optional.

- Frequency Steps: Selectable ½-octave and 1-octave; frequency accuracy within ±1% or better.

- Intensity Range:

-   Air conduction: minimum −10 dB HL to +120 dB HL (frequency-dependent).

-   Bone conduction: minimum −10 dB HL to +70 dB HL (frequency-dependent).

-   Sound-field: minimum −10 dB HL to +90 dB HL.

- Intensity Steps: Selectable 1, 2, and 5 dB steps.

- Level Accuracy: Within ±3 dB or better across the audiometric range.

- Stimulus Types: Pure tone (continuous / pulsed), warble tone, narrow-band noise, white noise, speech noise, external input (CD / MP3 / microphone / PC).

- Masking: Independent masking on the non-test ear with automatic or manual control; narrow-band, white, and speech-weighted noise.

- Speech Audiometry: Live-voice and recorded speech input; VU-meter or equivalent level indicator; word-list management.

- Test Protocols: Manual, semi-automatic, and automatic Hughson—Westlake threshold determination.

- Patient Response: Digital patient response indicator with automatic logging.

- Talk-Forward / Talk-Back / Monitor: Independent adjustable channels with volume control.

- Data Management: On-board memory for a minimum of 100 patient sessions or PC-based unlimited storage; audiogram export in standard formats (PDF, XML, or equivalent); NOAH-compatible interface.

- Connectivity: USB, RS-232, or LAN; wireless connectivity configurable/optional.

- Printer: Integrated thermal printer or PC-based printout.

- Power Supply: 100—240 VAC, 50/60 Hz; internal medical-grade PSU compliant with IEC 60601-1.

- Environmental: Operating temperature 15—35 °C; relative humidity 30—90 % non-condensing.

- Dimensions/Weight: Compact tabletop configuration; weight approximately 3—8 kg.

- Safety Class: Class II or Class I with Type B applied parts, or equivalent per IEC 60601-1.

- Applicable Standards: IEC 60645-1, IEC 60645-2, IEC 60601-1, IEC 60601-1-2 (EMC), ISO 389 series (reference calibration), ISO 8253-1/2/3, CE MDR, FDA or equivalent regulatory clearance, ISO 13485 for the manufacturer.

## 1.3 Standard Accessories

- Reusable Air-Conduction Headphones (TDH-39 or equivalent): Supra-aural audiometric headphones; frequency response compliant with ISO 389-1; cushioned earpads; adjustable headband; connector compatible with the audiometer; calibrated as a matched pair; calibration certificate traceable to national standards.

- Reusable Bone-Conduction Headset (B71 or equivalent): Piezoelectric or electromagnetic bone vibrator; frequency response compliant with ISO 389-3; steel spring headband with defined static force; calibration certificate traceable to national standards.

- Patient Response Signal Switch with Coiled Cable: Ergonomic push-button switch; medical-grade housing; coiled cable of adequate length; robust connector; cleanable/disinfectable surface.

- Free-Field Loudspeakers (1 pair): Active or passive audiometric loudspeakers; frequency response suitable for warble tone and speech; magnetically shielded; compatible with the audiometer's sound-field channels; mounting hardware included; calibrated for use at defined test distance.

- Talk-Back Microphone and Operator Monitor Headset: Sensitive omnidirectional microphone with boom or table mount; matched operator monitor headset for talk-back and monitoring; low-noise pre-amplification.

- Integrated Thermal Printer or USB PC Export Cable: Thermal printer with paper roll (or USB cable of adequate length for PC connection); print/export of audiogram, patient data, and test results.

- Audiometric Data Management Software License: Full clinical license, NOAH-compatible or equivalent; supports patient database, audiogram plotting, report generation, PDF export, and network connectivity; free version updates during warranty.

- Audiogram Chart Sheets: Standard pre-printed pack of audiogram forms; minimum 500 sheets or equivalent.

## 1.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz single-phase supply from UPS-protected medical socket with earth-continuity \< 0.2 Ω. Installation inside an audiometric sound-treated booth conforming to ISO 8253-1 ambient-noise limits, or an equivalent low-noise room. Ambient noise levels shall permit valid threshold measurements down to 0 dB HL. Adequate cable routing for headphones, bone vibrator, patient response switch, and loudspeakers. Ergonomic table/desk space and operator seating.

## 1.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate traceable to national/international standards, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), Preventive Maintenance Schedule, Installation & Commissioning Report — all in English; hard and soft copies.

## 1.6 Training

On-site clinical/application training for audiologists and ENT staff; biomedical engineering training on preventive maintenance, calibration verification, and basic troubleshooting; software training for data management.

## 1.7 Warranty

Comprehensive on-site warranty minimum two (2) years covering labour, spare parts, annual calibration, preventive maintenance, breakdown maintenance, software updates, and technical support.

# 2. Impedance Audiometer / Tympanometer

Quantity: 1

## 2.1 Clinical & Functional Specifications

Microprocessor-controlled, clinical-grade middle-ear analyzer intended for tympanometry, acoustic reflex testing (ipsilateral and contralateral), reflex decay, and Eustachian tube function testing in adult and pediatric patients. It shall automatically measure and plot tympanograms, determine ear canal volume, middle-ear compliance and pressure, and elicit acoustic reflex thresholds. It shall provide automatic and manual test modes with audible/visual guidance and shall generate clinically formatted printed reports.

## 2.2 Engineering Technical Specifications

- Type: Clinical diagnostic tympanometer conforming to IEC 60645-5 (Type 1) or equivalent.

- Probe Tone: 226 Hz standard; additional 678 Hz, 800 Hz, and 1000 Hz selectable for neonatal/pediatric applications.

- Pressure Range: Minimum −600 to +400 daPa, adjustable in configurable steps.

- Pressure Accuracy: Within ±10 daPa or ±10 %, whichever is greater.

- Pump Speed: Selectable/configurable (e.g., 50, 100, 200, 400 daPa/s or auto).

- Compliance Range: Minimum 0.1 to 6.0 ml (or mmho), with accuracy within ±5 % or ±0.1 ml.

- Reflex Stimulus: Pure tone 500, 1000, 2000, 4000 Hz; broadband noise (BBN); high-pass and low-pass noise; external input.

- Reflex Intensity Range: Ipsilateral up to minimum 100 dB HL; contralateral up to minimum 120 dB HL; adjustable in 1, 2, and 5 dB steps.

- Reflex Tests: Threshold, decay, latency (configurable).

- Eustachian Tube Function Tests: Williams (intact TM), Toynbee/Valsalva sequences (perforated TM).

- Display: Color graphic display, minimum 5-inch diagonal, or PC-based; minimum resolution 800 × 480 pixels.

- Data Storage: Internal storage of minimum 100 test sessions or PC-based unlimited.

- Printer: Integrated thermal strip printer; PC report export in PDF or equivalent.

- Connectivity: USB and/or RS-232/LAN; NOAH-compatible module or equivalent.

- Power Supply: External medical-grade adapter, 100—240 VAC, 50/60 Hz; internal rechargeable battery configurable/optional.

- Weight: Approximately 2—5 kg for handheld/portable; up to 8 kg for tabletop configuration.

- Safety Class: Class II or Class I, Type B applied parts, or equivalent.

- Applicable Standards: IEC 60645-5, IEC 60601-1, IEC 60601-1-2, ANSI S3.39, CE MDR, FDA or equivalent, ISO 13485.

## 2.3 Standard Accessories

- Handheld Impedance Probe with Status LEDs: Ergonomic probe handle with color/LED indicators showing seal status (leak / blocked / ready); shielded cable; medical-grade housing.

- Reusable Silicone Ear-Tip Starter Kit: Assorted diameters (typical range 3—22 mm) suitable for infant to adult ear canals; medical-grade silicone; latex-free; autoclavable or high-level disinfectable per manufacturer instructions.

- Contralateral Stimulation Earphone: Insert or supra-aural earphone; frequency response suitable for reflex stimulation up to minimum 120 dB HL; calibrated and traceable.

- Calibration Test Cavity (2 cc / 0.5 cc): Rigid metal test cavity of standardized volume for daily biological/electroacoustic verification per IEC 60645-5.

- Integrated Thermal Strip Printer with 2 Paper Rolls: Prints tympanogram and reflex results; standard thermal paper of specified width.

- PC Data Interface Cable: USB or equivalent; suitable length for tabletop use.

- Main Power Supply Adapter & Cable: Medical-grade adapter compliant with IEC 60601-1; regional plug per purchaser.

## 2.4 Installation Requirements

Single-phase 100—240 VAC, 50/60 Hz supply from UPS-protected medical socket; earth continuity \< 0.2 Ω. Quiet examination environment; ambient temperature 15—35 °C; RH 30—90 % non-condensing. Adequate desk area for tabletop unit or storage cradle for portable device.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, IEC 62353 Electrical Safety Report, PM Schedule, Installation & Commissioning Report.

## 2.6 Training

Clinical/application training for audiologists and ENT staff; biomedical engineering training on daily verification, calibration, and troubleshooting.

## 2.7 Warranty

Minimum two (2) years comprehensive on-site warranty including labour, spare parts, annual calibration, PM, breakdown maintenance, software updates, and technical support.

# 3. BERA Machine (Auditory Brainstem Response System)

Quantity: 1

## 3.1 Clinical & Functional Specifications

PC-based, multi-channel evoked-potential platform intended for auditory brainstem response (ABR/BERA), auditory steady-state response (ASSR), otoacoustic emissions (configurable/optional), electrocochleography (ECochG), and neonatal hearing screening. The system shall provide threshold-seeking and neuro-diagnostic ABR with click and tone-burst stimuli via air conduction (insert/supra-aural) and bone conduction. It shall support automated screening protocols with clear PASS/REFER indication as well as full diagnostic protocols with waveform averaging, latency-intensity plots, and clinical reporting.

## 3.2 Engineering Technical Specifications

- Channels: Minimum 2 EEG acquisition channels; expandable configurable/optional.

- Amplifier: Isolated, low-noise, high-CMRR pre-amplifier; input noise ≤ 25 nV/√Hz or equivalent; CMRR ≥ 110 dB; input impedance ≥ 10 MΩ.

- A/D Conversion: Minimum 24-bit resolution.

- Sampling Rate: Minimum 20 kHz/channel; higher configurable.

- Filters: Digital high-pass and low-pass filters, selectable within minimum 0.1 Hz to 5 kHz; notch filter 50/60 Hz.

- Artifact Rejection: Automatic and manual, adjustable thresholds.

- Stimulus Types: Click, tone burst (500, 1000, 2000, 4000 Hz and configurable), chirp; rarefaction, condensation, and alternating polarity.

- Stimulus Intensity: Up to minimum 100 dB nHL for air conduction; up to minimum 70 dB nHL for bone conduction; adjustable in 1, 5, 10 dB steps.

- Stimulus Rate: Adjustable, typically 0.1—90 Hz.

- Transducers Supported: Insert earphones, supra-aural earphones, bone vibrator, sound-field speakers.

- Software: Dedicated clinical acquisition and analysis software with automated peak marking, latency-intensity plots, waveform overlay, normative data, report generation, PDF export, and database management; multi-language user interface configurable.

- Screening Mode: Automated ABR (AABR) with objective PASS/REFER algorithm suitable for newborn screening.

- Data Storage: Unlimited on PC hard drive/server; encrypted patient database.

- Connectivity: USB, LAN; DICOM or HL7 export configurable/optional; NOAH-compatible configurable.

- Computer: Laptop or desktop with minimum quad-core processor, minimum 8 GB RAM, minimum 256 GB SSD, medical-grade or equivalent-quality display; suitable OS as supported by the software vendor.

- Power Supply: 100—240 VAC, 50/60 Hz via UPS.

- Safety Class: Type BF applied parts; patient-isolated amplifier compliant with IEC 60601-1.

- Applicable Standards: IEC 60645-3 (short-duration stimuli), IEC 60645-7 (ABR), IEC 60601-1, IEC 60601-1-2, IEC 62304 (software lifecycle), CE MDR, FDA or equivalent, ISO 13485, ISO 14971.

## 3.3 Standard Accessories

- Multichannel ABR/BERA Pre-amplifier and Isolation Module: Patient-isolated, low-noise amplifier; shielded housing; snap or DIN connectors; robust cable strain relief.

- High-Shielded Insert Earphones (ER-3A or equivalent): Frequency response suitable for click/tone-burst ABR; shielded to minimize electromagnetic artifact; disposable foam ear tips supplied in adequate starter quantity.

- Bone-Conduction Transducer Headset: Calibrated bone vibrator with steel spring headband; calibration certificate.

- Reusable Electrode Lead Wires with Snap/Clip Connectors: Color-coded lead wires; low-noise; touch-proof safety connectors; suitable length; cleanable/disinfectable.

- High-Performance Laptop/Desktop with Monitor: Specifications as per software vendor requirements; medical-grade or equivalent display; keyboard, mouse, and required cables.

- Mobile System Trolley with Articulated Pre-amp Arm: Medical-grade mobile cart; lockable antistatic casters; articulated arm for pre-amplifier positioning; cable management; shelves for PC and printer.

- Dedicated BERA/ABR Software Package: Full clinical license; unlimited patient records; free version updates during warranty.

- Uninterruptible Power Supply (UPS): Online double-conversion UPS sized to full system load; minimum 30 minutes runtime at full load; sine-wave output.

- Rigid Transit Carrying Case: Impact-resistant, foam-lined; transport case for portable system components.

## 3.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz single-phase supply from UPS; earth continuity \< 0.2 Ω; electrically quiet room, preferably with electromagnetic shielding or away from strong EMI sources (MRI, elevator motors, large transformers). Ambient temperature 18—28 °C, RH 30—70 %. Adequate space for trolley, patient bed/reclining chair, and operator workstation. Network connection for archiving/HIS integration configurable.

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, IEC 62353 Electrical Safety Report, PM Schedule, Installation & Commissioning Report, Software Licence documents.

## 3.6 Training

Comprehensive clinical/application training for audiologists and ENT physicians; biomedical engineering training on PM, electrode integrity checks, calibration verification, and troubleshooting; IT training on data backup and database management.

## 3.7 Warranty

Minimum two (2) years comprehensive on-site warranty including labour, spare parts, annual calibration, PM, breakdown maintenance, software updates, and technical support.

# 4. Rigid ENT Endoscope Set with Camera, Light Source and HD Monitor

Quantity: 2

## 4.1 Clinical & Functional Specifications

Complete rigid endoscopic visualization system intended for ENT diagnostic and therapeutic procedures including nasal endoscopy, sinus surgery, laryngoscopy, and otoscopy. The system shall provide high-definition, distortion-free imaging with excellent color reproduction; support still image capture, video recording, and PACS/HIS export; and be mounted on a mobile equipment tower for use in OT and OPD.

## 4.2 Engineering Technical Specifications

- Telescopes: Rigid, autoclavable, stainless steel; wide-angle field of view minimum 80°; high-resolution optical fibers.

- Camera Head: 1-chip or 3-chip CMOS/CCD HD or Full-HD; sensitivity ≥ 1 lux (typical); progressive scan; automatic and manual white balance; auto-exposure; sealed, submersible for reprocessing per manufacturer instructions.

- Camera Resolution: Minimum 1920 × 1080 (Full HD); 4K UHD configurable/optional.

- Video Processor / CCU: Digital signal processing with edge enhancement, noise reduction, and multiple image modes.

- Video Outputs: HD-SDI, DVI, HDMI, or DisplayPort; minimum two independent outputs.

- Capture: Still image and video capture to USB, SD, or network storage; DICOM export configurable/optional.

- Light Source: Medical-grade LED; luminous output equivalent to minimum 300 W xenon or better; color temperature approximately 5500—6500 K; adjustable intensity; lamp life ≥ 20,000 hours for LED; automatic fan cooling.

- Monitor: Medical-grade LCD/LED, minimum 24-inch diagonal; minimum resolution 1920 × 1080; brightness ≥ 400 cd/m²; contrast ratio ≥ 1000:1; DICOM Part 14 grayscale configurable; wide viewing angle ≥ 178°.

- Trolley: Medical-grade mobile cart with lockable antistatic casters, isolation transformer, shelves for CCU/light source/recorder, and cable management.

- Power Supply: 100—240 VAC, 50/60 Hz; isolation transformer integral to trolley.

- Reprocessing: Autoclavable and/or compatible with STERRAD/EtO/high-level disinfection per accessory instructions.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-18 (endoscopic equipment), ISO 8600 series (endoscopes), CE MDR, FDA or equivalent, ISO 13485.

## 4.3 Standard Accessories

- Rigid 0° ENT Telescope, Ø 4 mm: Stainless steel; autoclavable; working length suitable for sinus/nasal endoscopy (typically 175—180 mm); wide-angle optics; compatible with the specified light guide and camera coupler.

- Rigid 30° ENT Telescope, Ø 4 mm: Specifications as above with 30° oblique view.

- HD Medical-Grade Display Monitor (≥ 24 inch): Specifications as above.

- HD Camera Head with Integrated Coupler: 1-chip or 3-chip HD/Full-HD; ergonomic housing; sealed for reprocessing; C-mount or proprietary coupler compatible with the telescopes; manual and automatic controls.

- HD Video Processor + LED Light Source Console: Combined or separate units; DSP with edge enhancement, noise reduction; digital and analog outputs; still and video recording; USB/SD/network storage.

- Fiber-Optic Light Guide Cable, 2.5—3 m: High-transmission optical fiber bundle; heat-resistant sheath; universal adapters at both ends compatible with the light source and telescopes.

- Heavy-Duty Mobile Equipment Tower with Lockable Casters: Medical-grade cart with isolation transformer, cable management, and lockable antistatic wheels.

- Watertight Protective Cleaning/Immersion Caps: For camera and light-guide connectors during reprocessing; leak-proof.

- Perforated Stainless Steel Sterilization Tray: For telescope and accessory storage and steam/EtO/plasma sterilization; silicone inserts to protect optics.

## 4.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz supply from UPS-protected medical socket; earth continuity \< 0.2 Ω. OT/OPD floor space adequate for mobile tower; adjustable monitor viewing distance. Environmental temperature 15—30 °C; RH 30—75 %. Network access for image/video archiving configurable.

## 4.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, IEC 62353 Electrical Safety Report, Reprocessing instructions per accessory, PM Schedule, Installation & Commissioning Report.

## 4.6 Training

Clinical/application training for ENT surgeons and OT nurses; reprocessing training for CSSD staff; biomedical engineering training on PM and troubleshooting.

## 4.7 Warranty

Minimum two (2) years comprehensive on-site warranty including labour, spare parts, PM, breakdown maintenance, software updates, and technical support.

# 5. Flexible Nasopharyngoscope

Quantity: 2

## 5.1 Clinical & Functional Specifications

Flexible fiber-optic (or chip-on-tip) nasopharyngoscope intended for examination of nasal cavity, nasopharynx, oropharynx, hypopharynx, and larynx in adult and pediatric patients in OPD and bedside settings.

## 5.2 Engineering Technical Specifications

- Type: Flexible ENT scope, fiber-optic image bundle or distal CMOS chip-on-tip; compatible with the ENT tower's light source and camera or with dedicated portable LED handle.

- Insertion Tube Outer Diameter: Approximately 3.0—3.7 mm.

- Working Length: Approximately 300—350 mm.

- Field of View: Minimum 85°.

- Depth of Field: Approximately 3—50 mm.

- Angulation: Two-way; up minimum 130°, down minimum 130°.

- Image Resolution: Minimum 10,000 image pixels (fiber-optic) or 400 × 400 pixels distal chip; or equivalent.

- Reprocessing: Compatible with high-level disinfection and/or EtO/STERRAD per manufacturer instructions; leak-tested before every reprocessing cycle.

- Applicable Standards: IEC 60601-1, IEC 60601-2-18, ISO 8600 series, CE MDR, FDA or equivalent.

## 5.3 Standard Accessories

- Flexible Nasopharyngoscope Insertion Tube: Two-way angulation; high-resolution optics; leak-tight construction.

- Fiber-Optic Light Guide Adapter Plugs: Universal adapters compatible with common light sources.

- Heavy-Duty Fiber-Optic Light Guide Cable: High transmission; heat-resistant sheath; adequate length.

- Handheld Immersion Leak Tester with Pressure Gauge: Manual pump with pressure gauge; adapters for the scope.

- Watertight Cleaning/Sterilization Cap Set: Protects electrical/optical connectors during reprocessing.

- Rigid Padded Transport/Storage Case: Impact-resistant with foam inserts.

## 5.4 Installation Requirements

Access to compatible light source and camera or dedicated portable LED handle. CSSD/reprocessing room compliant with local infection-control guidelines.

## 5.5 Documentation

User Manual, Service Manual, Parts Catalogue, Reprocessing instructions, Leak-test procedure, PM Schedule.

## 5.6 Training

Clinical training for ENT physicians; reprocessing and leak-testing training for CSSD/OPD staff.

## 5.7 Warranty

Minimum two (2) years comprehensive on-site warranty.

# 6. FESS Instrument Set

Quantity: 1

## 6.1 Clinical & Functional Specifications

Complete set of surgical hand instruments for Functional Endoscopic Sinus Surgery, suitable for adult and pediatric procedures.

## 6.2 Engineering Technical Specifications

- Material: Surgical stainless steel conforming to ISO 7153-1 or equivalent; passivated; non-magnetic where clinically required.

- Finish: Satin or matte anti-glare finish suitable for endoscopic visualization.

- Reprocessing: Fully autoclavable per ISO 17665; compatible with common washer-disinfectors.

- Marking: Laser-etched manufacturer mark, catalogue number, and material designation.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 6.3 Standard Accessories

- Blakesley Nasal Forceps (straight, upturned 45°, upturned 90°): Stainless steel; multiple sizes; ergonomic ring handles; smooth ratchet action.

- Through-Cutting Nasal Forceps (various sizes): Sharp precision cutting jaws; multiple angulations.

- Backbiting Sinus Forceps (Antrum Punch): Reverse-angle cutting for maxillary sinus ostium.

- Sickle Knife & Sinus Elevators: Sharp, precision-ground; ergonomic handle.

- Maxillary Sinus Seekers & Ball-Tip Probes: Multiple angulations and tip configurations.

- Nasal Suction Cannulas (Frazier style, various Fr sizes): With thumb cut-off; smooth internal lumen.

- Hajek-Kofler Sphenoid Punch Forceps: Precision cutting jaws for sphenoidotomy.

- Perforated Stainless-Steel Sterilization Tray with Silicone Mats: Organized instrument holders; secure retention during transport and sterilization.

## 6.4 Installation Requirements

CSSD compatibility; storage cabinet.

## 6.5 Documentation

Instrument list with catalogue references, material certificate, reprocessing instructions.

## 6.6 Training

Reprocessing training for CSSD staff.

## 6.7 Warranty

Minimum two (2) years against manufacturing defects.

# 7. Mastoidectomy Instrument Set

Quantity: 1

## 7.1 Clinical & Functional Specifications

Complete surgical instrument set for cortical, canal-wall-up, and canal-wall-down mastoidectomy procedures.

## 7.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1; non-magnetic where applicable.

- Finish: Satin/matte anti-glare.

- Reprocessing: Fully autoclavable per ISO 17665.

- Marking: Laser-etched catalogue and material identification.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 7.3 Standard Accessories

- Lempert Mastoid Specula and Self-Retaining Retractors: Multiple sizes; secure ratchet locking.

- Jansen-Middleton Bone-Cutting Forceps: Heavy-duty cutting jaws; ergonomic handle.

- Mastoid Curettes (various sizes): Sharp cutting edges; assorted cup sizes.

- Periosteal Elevators (Lempert and Freer style): Double-ended; sharp and blunt edges.

- Bone Curettes and Volkmann Rongeurs: Multiple sizes and angulations.

- Micro-Suction Tubes with Finger Cut-Off Adapters: Various Fr sizes; smooth internal lumen.

- Stainless-Steel Mallet and Osteotome/Chisel Set: Multiple widths; ergonomic handle.

- Heavy-Duty Perforated Stainless-Steel Sterilization Tray: With silicone inserts.

## 7.4 Installation Requirements

CSSD compatibility; secure storage.

## 7.5 Documentation

Instrument list, material certificate, reprocessing instructions.

## 7.6 Training

Reprocessing training for CSSD staff.

## 7.7 Warranty

Minimum two (2) years against manufacturing defects.

# 8. Tympanoplasty Instrument Set

Quantity: 1

## 8.1 Clinical & Functional Specifications

Complete micro-instrument set for type I—V tympanoplasty procedures under operating microscope.

## 8.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1; non-magnetic.

- Finish: Satin/matte anti-glare for microscopic work.

- Precision: Micro-tip tolerances suitable for middle-ear microsurgery.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 8.3 Standard Accessories

- Wullstein Micro-Scissors (straight, curved right, curved left): Sharp micro-blades.

- Wullstein Micro-Alligator Forceps (straight, angled): Fine grasping jaws.

- Micro-Needles and Micro-Hooks (45°, 90°, oval): Assorted tip configurations.

- Micro-Graft Press Block Plates: Metal press for tragal/temporalis fascia graft preparation.

- Plester Suction Tubes and Micro-Tips (various gauges): Fine lumens; ergonomic.

- Reusable Self-Retaining Tympanoplasty Speculum Retractors: Multiple sizes.

- Perforated Stainless-Steel Micro-Instrument Sterilization Box with Silicone Inserts: Individual instrument retention.

## 8.4 Installation Requirements

CSSD compatibility.

## 8.5 Documentation

Instrument list, material certificate, reprocessing instructions.

## 8.6 Training

Reprocessing training.

## 8.7 Warranty

Minimum two (2) years.

# 9. Stapedectomy / Stapedotomy Instrument Set

Quantity: 1

## 9.1 Clinical & Functional Specifications

Ultra-precision micro-instrument set for stapes surgery under operating microscope.

## 9.2 Engineering Technical Specifications

- Material: High-grade surgical stainless steel per ISO 7153-1; non-magnetic where clinically required (near ossicular prostheses).

- Finish: Matte anti-glare.

- Tip Tolerance: Micron-level for footplate work.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 9.3 Standard Accessories

- Micro-Alligator Forceps with Extra-Fine Jaws: Precision grasping.

- Micro-Scissors with Ultra-Fine Blades: Sharp; controllable.

- Stapes Measuring Rods/Calipers (various lengths): Calibrated for prosthesis sizing.

- Stapes Hooks (0.2—0.4 mm tips, 45° and 90°): Assorted for footplate manipulation.

- Perforation Needles and Footplate Picks (straight, angled): Assorted geometries.

- Teflon/Stainless-Steel Piston Crimping Forceps: Controlled crimp; ergonomic.

- Dedicated Ultra-Fine Perforated Micro-Sterilization Case: Individual holders.

## 9.4 Installation Requirements

CSSD compatibility; secure storage.

## 9.5 Documentation

Instrument list, material certificate, reprocessing instructions.

## 9.6 Training

Reprocessing training.

## 9.7 Warranty

Minimum two (2) years.

# 10. ENT Operating Microscope

Quantity: 1

## 10.1 Clinical & Functional Specifications

High-precision surgical operating microscope providing stereoscopic, coaxially illuminated magnified visualization for otologic, rhinologic, and laryngeal microsurgery. Shall support smooth zoom, ergonomic positioning, integrated HD documentation, and configurable co-observation.

## 10.2 Engineering Technical Specifications

- Optical System: Apochromatic stereoscopic optics; coaxial illumination.

- Magnification: Continuous zoom or 5-step manual; total range approximately 4× to 25× with standard eyepieces.

- Objective Lens: Variable working distance approximately 200—400 mm (manual or motorized).

- Eyepieces: Widefield 10× or 12.5×; diopter adjustment ±5 D with locking mechanism; wide interpupillary adjustment (approximately 55—75 mm).

- Binocular Tube: Tilting 0°—180°; ergonomic.

- Illumination: High-intensity LED (or xenon) coaxial system; color temperature approximately 5500—6500 K; adjustable intensity; backup lamp module or dual-LED for uninterrupted operation.

- Documentation: HD/Full-HD video camera attachment; medical-grade monitor minimum 24 inch; still image and video recording; DICOM export configurable.

- Stand: Floor stand with mechanical/magnetic braking; balanced articulated arm; lockable antistatic casters; stability class suitable for microsurgery.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, ISO 10936-1 (surgical microscopes), CE MDR, FDA or equivalent, ISO 13485.

## 10.3 Standard Accessories

- Floor Stand with Mechanical/Magnetic Braking and Lockable Casters: Stable base; smooth motion; lockable brakes.

- Binocular Tube 0°—180° Variable Tilt: Ergonomic tilt for surgeon comfort.

- Widefield Eyepieces 10× or 12.5× with Diopter Locks: Individual diopter adjustment and locking.

- Manual or Motorized Zoom and Focus System with Variable Working Distance: Smooth continuous control.

- Integrated LED or Xenon Coaxial Illumination with Backup Lamp: Redundant illumination; adjustable intensity.

- HD Video Camera Attachment with Medical-Grade Monitor: Minimum Full-HD; minimum 24-inch monitor.

- Autoclavable Protective Caps/Covers for Adjustment Knobs: For sterile intraoperative handling.

## 10.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz supply from UPS-protected medical socket; earth continuity \< 0.2 Ω. OT floor space and vibration-free environment; adequate ceiling clearance. Environmental 15—30 °C; RH 30—75 %.

## 10.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, IEC 62353 Report, PM Schedule, Installation & Commissioning Report.

## 10.6 Training

Clinical training for ENT surgeons; biomedical engineering training on PM, lamp/LED replacement, balancing, and troubleshooting.

## 10.7 Warranty

Minimum two (2) years comprehensive on-site warranty.

# 11. Battery-Operated ENT Micro-Drill

Quantity: 1

## 11.1 Clinical & Functional Specifications

High-speed, battery-powered surgical drill system for mastoid, temporal bone, and sinus procedures.

## 11.2 Engineering Technical Specifications

- Motor: Brushless DC micro-motor; high torque; low vibration and noise.

- Speed Range: Adjustable from minimum 1,000 rpm to maximum 80,000 rpm or better via foot pedal.

- Direction: Forward/reverse selectable.

- Battery: Rechargeable lithium-ion; minimum 2 hours continuous surgical use per charge; charging time maximum 3 hours.

- Sterilization: Handpiece, attachments, and battery housing autoclavable per ISO 17665 (as applicable).

- Irrigation: Integrated peristaltic irrigation pump with adjustable flow rate.

- Foot Pedal: Waterproof, IP rating minimum IPX8; multi-function.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-2 (where applicable), CE MDR, FDA or equivalent, ISO 13485.

## 11.3 Standard Accessories

- High-Torque Brushless Micro-Motor Handpiece: Ergonomic; autoclavable; low vibration.

- Straight Micro-Drill Attachment Sleeve: Precision-machined; autoclavable.

- Angled Micro-Drill Attachment Sleeve: Angled for anatomical access.

- Rechargeable Lithium-Ion Battery Packs (2): Minimum 2 hours runtime each; interchangeable.

- Universal Multi-Slot Automatic Battery Charger: Simultaneous charging; charge status indicators.

- Integrated Irrigation Pump Module with Tube Bracket: Adjustable flow.

- Waterproof Multi-Function Foot Pedal: IPX8; ergonomic.

- Autoclavable Sterilization/Storage Tray: With instrument retention.

## 11.4 Installation Requirements

100—240 VAC, 50/60 Hz for charger; OT compatible.

## 11.5 Documentation

User Manual, Service Manual, Parts Catalogue, IEC 62353 Report, PM Schedule.

## 11.6 Training

Clinical training for surgeons and OT staff; biomedical training on battery care and PM.

## 11.7 Warranty

Minimum two (2) years comprehensive; batteries covered minimum one (1) year.

# 12. ENT Examination Unit

Quantity: 1

## 12.1 Clinical & Functional Specifications

Integrated ENT OPD workstation providing suction, compressed air, illumination, instrument warming, ear irrigation, scope disinfection, and motorized patient chair — all in a single ergonomic console.

## 12.2 Engineering Technical Specifications

- Console: Anti-bacterial work surfaces; corrosion-resistant construction.

- Suction: Integrated high-vacuum pump; vacuum minimum −70 kPa; flow minimum 20 L/min; secretion jar with overflow protection.

- Compressed Air: Oil-free compressor; pressure minimum 2 bar; atomizing spray bottles.

- Illumination: LED fiber-optic light source; color temperature approximately 5500—6500 K; adjustable intensity; scope holder hooks.

- Patient Chair: Motorized; height, backrest, and headrest adjustable; foot-switch operated; auto-return; safe working load minimum 135 kg.

- Water System: Ear-irrigation module with temperature-controlled water (approximately 37 °C ± 2 °C); flow control.

- Scope Disinfection Holders: Integrated tubes/holders for rigid/flexible scopes.

- Power Supply: 100—240 VAC, 50/60 Hz; internal isolation transformer.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-52 (where applicable to chair), CE MDR, ISO 13485.

## 12.3 Standard Accessories

- Main Console Workstation with Anti-Bacterial Surfaces: Ergonomic; smooth cleanable finish.

- Integrated High-Vacuum Suction with Secretion Jar: Autoclavable jar; overflow protection.

- Integrated Compressed Air System with Atomizing Spray Bottles (2): Adjustable pressure.

- Integrated LED Fiber-Optic Light Source with Scope Holder Hooks: Multiple scope holders.

- Motorized Patient Chair with Foot Switch and Auto-Return: Multi-axis adjustment.

- Integrated Rigid/Flexible Scope Disinfection Tubes/Holders: For quick turnaround.

- Built-In Ear-Irrigation Module with Temperature-Controlled Water Line: Precise temperature.

- Medical-Grade Waste Bin and Pull-Out Storage Drawers: Ergonomic layout.

## 12.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz supply from UPS/medical socket; earth continuity \< 0.2 Ω; potable water supply and drainage; adequate OPD floor space; ventilation for compressor heat.

## 12.5 Documentation

User Manual, Service Manual, Parts Catalogue, IEC 62353 Report, PM Schedule, Installation & Commissioning Report.

## 12.6 Training

Clinical training for ENT physicians and nurses; biomedical engineering training on PM (compressor, suction, water system) and troubleshooting.

## 12.7 Warranty

Minimum two (2) years comprehensive on-site warranty.

# 13. ENT Diagnostic and Minor-Procedure Set

Quantity: 7

## 13.1 Clinical & Functional Specifications

Standard set of ENT hand instruments for OPD diagnostic examination and minor procedures.

## 13.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1; non-magnetic where applicable.

- Finish: Satin/mirror as appropriate.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 13.3 Standard Accessories

- Hartmann Ear Specula (set of 3 sizes): Assorted diameters for adult/pediatric.

- Killian Nasal Specula (various lengths): Multiple blade lengths.

- Lucae Ear Forceps and Wilde Alligator Forceps: For foreign body removal.

- Jobson-Horne Ear Probes/Cotton Applicators: Double-ended.

- Indirect Laryngoscopy Mirrors (various sizes) with Universal Handle: Multiple diameters; heatable configurable.

- Politzer Air Bag with Universal Nose Pieces: For Eustachian tube inflation.

- Stainless-Steel Tongue Depressors and Wax Hooks/Curettes: Assorted geometries.

- Perforated Stainless-Steel Instrument Storage Case: For each set.

## 13.4 Installation Requirements

CSSD compatibility; OPD storage.

## 13.5 Documentation

Instrument list per set; material certificate; reprocessing instructions.

## 13.6 Training

Reprocessing training.

## 13.7 Warranty

Minimum two (2) years against manufacturing defects.

# 14. ENT Microscope for OPD

Quantity: 1

## 14.1 Clinical & Functional Specifications

Compact operating microscope for OPD ear examinations, cerumen removal, foreign-body extraction, and minor otological procedures.

## 14.2 Engineering Technical Specifications

- Optical System: Stereoscopic; apochromatic or equivalent.

- Magnification: 3-step or 5-step manual changer; total range approximately 4×—20×.

- Objective Lens: Fixed working distance 250 mm or 300 mm.

- Eyepieces: Widefield 10×; diopter adjustable.

- Illumination: Integrated LED; adjustable intensity; color temperature approximately 5500—6500 K.

- Stand: Compact floor stand with smooth-rolling lockable casters or wall/table mount configurable.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, ISO 10936-1, CE MDR, ISO 13485.

## 14.3 Standard Accessories

- Compact Floor Stand with Lockable Casters: Smooth rolling; stable.

- Fixed or Inclinable Binocular Head with 10× Eyepieces: Ergonomic viewing.

- 3-Step or 5-Step Manual Magnification Changer: Positive detents.

- Integrated LED Light Source with Brightness Control: Long lamp life.

- Fixed Objective Lens (f = 250 mm or 300 mm): Suitable working distance.

- Autoclavable Detachable Handle Covers: For sterile handling.

## 14.4 Installation Requirements

100—240 VAC, 50/60 Hz supply; OPD floor space; adequate lighting environment.

## 14.5 Documentation

User Manual, Service Manual, Parts Catalogue, IEC 62353 Report, PM Schedule.

## 14.6 Training

Clinical training; biomedical training on PM.

## 14.7 Warranty

Minimum two (2) years comprehensive on-site warranty.

# 15. Microlaryngeal Instruments Set

Quantity: 1

## 15.1 Clinical & Functional Specifications

Precision micro-instruments for microlaryngoscopy and phonosurgery under operating microscope.

## 15.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1; non-magnetic where required.

- Working Length: Extra-long shaft suitable for microlaryngoscopy (typically 220—250 mm).

- Finish: Matte anti-glare.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 15.3 Standard Accessories

- Kleinsasser Microlaryngeal Scissors (straight, curved right, curved left): Micro-blades.

- Kleinsasser Microlaryngeal Grasping Forceps (straight, upturned, angled): Precision jaws.

- Microlaryngeal Cup Forceps (straight, angled): For biopsy.

- Microlaryngeal Probes, Hooks, and Cotton Carriers: Assorted geometries.

- Suction Cannulas with Integrated Thumb Regulator (various lengths): Smooth lumens.

- Rigid Protective Micro-Instrument Sterilization Case: With individual holders.

## 15.4 Installation Requirements

CSSD compatibility.

## 15.5 Documentation

Instrument list, material certificate, reprocessing instructions.

## 15.6 Training

Reprocessing training.

## 15.7 Warranty

Minimum two (2) years.

# 16. Telescope for Microlaryngoscopy

Quantity: 1

## 16.1 Clinical & Functional Specifications

Rigid angled telescope for diagnostic and operative microlaryngoscopy providing wide-angle high-resolution visualization of the larynx and vocal folds.

## 16.2 Engineering Technical Specifications

- Type: Rigid, autoclavable, stainless steel.

- Angle: 70° or 90°.

- Working Length: Extra-long, typically 170—200 mm.

- Outer Diameter: Approximately 5—10 mm (selectable/configurable).

- Field of View: Minimum 80°.

- Optical Quality: Wide-angle high-resolution; compatible with HD/Full-HD camera couplers.

- Reprocessing: Autoclavable per manufacturer instructions.

- Applicable Standards: IEC 60601-2-18, ISO 8600 series, CE MDR, ISO 13485.

## 16.3 Standard Accessories

- Rigid 70° or 90° Laryngoscope Telescope (Extra-Long): As specified above.

- Fiber-Optic Light Guide Cable Adapter Plugs: Universal.

- Watertight Protective Cleaning Cap Set: For connectors.

- Rigid Perforated Sterilization/Containment Tray: With silicone inserts.

## 16.4 Installation Requirements

Compatible light source and camera required.

## 16.5 Documentation

User Manual, Reprocessing instructions.

## 16.6 Training

Clinical training; reprocessing training.

## 16.7 Warranty

Minimum two (2) years.

# 17. Rigid Oesophagoscope Set

Quantity: 1

## 17.1 Clinical & Functional Specifications

Rigid oesophagoscopy set for diagnostic examination, foreign-body removal, and biopsy in adult and pediatric patients.

## 17.2 Engineering Technical Specifications

- Type: Rigid, stainless steel; autoclavable.

- Configurations: Adult and pediatric sizes; oval or round cross-section.

- Lengths: Adult approximately 45—53 cm; pediatric approximately 25—30 cm.

- Illumination: Fiber-optic light carrier compatible with standard light sources.

- Applicable Standards: IEC 60601-2-18, ISO 8600 series, CE MDR, ISO 13485.

## 17.3 Standard Accessories

- Rigid Oesophagoscope Tubes (Adult and Pediatric): Multiple sizes.

- Fiber-Optic Light Carrier Inserts/Bars: Compatible with tubes.

- Heavy-Duty Fiber-Optic Light Guide Cable: Adequate length.

- Esophageal Grasping and Biopsy Forceps (Extra-Long): Assorted jaws.

- Rigid Esophageal Suction Tubes (Various Lengths): Smooth lumens.

- Insulated Tooth Protector Guards: Cushioned; secure fit.

- Long Perforated Sterilization/Containment Box: With retention inserts.

## 17.4 Installation Requirements

Compatible light source required; CSSD compatibility.

## 17.5 Documentation

User Manual, Reprocessing instructions.

## 17.6 Training

Clinical training; reprocessing training.

## 17.7 Warranty

Minimum two (2) years.

# 18. Tonsillectomy and Adenoidectomy Instrument Set

Quantity: 2

## 18.1 Clinical & Functional Specifications

Complete instrument set for conventional tonsillectomy and adenoidectomy.

## 18.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1.

- Finish: Satin/mirror.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 18.3 Standard Accessories

- Boyle-Davis Mouth Gag Frame with Interchangeable Tongue Blades (set of 4 sizes): Adjustable frame; assorted blade sizes for adult/pediatric.

- Negus Tonsil Artery Forceps (straight, curved): Ratchet locking.

- Denis-Browne Tonsil Holding Forceps: Secure grip.

- Beckmann Adenoid Curettes (various sizes): Sharp cutting edges.

- Tonsil Snare Assembly with Replacement Wire Pack: Adjustable; wire replacements included.

- Tonsil Dissector and Pillar Retractor: Ergonomic handles.

- Heavy-Duty Perforated Stainless-Steel Sterilization Tray: With silicone inserts.

## 18.4 Installation Requirements

CSSD compatibility.

## 18.5 Documentation

Instrument list per set, material certificate, reprocessing instructions.

## 18.6 Training

Reprocessing training.

## 18.7 Warranty

Minimum two (2) years.

# 19. Septoplasty Instrument Set

Quantity: 2

## 19.1 Clinical & Functional Specifications

Complete instrument set for septoplasty and submucous resection procedures.

## 19.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1.

- Finish: Satin/mirror; matte anti-glare where required.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 19.3 Standard Accessories

- Killian Nasal Specula (long and short blades): Multiple sizes.

- Freer Double-Ended Periosteal Elevator: Sharp and blunt ends.

- Ballenger Swivel Knife: Sharp swivel blade.

- Cottle Nasal Cartilage Speculum and Septal Speculum: Assorted.

- Jansen-Middleton Septal Punch Forceps: Precision cutting jaws.

- Asch Septal Straightening Forceps: Strong grip.

- Stainless-Steel Surgical Mallet and Nasal Chisels/Gouges: Multiple widths.

- Perforated Stainless-Steel Instrument Disinfection Tray: With silicone inserts.

## 19.4 Installation Requirements

CSSD compatibility.

## 19.5 Documentation

Instrument list per set, material certificate, reprocessing instructions.

## 19.6 Training

Reprocessing training.

## 19.7 Warranty

Minimum two (2) years.

# 20. Rhinoplasty Instrument Set

Quantity: 1

## 20.1 Clinical & Functional Specifications

Complete instrument set for functional and aesthetic rhinoplasty procedures.

## 20.2 Engineering Technical Specifications

- Material: Surgical stainless steel per ISO 7153-1; non-magnetic where required.

- Finish: Matte anti-glare for fine work.

- Reprocessing: Autoclavable per ISO 17665.

- Applicable Standards: ISO 7153-1, ISO 13485, CE MDR.

## 20.3 Standard Accessories

- Aufricht Nasal Retractors (with and without fiber-optic light guide): With light guide port option.

- Joseph Nasal Saws (left and right) and Rasps (various grit sizes): Sharp edges; assorted grits.

- Cinelli Osteotomes and Rubin Septal Morselizers: Precision cutting.

- Fomon Dorsal Scissors and Fine-Tipped Plastic Surgery Scissors: Sharp micro-blades.

- Adson-Brown Delicate Tissue Forceps: Fine teeth.

- Cottle Cartilage Crushing Box with Mallet: Solid base; smooth crushing surface.

- Heavy-Duty Perforated Stainless-Steel Containment Tray with Silicone Padding: Instrument retention.

## 20.4 Installation Requirements

CSSD compatibility.

## 20.5 Documentation

Instrument list, material certificate, reprocessing instructions.

## 20.6 Training

Reprocessing training.

## 20.7 Warranty

Minimum two (2) years.

# 21. Rigid Bronchoscope Set

Quantity: 1

## 21.1 Clinical & Functional Specifications

Rigid bronchoscopy set for adult and pediatric diagnostic and therapeutic bronchoscopy, including foreign-body removal and biopsy, with support for jet or conventional ventilation.

## 21.2 Engineering Technical Specifications

- Type: Rigid stainless-steel bronchoscope tubes; autoclavable.

- Sizes: Adult and pediatric range; outer diameter approximately 3.5—9 mm; corresponding lengths.

- Illumination: Prismatic light deflector or fiber-optic light carrier; compatible with standard light sources.

- Ventilation Ports: Side ports for ventilation; sealing caps for artificial ventilation.

- Applicable Standards: IEC 60601-2-18, ISO 8600 series, CE MDR, ISO 13485.

## 21.3 Standard Accessories

- Rigid Bronchoscope Tubes (various diameters/lengths, adult and pediatric): Full size range.

- Prismatic Light Deflector or Fiber-Optic Light Carrier Inserts: Compatible with tubes.

- Optical Biopsy Forceps and Optical Foreign-Body Grasping Forceps: Integrated telescope compatibility; assorted jaws.

- Universal Injection Cannula and Rigid Suction Tubes: Various lengths.

- Glass Slide Windows/Caps for Ventilation Attachment Closure: Sealing caps.

- Heavy-Duty Fiber-Optic Light Transmission Cable: Adequate length.

- Long Perforated Sterilization/Containment Tray: With silicone retention inserts.

## 21.4 Installation Requirements

Compatible light source and anesthesia/ventilation interface required; CSSD compatibility.

## 21.5 Documentation

User Manual, Reprocessing instructions, Material certificate.

## 21.6 Training

Clinical training for ENT/anesthesia teams; reprocessing training.

## 21.7 Warranty

Minimum two (2) years comprehensive.