HOSPITAL TECHNICAL TENDER SPECIFICATIONS

Emergency, Resuscitation & Patient Handling Equipment

Table of Contents

1. BIPHASIC DEFIBRILLATOR WITH EXTERNAL PADDLES + AED MODE + EXTERNAL PACING	4

1.1  Clinical & Functional Specifications	4

1.2  Engineering Technical Specifications	4

1.3  Standard Accessories	7

1.4  Installation Requirements	8

1.5  Documentation	9

1.6  Training	9

1.7  Warranty	9

2. DEFIBRILLATOR WITH INTERNAL PADDLES	10

2.1  Clinical & Functional Specifications	10

2.2  Engineering Technical Specifications	10

2.3  Standard Accessories	11

2.4  Installation Requirements	13

2.5  Documentation	13

2.6  Training	13

2.7  Warranty	13

3. AUTOMATIC EXTERNAL DEFIBRILLATOR (AED)	14

3.1  Clinical & Functional Specifications	14

3.2  Engineering Technical Specifications	14

3.3  Standard Accessories	15

3.4  Installation Requirements	16

3.5  Documentation	16

3.6  Training	17

3.7  Warranty	17

4. CRASH CART	18

4.1  Clinical & Functional Specifications	18

4.2  Engineering Technical Specifications	18

4.3  Standard Accessories	19

4.4  Installation Requirements	21

4.5  Documentation	21

4.6  Training	21

4.7  Warranty	21

5. RESUSCITATION BED / TROLLEY WITH X-RAY TRANSLUCENT TOP	22

5.1  Clinical & Functional Specifications	22

5.2  Engineering Technical Specifications	22

5.3  Standard Accessories	23

5.4  Installation Requirements	24

5.5  Documentation	24

5.6  Training	25

5.7  Warranty	25

6. PATIENT TRANSFER STRETCHER	26

6.1  Clinical & Functional Specifications	26

6.2  Engineering Technical Specifications	26

6.3  Standard Accessories	27

6.4  Installation Requirements	28

6.5  Documentation	28

6.6  Training	28

6.7  Warranty	28

7. TRAUMA TROLLEY	29

7.1  Clinical & Functional Specifications	29

7.2  Engineering Technical Specifications	29

7.3  Standard Accessories	30

7.4  Installation Requirements	31

7.5  Documentation	31

7.6  Training	31

7.7  Warranty	31

8. SPINAL BOARDS, CERVICAL COLLARS & SPLINTS (IMMOBILISATION SET)	32

8.1  Clinical & Functional Specifications	32

8.2  Engineering Technical Specifications	32

8.3  Standard Accessories	33

8.4  Installation Requirements	34

8.5  Documentation	34

8.6  Training	34

8.7  Warranty	34

9. EYE-WASH STATION	35

9.1  Clinical & Functional Specifications	35

9.2  Engineering Technical Specifications	35

9.3  Standard Accessories	36

9.4  Installation Requirements	37

9.5  Documentation	38

9.6  Training	38

9.7  Warranty	38

# 1. BIPHASIC DEFIBRILLATOR WITH EXTERNAL PADDLES + AED MODE + EXTERNAL PACING

Quantity: 12

## 1.1  Clinical & Functional Specifications

The equipment shall be a portable, microprocessor-controlled, biphasic defibrillator-monitor intended for use in emergency departments, intensive care units, coronary care units, operating theatres, cardiac catheterization laboratories, and during intra-hospital patient transport. It shall be suitable for adult, pediatric, and neonatal patient populations.

The unit shall provide the following clinical operating modes, selectable through the user interface:

- Manual Defibrillation Mode --- for use by trained clinicians with user-selectable energy levels.

- Synchronized Cardioversion Mode --- for treatment of atrial and ventricular tachyarrhythmias.

- Advisory / AED Mode --- with automatic ECG rhythm analysis, shock advisory, and voice-guided operator prompts suitable for use by basic life support providers.

- Non-Invasive External Transcutaneous Pacing Mode --- with demand and fixed (asynchronous) pacing.

- ECG Monitoring Mode --- through paddles, pads, and patient cable.

- Multi-Parameter Patient Monitoring Mode --- including SpO₂ and NIBP (where integrated).

The equipment shall support defibrillation via reusable external paddles, disposable multi-function pads, and shall accept pediatric paddle adapters or dedicated pediatric paddles. The AED mode shall incorporate a validated rhythm-analysis algorithm capable of detecting shockable rhythms (ventricular fibrillation and pulseless ventricular tachycardia) with high sensitivity and specificity in accordance with international resuscitation guidelines (AHA/ERC/ILCOR).

The device shall provide automatic energy escalation protocols configurable by the biomedical department, event recording with time-stamped audit trail, and data export capability for post-event review and quality assurance.

## 1.2  Engineering Technical Specifications

Waveform

- Biphasic truncated exponential, biphasic rectilinear, or equivalent clinically validated biphasic waveform

- Automatic impedance compensation across a patient impedance range of at least 25--175 ohms

Energy Output (Defibrillation)

- Adult selectable energy range: minimum 2 to 200 Joules or wider

- Pediatric selectable energy range: minimum 1 to 50 Joules or wider

- Selectable energy steps suitable for progressive escalation protocols

- Charge time to maximum energy: not exceeding 10 seconds from fully charged battery

- Energy delivery accuracy: within ±15% or better across the operating range

External Pacing

- Pacing modes: Demand (synchronous) and Fixed (asynchronous), user-selectable

- Pacing rate range: minimum 30 to 180 pulses per minute, adjustable

- Pacing current output: minimum 10 to 200 mA, adjustable in defined steps

- Pulse width: fixed or selectable, typically 20--40 ms

ECG Monitoring

- Lead configuration: 3-lead and 5-lead selectable via patient cable; monitoring through paddles and pads

- ECG bandwidth (diagnostic): minimum 0.05--150 Hz

- ECG bandwidth (monitoring): minimum 0.5--40 Hz

- Heart-rate measurement range: minimum 20 to 300 bpm

- Heart-rate accuracy: within ±2 bpm or ±2%, whichever is greater

- Defibrillation-protected inputs conforming to IEC 60601-2-27

- Sampling rate: minimum 500 samples/second/channel

SpO₂ (where integrated)

- Measurement range: 0--100%

- Accuracy: within ±2% in the 70--100% range for adults/pediatrics; ±3% for neonates

- Pulse rate range: minimum 30--250 bpm

- Motion-tolerant, low-perfusion capable technology

NIBP (where integrated)

- Measurement technique: oscillometric

- Systolic range: minimum 30--260 mmHg

- Diastolic range: minimum 15--220 mmHg

- MAP range: minimum 20--235 mmHg

- Accuracy: within ±3 mmHg or ±2%

- Adult, pediatric, and neonatal modes

Display

- Color TFT LCD or equivalent display technology

- Minimum diagonal viewing area: 5 inches

- Minimum resolution: 320 × 240 pixels or better

- Simultaneous display of at least 2 waveforms and all active numeric parameters

- Sunlight-readable, wide viewing angle

Data Management & Event Recording

- Internal event memory: minimum 4 hours of continuous ECG and event data

- Time-stamped storage of shocks, alarms, pacing events, and monitoring data

- Data export via USB, SD card, Ethernet, or wireless (Wi-Fi/Bluetooth) --- configurable

- Report generation in PDF or equivalent portable format

Printer / Recorder

- Integrated thermal strip recorder (where specified in BOQ)

- Paper width: minimum 50 mm

- Print speed: selectable, typically 12.5, 25, and 50 mm/s

- Automatic printing on shock delivery and alarm events

Alarms

- Multi-level audible and visual alarms with adjustable priority

- User-configurable upper/lower limits for all monitored parameters

- Alarm silence, pause, and history log functions

User Interface

- Dedicated hard keys for critical functions (Charge, Shock, Energy Select, Pacer)

- Rotary selector or membrane keypad, chemical- and fluid-resistant

- Multilingual interface capability

Power Supply

- Input: 100--240 VAC, 50/60 Hz, auto-ranging

- Integrated rechargeable battery: lithium-ion or sealed lead-acid

- Battery runtime: minimum 3 hours of monitoring, or minimum 50 full-energy discharges on a fully charged battery

- Charging time: maximum 4 hours to 90% capacity

- Battery status indicator with capacity display

- Hot-swappable battery (preferred)

Physical Characteristics

- Portable, ruggedized construction with integrated carry handle

- Weight: approximately 5--10 kg with battery

- Dimensions: compact portable configuration suitable for bedside and transport use

- Ingress protection: minimum IP44 or equivalent for splash resistance

- Drop resistance: minimum 0.75 m onto hard surface

Environmental Conditions

- Operating temperature: 0--45 °C

- Storage temperature: −20 to +60 °C

- Operating humidity: 10--95% RH, non-condensing

- Altitude: operational up to 3,000 m or better

Connectivity

- USB, Ethernet (RJ45), and optional Wi-Fi/Bluetooth

- Compatible with hospital information systems via HL7 or equivalent

Applicable Standards

- IEC 60601-1 (General safety)

- IEC 60601-1-2 (EMC)

- IEC 60601-2-4 (Defibrillators)

- IEC 60601-2-25 / 60601-2-27 (ECG monitoring)

- IEC 60601-2-49 (Multi-parameter monitors, where applicable)

- IEC 62304 (Software lifecycle)

- ISO 13485, ISO 14971

- CE MDR / FDA 510(k) clearance or equivalent

## 1.3  Standard Accessories

Adult Reusable External Paddles (1 pair)

- Anterior/apex configuration with integrated Charge and Shock controls

- Stainless steel or equivalent conductive electrode surface

- Minimum electrode area: 70 cm² adult

- Ergonomic insulated handles with cable strain relief

- Compatible with pediatric adapters (slide-over or plug-in)

Pediatric Paddle Adapters or Pediatric Paddles (1 pair)

- Minimum electrode area: 15 cm² pediatric

- Sliding, latching, or plug-in mechanism ensuring positive electrical contact

- Compatible with the supplied adult paddles or as dedicated pediatric paddles

Disposable Adult Defibrillation Pads (minimum 10 pairs)

- Pre-gelled, radiotranslucent, single-use adhesive electrodes

- Shelf life: minimum 24 months from date of supply

- Compatible with defibrillation, pacing, and ECG monitoring

- Individually sealed in moisture-proof packaging

Pediatric AED Pads (if AED pediatric mode supported)

- Pre-gelled disposable pads sized for pediatric patients (\<25 kg or \<8 years)

- Shelf life: minimum 24 months

- Individually sealed

3-Lead and 5-Lead ECG Patient Cable

- Defibrillation-protected, shielded, medical-grade cable

- IEC or AHA color coding, clearly labeled

- Snap or clip terminations, latex-free

- Cable length: minimum 2.5 m

ECG Monitoring Electrodes

- Disposable, pre-gelled, radiolucent

- Adult, pediatric, and neonatal sizes as applicable

- Shelf life: minimum 12 months

SpO₂ Sensor (if integrated)

- Reusable adult finger sensor with cable length ≥2 m

- Compatible with the host monitor

- Latex-free, biocompatible

NIBP Cuff with Hose (if integrated)

- Reusable adult cuff with matched hose

- Latex-free, wipe-cleanable bladder and sleeve

- Marking for correct sizing and artery position

Rechargeable Battery (minimum 2 units)

- Lithium-ion or sealed lead-acid, compatible with the host device

- Capacity meeting or exceeding the runtime specifications above

- Charge cycle life: minimum 300 full cycles

AC Power Cable

- Hospital-grade, appropriate country plug, minimum 2.5 m length

- Compliant with IEC 60601-1

Battery Charger (if external)

- Bench-top or wall-mount, capable of charging at least one battery simultaneously

- Charge status indication and end-of-life detection

ECG Recording Paper (if printer available)

- Thermal paper compatible with the recorder

- Minimum 10 rolls supplied

Conductive Gel for Paddles

- Water-soluble, non-irritant, electrically conductive

- Minimum 4 tubes/bottles supplied

Operator Manual

- Printed and electronic copies in English (and local language if specified)

Carrying Case (portable models)

- Rigid or semi-rigid impact-resistant case with foam inserts

- Storage compartments for accessories and cables

## 1.4  Installation Requirements

- Electrical supply: 100--240 VAC, 50/60 Hz single-phase, connected to essential/UPS-backed power circuit

- Dedicated hospital-grade earthing with earth-leakage protection

- Compatible with hospital UPS and emergency generator supply

- Ambient conditions: air-conditioned clinical environment, 18--26 °C recommended

- Wall or trolley mounting where applicable

- Network point (RJ45) for connectivity to hospital HIS/EMR (where applicable)

- Space requirement: compact bedside footprint

## 1.5  Documentation

Supplier shall provide, in English (and local language if specified):

- User/Operator Manual

- Service and Maintenance Manual with schematics

- Spare Parts Catalogue with part numbers

- Factory Test Certificate

- Electrical Safety Test Report per IEC 62353

- Calibration Certificate (traceable to national/international standard)

- Preventive Maintenance Schedule

- Installation and Commissioning Report

- Declaration of Conformity (CE / FDA / equivalent)

## 1.6  Training

Supplier shall provide, on-site, at no additional cost:

- Clinical/Application Training for medical and nursing users

- Biomedical Engineering Training on operation, calibration, and safety testing

- Preventive Maintenance Training

- Basic Troubleshooting and first-line service training

- Training documentation and attendance certificates

## 1.7  Warranty

- Minimum comprehensive warranty: 2 years from date of commissioning, or as specified by purchaser

- Warranty shall include: labour, spare parts, batteries, preventive maintenance visits (minimum 2 per year), breakdown maintenance, software updates and patches, and technical support

- Guaranteed availability of spare parts for a minimum of 10 years from date of supply

# 2. DEFIBRILLATOR WITH INTERNAL PADDLES

Quantity: 01

## 2.1  Clinical & Functional Specifications

The equipment shall be a microprocessor-controlled biphasic defibrillator designed for use in operating theatres, cardiac surgery suites, and cardiothoracic intensive care units, providing both external and open-chest (internal) defibrillation capability. It shall be suitable for adult and pediatric cardiac surgical patients.

The unit shall support:

- Manual Defibrillation with internal (open-chest) paddles

- Manual Defibrillation with external paddles

- Synchronized Cardioversion

- ECG Monitoring through paddles, pads, and patient cable

- Event recording and data export

The equipment shall provide a dedicated internal-paddle safety interlock automatically limiting the maximum deliverable energy when internal paddles are connected, in accordance with international clinical practice.

## 2.2  Engineering Technical Specifications

Waveform

- Biphasic truncated exponential, biphasic rectilinear, or equivalent

- Impedance-compensated across 25--175 ohms

Energy Output --- External Paddles/Pads

- Adult range: minimum 2--200 Joules or wider

- Pediatric range: minimum 1--50 Joules or wider

Energy Output --- Internal Paddles

- Adult range: minimum 1--50 Joules

- Pediatric range: minimum 1--30 Joules

- Automatic maximum-energy limitation when internal paddles detected

- Energy delivery accuracy: within ±15%

- Charge time to maximum energy: not exceeding 10 seconds

ECG Monitoring

- 3-lead and 5-lead configurations

- Bandwidth (diagnostic): minimum 0.05--150 Hz

- HR range: 20--300 bpm; accuracy ±2 bpm or ±2%

- Defibrillation-protected inputs per IEC 60601-2-27

Display

- Color TFT LCD or equivalent

- Minimum diagonal viewing area: 5 inches

- Minimum resolution: 320 × 240 pixels

- Simultaneous waveform and numeric display

Synchronized Cardioversion

- R-wave detection with visual sync markers

- Sync-mode retention or automatic reset as configurable

Event Recording

- Internal memory: minimum 4 hours of continuous ECG

- Time-stamped shock and event log

- Data export via USB or equivalent

Alarms

- Multi-level audible/visual alarms with configurable limits

Power Supply

- Input: 100--240 VAC, 50/60 Hz

- Rechargeable battery: minimum 2 hours monitoring or minimum 30 full-energy discharges

- Charging time: maximum 4 hours to 90%

Physical Characteristics

- Portable/trolley-mountable configuration

- Weight: approximately 6--12 kg

- Ingress protection: minimum IP44 or equivalent

Environmental

- Operating: 0--45 °C; 10--95% RH non-condensing

- Storage: −20 to +60 °C

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-2-4, 60601-2-27

- IEC 62304, ISO 13485, ISO 14971

- CE MDR / FDA 510(k) or equivalent

## 2.3  Standard Accessories

Adult Internal Paddles (1 pair)

- Autoclavable, spoon-shaped stainless steel electrodes

- Minimum electrode diameter: 65--75 mm adult

- Insulated handles with integrated discharge control

- Cable length: minimum 3 m

- Fully immersible and steam-sterilizable at 134 °C

Pediatric Internal Paddles (1 pair)

- Autoclavable pediatric spoon electrodes

- Minimum electrode diameter: 25--45 mm

- Insulated handles with discharge control

- Steam-sterilizable at 134 °C

Paddle Sterilization Tray/Container

- Stainless steel or medical-grade equivalent

- Perforated design suitable for steam sterilization

- Sized to accommodate all supplied internal paddles and cables

Paddle Cables

- Defibrillation-rated, autoclavable or reusable per manufacturer instructions

- Length: minimum 3 m

- Strain-relieved connectors

External Paddles (1 pair)

- Reusable adult paddles as per Item 1 specifications

- Integrated Charge/Shock controls

ECG Patient Cable

- 3-lead or 5-lead, defibrillation-protected, IEC/AHA coded

- Length: minimum 2.5 m

Disposable Defibrillation Pads

- Multi-function pads for defibrillation and monitoring

- Minimum 10 pairs supplied; shelf life ≥24 months

Rechargeable Battery

- Minimum 2 units; compatible with host device

- Charge cycle life: minimum 300 cycles

AC Power Cord

- Hospital-grade, minimum 2.5 m, compliant with IEC 60601-1

ECG Paper Roll (if recorder present)

- Thermal paper, minimum 10 rolls

Conductive Gel

- Water-soluble, non-irritant; minimum 4 units

User Manual

- Printed and electronic in English

Trolley or Mounting Bracket (if specified)

- Mobile trolley with lockable castors, accessory basket, and cable management

## 2.4  Installation Requirements

- 100--240 VAC, 50/60 Hz, UPS-backed essential supply

- Hospital-grade earthing

- Air-conditioned OT/ICU environment

- Space for trolley or shelf-mount configuration

- Compatible with hospital sterilization services for internal paddles

## 2.5  Documentation

Same as Item 1 including sterilization validation for internal paddles.

## 2.6  Training

Same scope as Item 1, with additional training on internal paddle handling, sterilization, and OT safety protocols.

## 2.7  Warranty

Minimum 2 years comprehensive, including internal paddle sets and spare batteries; spare-parts availability ≥10 years.

# 3. AUTOMATIC EXTERNAL DEFIBRILLATOR (AED)

Quantity: 09

## 3.1  Clinical & Functional Specifications

The equipment shall be a portable, semi-automatic external defibrillator intended for use by trained lay responders, first responders, and healthcare personnel for the treatment of sudden cardiac arrest. It shall be suitable for adult and pediatric use through appropriate electrode pad or key selection.

The AED shall provide:

- Automatic ECG rhythm analysis with shockable/non-shockable determination

- Voice and visual prompts guiding the operator through the resuscitation sequence

- CPR coaching including metronome and, where available, real-time CPR feedback

- Automatic self-testing (daily/weekly) with visible status indicator

- Event recording with time-stamped ECG and audio (where applicable)

The rhythm-analysis algorithm shall comply with the latest AHA/ERC/ILCOR resuscitation guidelines with high sensitivity and specificity for VF and pulseless VT detection.

## 3.2  Engineering Technical Specifications

Waveform

- Biphasic truncated exponential, biphasic rectilinear, or equivalent

- Impedance-compensated

Energy Output

- Adult: fixed or escalating protocol within 120--360 Joules (typical) or manufacturer\'s clinically validated protocol

- Pediatric: reduced energy protocol via pediatric pads or key

- Charge time: not exceeding 10 seconds from shock advisory to ready state

Rhythm Analysis

- Sensitivity for shockable rhythms: ≥90%

- Specificity for non-shockable rhythms: ≥95%

- Analysis time: not exceeding 15 seconds

Voice & Visual Prompts

- Multilingual voice prompts (configurable at supply)

- LED or LCD visual prompts

CPR Coaching

- Integrated metronome

- Real-time CPR quality feedback (rate and/or depth) --- where technology permits

Display (if fitted)

- Monochrome or color LCD; minimum 2 inches diagonal

- ECG waveform display capability (optional)

Event Recording

- Minimum 30 minutes of ECG and event data

- Data export via USB, IrDA, Bluetooth, or memory card

Self-Test

- Automatic daily self-test of battery, electronics, and pad connection

- Clear visible status indicator (ready/not ready)

Power Supply

- Long-life primary lithium battery or rechargeable battery

- Standby life (primary battery): minimum 4 years

- Shock capacity per battery: minimum 200 shocks at maximum energy, or minimum 6 hours of continuous monitoring

Physical Characteristics

- Weight: approximately 1.5--3.5 kg with battery and pads

- Dimensions: compact, single-hand carry

- Ingress protection: minimum IP54 or equivalent

- Drop resistance: minimum 1 m onto hard surface

Environmental

- Operating: 0--50 °C; 5--95% RH non-condensing

- Storage: −20 to +60 °C

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-2-4

- IEC 62304, ISO 13485, ISO 14971

- CE MDR / FDA 510(k) or equivalent

## 3.3  Standard Accessories

Adult Adhesive Electrode Pads (minimum 2 sets)

- Pre-gelled, single-use, individually sealed

- Shelf life: minimum 24 months

- Radiotranslucent, latex-free

Pediatric Pads or Pediatric Key (where applicable)

- Pediatric-sized pre-gelled adhesive pads OR pediatric mode key

- Shelf life: minimum 24 months (pads); indefinite (key)

Rechargeable Battery or Long-Life Lithium Battery

- Minimum standby life 4 years (primary) or 2 years (rechargeable)

- Shock/monitoring capacity as specified above

Wall Mounting Bracket

- Impact-resistant plastic or metal

- Universal fitting with screws and wall anchors

Carry Case

- Soft or semi-rigid, water-resistant

- Storage for AED, pads, and rescue kit

Rescue Kit

- CPR pocket mask with one-way valve

- Trauma scissors

- Disposable razor

- Nitrile examination gloves (minimum 2 pairs)

- Absorbent towel/gauze

- Alcohol-free wipes

Status Indicator Labels

- Adhesive labels indicating device status and last self-test

User Manual

- Printed and electronic, quick-reference guide included

AED Signage

- ISO-standard universal AED wall sign, high-visibility

Battery Charger (if rechargeable)

- Compatible with host battery, with charge status indication

## 3.4  Installation Requirements

- Wall-mounted in publicly accessible location (as per hospital plan)

- Ambient temperature within operating range

- No dedicated electrical supply required (battery-powered)

- Optional network connectivity for remote status monitoring

- Signage installed per ISO standards

## 3.5  Documentation

Supplier shall provide:

- User Manual and Quick Reference Guide

- Service Manual

- Self-Test Log Procedure

- Factory Test Certificate

- Electrical Safety Test Report per IEC 62353

- Preventive Maintenance Schedule

- Declaration of Conformity

## 3.6  Training

Supplier shall provide:

- BLS/AED clinical training for designated staff

- Biomedical engineering training on self-test verification and battery/pad management

- Basic troubleshooting training

- Refresher training scheduling

## 3.7  Warranty

Minimum 2 years comprehensive warranty on the AED unit, minimum 4 years pro-rated on primary battery (or as specified), including pads shelf-life replacement, technical support, and software updates. Spare-parts availability ≥8 years.

# 4. CRASH CART

Quantity: 24

## 4.1  Clinical & Functional Specifications

The equipment shall be a mobile emergency resuscitation cart designed to store and transport life-saving medications, consumables, defibrillator, monitoring equipment, and airway management supplies to the site of medical emergencies within the hospital. It shall provide rapid, organized access to emergency supplies during code-blue situations.

The cart shall support:

- Secure storage with tamper-evident locking

- Rapid single-action access during emergencies

- Mounting of defibrillator/monitor on top surface

- Mounting of IV pole and oxygen cylinder

- Compartmentalized storage for drugs, consumables, and airway equipment

- Mobility across hospital corridors, elevators, and door openings

## 4.2  Engineering Technical Specifications

Construction

- Body material: high-impact ABS plastic, powder-coated steel, or medical-grade aluminum

- Corrosion-resistant, non-porous, easy-to-clean surfaces

- Chemical-resistant finish compatible with hospital-grade disinfectants

- Rounded corners and edges for user safety

Dimensions (generalized)

- Height (excluding accessories): approximately 90--110 cm

- Width: approximately 55--75 cm

- Depth: approximately 45--65 cm

- Working top surface: minimum 40 × 55 cm

- Configuration suitable for standard hospital door widths

Weight

- Empty weight: approximately 40--75 kg depending on configuration

- Safe working load: minimum 150 kg total

Drawers

- Minimum 4 drawers of varied depth

- Full-extension ball-bearing or telescopic drawer slides

- Load capacity per drawer: minimum 15 kg

- Central locking system with breakaway plastic seal (numbered)

Locking System

- Single-point central locking of all drawers

- Numbered breakaway seals with tamper evidence

- Optional keyed or electronic lock

Mobility

- Minimum 4 swivel castors, diameter 100--150 mm

- Minimum 2 castors with total-lock brakes (wheel + swivel)

- Anti-static, non-marking castors suitable for hospital flooring

- Optional 5th directional castor for improved steering

Top Working Surface

- Reinforced to accept defibrillator/monitor weight (minimum 25 kg)

- Non-slip surface with raised edge or retaining lip

- Integrated cable management (where applicable)

Side Rails

- Stainless steel or equivalent, on both sides

- Removable or fixed, sized to accept standard hospital accessories

Environmental & Safety

- Antistatic castors and construction

- Fluid-resistant surfaces

- Compliance with fire-safety materials standards (as applicable)

Applicable Standards

- ISO 13485 (manufacturer QMS)

- Compliance with applicable furniture/medical device regulations

- CE marking or equivalent

## 4.3  Standard Accessories

CPR Backboard

- Rigid, radiolucent, high-density polymer or equivalent

- Dimensions: minimum 55 × 45 cm

- Integrated handles or attachment clips to the cart

- Wipe-cleanable, non-porous surface

IV Pole

- Stainless steel telescopic pole

- Minimum 4 hooks

- Adjustable height range: minimum 130--210 cm

- Locking mechanism at any position

Oxygen Cylinder Holder

- Sized for D or E-type cylinder (as specified)

- Secure retaining strap or clamp

- Mounted on side or rear of cart

Cardiac Board

- Rigid CPR board, radiolucent

- Storage location on cart side or dedicated slot

Utility Bins/Drawers

- Removable dividers and modular bins

- Transparent or labeled for rapid identification

- Autoclavable or wipe-cleanable

Medication Trays

- Removable, compartmentalized trays for ampoules/vials

- Anti-slip surface, labeled slots

- Compatible with color-coded drug organization systems

Locking System with Breakaway Seals

- Central lock with pack of minimum 50 numbered breakaway plastic seals supplied

Waste Bin

- Attached sharps or general-waste bin with secure lid

- Wipe-cleanable, chemical-resistant

Suction Pump Mounting Bracket (if applicable)

- Adjustable bracket sized for standard portable suction units

Monitor/Defibrillator Shelf

- Reinforced top shelf or pull-out shelf

- Load capacity: minimum 25 kg

- Retaining strap or bracket

Catheter Holders

- Side-mounted transparent tubes or clips for airway/catheter organization

Drawer Dividers

- Adjustable, modular, autoclavable or wipe-cleanable

Side Rails

- As specified in the engineering section above

Writing Board

- Fold-out or side-mounted clipboard/writing surface

- Wipe-cleanable

Caster Wheels with Brakes

- As specified in the engineering section above

## 4.4  Installation Requirements

- No electrical supply required (unless electronic lock specified --- then 100--240 VAC or battery-powered)

- Storage location with clearance for cart footprint and door access

- Compatible with hospital elevator and corridor widths

- Ambient hospital environmental conditions

## 4.5  Documentation

- User Manual with cleaning and disinfection procedures

- Assembly instructions

- Spare Parts Catalogue

- Certificate of Conformity

- Material Safety and Cleaning Compatibility Guide

## 4.6  Training

- User orientation on locking, seal replacement, and inventory management

- Cleaning and disinfection protocols

- Basic maintenance (castor lubrication, seal replacement)

## 4.7  Warranty

Minimum 2 years comprehensive warranty covering structural integrity, castors, locking mechanism, and drawer slides. Spare parts availability ≥10 years.

# 5. RESUSCITATION BED / TROLLEY WITH X-RAY TRANSLUCENT TOP

Quantity: 03

## 5.1  Clinical & Functional Specifications

The equipment shall be a mobile resuscitation trolley designed for emergency department resuscitation bays and shock rooms, providing a stable, X-ray translucent patient platform enabling imaging (chest X-ray, C-arm fluoroscopy) without patient transfer. It shall be suitable for adult and pediatric patients during resuscitation, initial stabilization, and diagnostic imaging.

The trolley shall support:

- Continuous X-ray imaging through the mattress platform

- Cassette insertion from head, foot, or side

- CPR-compatible rigid platform

- Height and backrest (Fowler) adjustment for airway management and clinical procedures

- Mobility with central braking and directional steering

- Attachment of monitoring, IV, and oxygen accessories

## 5.2  Engineering Technical Specifications

Construction

- Frame: powder-coated steel, stainless steel, or medical-grade aluminum

- Corrosion-resistant, wipe-cleanable, non-porous surfaces

- Chemical-resistant to hospital disinfectants

Patient Platform

- X-ray translucent material (carbon fiber, high-density polymer, or equivalent) covering full or major portion of platform

- Cassette tray with insertion from at least two sides (head/side)

- Compatible with standard imaging cassettes and C-arm

- CPR-rigid platform providing firm support during chest compressions

Dimensions (generalized)

- Overall length: approximately 200--220 cm

- Overall width: approximately 70--85 cm

- Mattress platform: minimum 190 × 60 cm

- Height range (adjustable): minimum 60--95 cm

Safe Working Load

- Minimum patient load capacity: 200 kg

- Total safe working load including accessories: minimum 250 kg

Adjustments

- Backrest (Fowler): 0° to at least 75°, gas-spring or hydraulic assisted

- Trendelenburg / Reverse Trendelenburg: minimum ±12°, mechanical or hydraulic

- Height adjustment (where specified): hydraulic or electric

Mobility

- Minimum 4 swivel castors, diameter 150--200 mm

- Central braking system operable from both sides (foot pedal)

- Directional steering (5th wheel) for corridor navigation

- Anti-static, non-marking castors

Bumpers

- Full-perimeter or corner bumpers to protect walls and equipment

Environmental

- Suitable for hospital ambient conditions

- Fluid-resistant construction

Applicable Standards

- IEC 60601-2-52 (Medical beds) where applicable

- ISO 13485, ISO 14971

- CE marking or equivalent

## 5.3  Standard Accessories

Mattress

- High-density foam with radiolucent core

- Fluid-proof, seamless, wipe-cleanable cover

- Antibacterial, antifungal, flame-retardant

- Thickness: minimum 70 mm

- Removable cover with waterproof zip

Collapsible Side Rails

- Full-length or split, on both sides

- Drop-down or fold-away mechanism, single-hand operable

- Chrome-plated or stainless steel

- Locking in raised position

IV Pole

- Stainless steel telescopic pole with minimum 4 hooks

- Height-adjustable: minimum 130--210 cm

- Corner-mounted with locking socket

Oxygen Cylinder Holder

- Sized for D/E-type cylinder

- Secure retaining mechanism, integrated to trolley frame

Utility Basket

- Removable stainless steel or plastic basket, minimum 40 × 30 × 15 cm

Drainage Hooks

- Minimum 4 hooks on each side, stainless steel

X-ray Cassette Holder (if applicable)

- Full-length or partial cassette tray

- Insertion from head and/or side

- Compatible with standard cassette sizes

Steering Handle

- At head end, ergonomic, chrome-plated or stainless steel

Central Braking System

- Foot-pedal operable from both sides

- Three-position: brake / neutral / directional steering

Bumpers

- Full-perimeter polymer bumpers

Patient Restraint Straps

- Minimum 2 pairs, wipe-cleanable, quick-release buckle

Mattress Cover

- Additional spare fluid-proof cover supplied

Foot/End Boards (where applicable)

- Removable, radiolucent, wipe-cleanable

## 5.4  Installation Requirements

- No electrical supply (mechanical/hydraulic) or 100--240 VAC UPS-backed supply (electric height adjustment)

- Clearance for full-length trolley in ED bays and imaging rooms

- Compatible with standard C-arm and portable X-ray equipment

- Compatible with hospital elevator dimensions

## 5.5  Documentation

- User Manual with cleaning/disinfection protocols

- Service Manual with hydraulic/mechanical schematics

- Spare Parts Catalogue

- Load Test Certificate

- X-ray Translucency Compliance Statement

- Declaration of Conformity

## 5.6  Training

- User training on adjustments, cassette insertion, and safe patient handling

- Preventive maintenance training (hydraulic checks, castor lubrication)

- Cleaning and disinfection protocols

## 5.7  Warranty

Minimum 2 years comprehensive warranty covering frame, castors, hydraulics, and mattress. Spare-parts availability ≥10 years.

# 6. PATIENT TRANSFER STRETCHER

Quantity: 17

## 6.1  Clinical & Functional Specifications

The equipment shall be a mobile patient transfer stretcher designed for safe intra-hospital transport of patients between wards, imaging, emergency department, operating theatre, and ICU. It shall support adult and pediatric patients under monitoring and IV therapy during transport.

The stretcher shall support:

- Height and backrest (Fowler) adjustment

- Trendelenburg / Reverse Trendelenburg positioning

- Side-rail patient protection

- IV, oxygen, and monitoring equipment mounting

- Central braking and directional steering

- Transfer board and CPR compatibility

## 6.2  Engineering Technical Specifications

Construction

- Frame: powder-coated steel or medical-grade aluminum

- Chemical-resistant, wipe-cleanable finish

Patient Platform

- Multi-section platform with radiolucent or standard construction (per BOQ)

- CPR-rigid backrest section

Dimensions (generalized)

- Overall length: approximately 195--215 cm

- Overall width: approximately 65--80 cm

- Platform: minimum 190 × 55 cm

- Height range: minimum 55--90 cm

Safe Working Load

- Minimum 200 kg patient load

Adjustments

- Backrest: 0° to at least 70°, gas-spring assisted

- Trendelenburg / Reverse Trendelenburg: minimum ±12°

- Height adjustment: hydraulic (foot pump) or electric per BOQ

Mobility

- Minimum 4 swivel castors, diameter 150--200 mm

- Central braking from both sides

- Directional (5th) wheel for corridor navigation

- Anti-static, non-marking castors

Bumpers

- Full-perimeter or corner bumpers

Environmental

- Standard hospital ambient conditions

Applicable Standards

- IEC 60601-2-52 (where applicable)

- ISO 13485

- CE marking or equivalent

## 6.3  Standard Accessories

Mattress

- High-density foam, minimum 60 mm thickness

- Fluid-proof, antibacterial, seamless cover with waterproof zip

Side Rails

- Full-length or split, collapsible, chrome-plated or stainless steel

- Single-hand operable with positive locking

IV Pole

- Telescopic stainless steel, minimum 4 hooks

- Adjustable 130--210 cm

Oxygen Cylinder Holder

- Sized for D/E-type cylinder, with secure retention

Patient Safety Belts (minimum 2)

- Wipe-cleanable, quick-release buckle

- Adjustable length

Central Brake System

- Foot-pedal operable from both sides

Steering Mechanism

- 5th directional wheel or equivalent

Storage Basket

- Removable stainless steel or plastic basket

Push Handles

- At head end, ergonomic, foldable or fixed

Caster Wheels

- As specified above

Bumpers

- Full-perimeter polymer bumpers

Transfer Board Compatibility Accessories

- Side clearance and mattress design compatible with standard transfer boards

Mattress Cover

- Spare fluid-proof cover supplied

## 6.4  Installation Requirements

- No electrical supply (mechanical/hydraulic) or 100--240 VAC (electric variants)

- Compatible with hospital corridors, elevators, and doorways

- No special environmental requirements

## 6.5  Documentation

- User Manual and cleaning guide

- Service Manual

- Spare Parts Catalogue

- Load Test Certificate

- Declaration of Conformity

## 6.6  Training

- User training on adjustments and safe transfer techniques

- Preventive maintenance training

## 6.7  Warranty

Minimum 2 years comprehensive warranty; spare-parts availability ≥10 years.

# 7. TRAUMA TROLLEY

Quantity: 01

## 7.1  Clinical & Functional Specifications

The equipment shall be a heavy-duty, X-ray translucent trauma trolley designed for use in emergency department trauma bays, receiving patients with polytrauma, spinal injury, and cardiovascular collapse. It shall enable rapid resuscitation, imaging, and stabilization without patient transfer.

The trolley shall support:

- X-ray imaging through the patient platform

- CPR-compatible rigid platform

- Backrest and Trendelenburg positioning

- Mounting of IV, oxygen, monitor, and drainage accessories

- Rapid mobility with central braking

## 7.2  Engineering Technical Specifications

Construction

- Heavy-duty frame: powder-coated steel or medical-grade aluminum

- Corrosion-resistant, chemical-resistant finish

Patient Platform

- X-ray translucent material (carbon fiber, high-density polymer, or equivalent)

- Cassette tray insertable from head, foot, or side

- CPR-rigid throughout

Dimensions (generalized)

- Length: approximately 200--220 cm

- Width: approximately 70--85 cm

- Height range: minimum 60--95 cm

Safe Working Load

- Minimum 250 kg patient load

Adjustments

- Backrest: 0° to at least 75°

- Trendelenburg / Reverse Trendelenburg: minimum ±15°

- Height adjustment: hydraulic or electric

Mobility

- 4 swivel castors, diameter 150--200 mm

- Central braking system from both sides

- Directional 5th wheel

- Anti-static, non-marking castors

Environmental

- Standard hospital ambient conditions

Applicable Standards

- IEC 60601-2-52 (where applicable)

- ISO 13485

- CE marking or equivalent

## 7.3  Standard Accessories

X-ray Translucent Mattress Platform

- Full-length radiolucent surface

- Compatible with standard imaging cassettes and C-arm

Mattress

- High-density foam with radiolucent core

- Minimum 70 mm thickness

- Fluid-proof, antibacterial, seamless cover

Side Rails

- Collapsible full-length or split rails on both sides

- Chrome-plated or stainless steel with positive locking

IV Pole

- Telescopic stainless steel, minimum 4 hooks, 130--210 cm range

Oxygen Cylinder Holder

- Sized for D/E-type cylinder with secure retention

Patient Restraint Straps

- Minimum 3 pairs, wipe-cleanable, quick-release buckle

Drainage Hooks

- Minimum 4 per side, stainless steel

Steering Handles

- Ergonomic head-end handles

Central Brake System

- Foot-pedal operable from both sides, three-position (brake/neutral/steer)

Caster Wheels

- As specified above

Bumpers

- Full-perimeter polymer bumpers

Utility Basket

- Removable, minimum 40 × 30 × 15 cm

CPR Board Attachment (if applicable)

- Integrated CPR-rigid platform or removable CPR board with attachment clips

## 7.4  Installation Requirements

- No electrical supply (mechanical/hydraulic) or 100--240 VAC UPS-backed (electric variants)

- Compatible with C-arm, portable X-ray, and CT imaging environments (as applicable)

- Compatible with hospital corridor and elevator dimensions

## 7.5  Documentation

- User Manual with cleaning/disinfection protocols

- Service Manual

- Spare Parts Catalogue

- Load Test Certificate

- X-ray Translucency Compliance Statement

- Declaration of Conformity

## 7.6  Training

- User training on adjustments, imaging cassette use, and patient handling

- Preventive maintenance training

## 7.7  Warranty

Minimum 2 years comprehensive warranty; spare parts availability ≥10 years.

# 8. SPINAL BOARDS, CERVICAL COLLARS & SPLINTS (IMMOBILISATION SET)

Quantity: 02

## 8.1  Clinical & Functional Specifications

The equipment shall be a complete pre-hospital and in-hospital immobilization set intended for stabilization and transport of patients with suspected spinal injury, cervical injury, and long-bone fractures. It shall be suitable for adult, pediatric, and infant patients.

The set shall support:

- Full-body spinal immobilization

- Cervical spine stabilization across all patient sizes

- Long-bone (upper and lower limb) splinting

- Compatibility with standard ambulance and hospital transfer equipment

- X-ray, CT, and MRI compatibility (radiolucent materials)

## 8.2  Engineering Technical Specifications

Long Spine Board

- Material: high-density polyethylene, polymer composite, or equivalent radiolucent material

- X-ray, CT, and MRI compatible (metal-free construction)

- Dimensions: minimum length 180 cm, width 40--45 cm, thickness 4--5 cm

- Safe working load: minimum 150 kg (typical 180--250 kg)

- Water-resistant, floats in water (positive buoyancy)

- Non-porous, chemical-resistant, wipe-cleanable

- Minimum 8 handholds (4 per side) with pin holes for strap attachment

- Compatible with standard restraint straps

Cervical Collars

- Rigid or semi-rigid polymer construction

- Radiolucent, MRI-compatible

- Adjustable or fixed sizing covering: infant, pediatric, small adult, medium adult, large adult

- Foam-padded chin and posterior support

- Tracheal opening for airway access and pulse check

- Aperture for posterior palpation

Head Immobilizer

- Two lateral foam blocks with hook-and-loop straps

- Radiolucent

- Wipe-cleanable, fluid-resistant cover

Restraint Straps

- Minimum 3 body straps + 2 head straps

- Quick-release buckles, color-coded

- Adjustable length, minimum 150 cm

- Wipe-cleanable

Limb Splints

- Padded, malleable aluminum, polymer, or vacuum splint construction

- Radiolucent (aluminum-free variants for CT/MRI where specified)

- Sizes for arm and leg (adult and pediatric)

- Adjustable, secured by hook-and-loop or strap closures

Environmental

- Operating temperature range: −30 to +60 °C

- Fluid-resistant, non-porous surfaces

- Compatible with standard hospital disinfectants

Applicable Standards

- ISO 13485

- Compliance with EN 1789 (Ambulance equipment, where applicable)

- CE marking or equivalent

## 8.3  Standard Accessories

Long Spine Board with Handholds

- As specified above; supplied with strap-attachment pin holes

Head Immobilizer Blocks

- Pair of foam lateral blocks with base plate; hook-and-loop attachment

Head Straps

- Minimum 2 straps (forehead and chin), wipe-cleanable, quick-release

Body Restraint Straps (minimum 3)

- Adjustable, color-coded, quick-release buckles

Cervical Collars (complete size set)

- Infant, pediatric, small adult, medium, large --- one of each minimum

- Rigid, radiolucent, MRI-compatible

Limb Splints (arm and leg)

- Adjustable, radiolucent (where specified)

- Arm splint set: minimum 2 sizes

- Leg splint set: minimum 2 sizes

Carrying Bag/Storage Case

- Durable, water-resistant, with compartments for all components

- Carry handles and shoulder strap

Instruction Manual

- Printed, illustrated, in English (and local language if specified)

Cleaning/Disinfection Guide

- Printed guide including approved chemical compatibility list

## 8.4  Installation Requirements

- Storage location in ED, ambulance, or ward with easy access

- No electrical supply required

- No special environmental requirements beyond hospital ambient

## 8.5  Documentation

- User Manual

- Cleaning and Disinfection Compatibility Guide

- Material Safety Data Sheets

- Certificate of Conformity

- Load Test Certificate for spine board

## 8.6  Training

- User training on immobilization technique, collar sizing, and strap application

- Cleaning and disinfection training

## 8.7  Warranty

Minimum 2 years comprehensive warranty on all rigid components; minimum 1 year on straps and consumable soft components. Replacement parts availability ≥5 years.

# 9. EYE-WASH STATION

Quantity: 01

## 9.1  Clinical & Functional Specifications

The equipment shall be an emergency eye-wash station intended for immediate decontamination of eyes and face following accidental exposure to chemicals, biological hazards, or foreign bodies in laboratories, pharmacy, sterile services, and clinical areas. Two configurations shall be supplied per BOQ requirement:

- Plumbed Eye-Wash Station --- permanently connected to potable water supply

- Portable Eye-Wash Station --- self-contained gravity-fed unit with buffered saline solution

The station shall support:

- Simultaneous irrigation of both eyes and face

- Hands-free continuous operation upon activation

- Activation within 1 second of user actuation

- Continuous flow for a minimum of 15 minutes per ANSI Z358.1

## 9.2  Engineering Technical Specifications

Compliance

- ANSI/ISEA Z358.1 (Emergency Eyewash and Shower Equipment) or equivalent international standard

- EN 15154-2 (where applicable)

Flow Rate

- Plumbed station: minimum 1.5 L/min (0.4 GPM) continuous for minimum 15 minutes

- Portable station: minimum 1.5 L/min continuous for minimum 15 minutes

Water Pressure (plumbed)

- Operating range: 210--620 kPa (30--90 psi) or as applicable

Water Temperature (plumbed)

- Tepid water: 16--38 °C (60--100 °F) delivered flow, per ANSI Z358.1

Spray Heads

- Dual anti-microbial spray heads

- Aerated or laminar flow producing controlled, gentle rinse

- Dust covers (flip-off or pop-off) automatically removed upon activation

Activation

- Push-flag/paddle handle and/or foot pedal, hands-free continuous flow

- Activation force: low-effort, single-motion

- Valve remains open until manually closed

Bowl / Drainage Basin

- Corrosion-resistant material: ABS, stainless steel, or polypropylene

- Minimum diameter/width: 25 cm

- Drain outlet with tail-piece for plumbed connection

Materials

- Wetted parts: corrosion-resistant, chemical-resistant, lead-free

- External surfaces: high-visibility yellow-green color per ANSI standard

Portable Unit Specifications

- Reservoir capacity: minimum 22 litres for 15 minutes of continuous flow

- Reservoir material: high-density polyethylene, UV-resistant

- Wall-mountable or floor-standing configuration

- Solution: buffered preservative saline or approved eyewash solution

- Shelf life of sealed solution: minimum 18 months

Signage

- ANSI-compliant high-visibility identification sign

- Inspection tag holder

Environmental

- Operating temperature: 5--45 °C

- Frost-free installation (plumbed units)

Applicable Standards

- ANSI/ISEA Z358.1

- EN 15154-2 (where applicable)

- ISO 13485 (manufacturer QMS)

- CE marking or equivalent

## 9.3  Standard Accessories

Dual Spray Heads

- Anti-microbial, aerated or laminar-flow, corrosion-resistant

- Integrated flow regulators

Dust Covers

- Automatic flip-off or pop-off type

- Retained by lanyard to prevent loss

Activation Handle/Foot Pedal

- Push-flag/paddle handle for hand activation

- Foot pedal (where model specifies), hands-free

- Positive-lock open position

Drainage Basin

- Corrosion-resistant bowl with drain outlet

- Minimum diameter 25 cm

Mounting Bracket

- Wall-mount or pedestal bracket with stainless steel or corrosion-resistant fixings

- Compatible with standard hospital wall construction

Plumbing Connection Kit (plumbed models)

- Inlet supply fitting (½-inch NPT or metric equivalent)

- Outlet drain tail-piece

- Isolation valve

- Flexible or rigid supply piping and fittings

Eyewash Solution Bottles (portable models)

- Buffered, preserved sterile eyewash solution

- Total supplied volume matching reservoir capacity plus one refill

- Shelf life minimum 18 months

Inspection Tag

- Weekly inspection tag with tie or clip attachment

- Preprinted inspection log format

Warning Signage

- ANSI-compliant identification sign with universal eye-wash symbol

- Fluorescent/high-visibility background

User Manual

- Printed and electronic, including installation, activation, weekly testing, and maintenance procedures

Maintenance Kit

- Replacement spray heads (minimum 1 set)

- Replacement dust covers (minimum 1 set)

- Inspection tags (minimum 12 units)

- Cleaning brush and disinfection wipes as applicable

## 9.4  Installation Requirements

Plumbed Units

- Potable water supply: ½-inch NPT or metric-equivalent connection

- Water pressure: 210--620 kPa (30--90 psi)

- Tepid water mixing valve (thermostatic) where required to deliver 16--38 °C

- Drain connection to hospital wastewater system

- Wall or pedestal mounting per manufacturer requirements

- Location: within 10 seconds unobstructed travel from hazard, per ANSI Z358.1

- Minimum clearance: 40 cm from wall on activation side

Portable Units

- Wall or floor mounting per manufacturer instructions

- Ambient temperature within operating range (frost-free)

General

- High-visibility signage installed at station and above eye-level

- Unobstructed access, clearly marked path

- Adequate lighting at station location

- No electrical supply required

## 9.5  Documentation

- User Manual with installation, operation, and weekly inspection procedures

- Compliance Certificate for ANSI/ISEA Z358.1 or equivalent

- Flow Rate Test Certificate

- Material Compatibility and Chemical Resistance Data

- Preventive Maintenance Schedule

- Weekly Inspection Log Template

- Declaration of Conformity

## 9.6  Training

- User training on activation, correct rinsing technique (15-minute minimum), and emergency response protocol

- Facility staff training on weekly inspection and monthly testing per ANSI Z358.1

- Biomedical/Facilities training on preventive maintenance, spray head replacement, and solution replenishment

## 9.7  Warranty

Minimum 2 years comprehensive warranty covering valves, spray heads, and structural components. Replacement solution and consumables supplied for warranty period. Spare-parts availability ≥10 years.