Hospital Endoscopy Suite — Technical Tender Specifications

Technical Tender Document

Table of Contents

1. Flexible Video Bronchoscope with Light Source and Processor	4

1.1 Clinical & Functional Specifications	4

1.2 Engineering Technical Specifications	4

1.3 Standard Accessories	6

1.4 Installation Requirements	7

1.5 Documentation	7

1.6 Training	7

1.7 Warranty	7

2. Upper GI Video Endoscope (Gastroscope)	8

2.1 Clinical & Functional Specifications	8

2.2 Engineering Technical Specifications	8

2.3 Standard Accessories	9

2.4 Installation Requirements	10

2.5 Documentation	10

2.6 Training	10

2.7 Warranty	10

3. Lower GI Video Endoscope (Colonoscope)	11

3.1 Clinical & Functional Specifications	11

3.2 Engineering Technical Specifications	11

3.3 Standard Accessories	12

3.4 Installation Requirements	12

3.5 Documentation	13

3.6 Training	13

3.7 Warranty	13

4. Duodenoscope (Side-Viewing Therapeutic, for ERCP)	14

4.1 Clinical & Functional Specifications	14

4.2 Engineering Technical Specifications	14

4.3 Standard Accessories	15

4.4 Installation Requirements	16

4.5 Documentation	16

4.6 Training	16

4.7 Warranty	17

# 1. Flexible Video Bronchoscope with Light Source and Processor

Quantity: 2

## 1.1 Clinical & Functional Specifications

The system shall be a complete flexible video bronchoscopy platform intended for diagnostic and therapeutic visualization of the tracheobronchial tree in adult and, where configured, pediatric patients. The system shall enable direct visual airway examination, bronchoalveolar lavage (BAL), transbronchial biopsy, foreign body retrieval, difficult airway management, percutaneous tracheostomy guidance, endotracheal tube placement confirmation, aspiration of secretions, and guided therapeutic interventions. The complete platform shall consist of a flexible video insertion scope, a digital video image processor, an integrated or standalone high-intensity light source, a medical-grade display monitor, and a mobile workstation cart with cable management. The system shall support real-time high-definition imaging, image freezing, digital capture (still and video), electronic patient record annotation, and DICOM-compliant export to hospital PACS/HIS. The device shall support both bedside/ICU/OT use and dedicated bronchoscopy suite operation.

## 1.2 Engineering Technical Specifications

Video Insertion Scope (Bronchoscope)

- Image sensor: distal-tip CMOS or CCD, high-definition, color, equivalent or better

- Optical resolution: minimum HD (1280 × 720) at scope tip; system output HD/Full HD compliant

- Field of view: minimum 110°, adjustable/wide-angle equivalent or better

- Depth of field: approximately 3—50 mm

- Distal tip outer diameter: adult configuration approximately 4.9—6.2 mm; slim/intubation configuration approximately 3.8—4.5 mm (as specified by purchaser)

- Working channel diameter: minimum 2.0 mm (adult); minimum 1.2 mm (slim) suitable for suction and standard bronchoscopic accessories

- Insertion tube working length: minimum 550—600 mm

- Total scope length: minimum 850—900 mm

- Angulation range: up minimum 180°, down minimum 130° (four-way or two-way as per clinical configuration)

- Illumination: integrated fiber-optic light guide bundle

- Watertight/immersible construction for full reprocessing

- Autoclave/ETO/high-level disinfection compatibility per manufacturer's validated cycle

Video Image Processor

- Native HD/Full HD digital signal processing (1920 × 1080) or better

- Digital image enhancement: contrast, edge enhancement, structure enhancement, and narrow-band/spectral imaging or equivalent

- Multiple observation modes: standard white light plus at least one advanced imaging mode (spectral/narrow-band/color-enhancement) selectable from control panel

- Video outputs: minimum HD-SDI, DVI, HDMI, and USB (any equivalent set)

- Still image capture: JPEG/TIFF or equivalent

- Video recording: MPEG-4/H.264 or equivalent, to internal storage and/or external USB

- Internal storage: minimum 250 GB or removable media support

- Patient data entry via keyboard, touchscreen, or barcode

- DICOM 3.0 compliant output; HL7 worklist support

- Network: Ethernet 10/100/1000, wired; optional wireless

Light Source

- LED or Xenon high-intensity medical light source, equivalent or better

- Lamp life: LED minimum 20,000 hours; Xenon minimum 500 hours

- Color temperature: approximately 5,500—6,500 K (daylight equivalent)

- Adjustable brightness: continuous or multi-step, minimum 8 levels

- Automatic and manual brightness control

- Standby mode and lamp-hour counter

Display Monitor

- Medical-grade LCD/LED or OLED, minimum 24-inch diagonal

- Native resolution: minimum Full HD 1920 × 1080

- Brightness: minimum 400 cd/m²; contrast ratio minimum 1000:1

- Viewing angle: minimum 170° horizontal/vertical

- Inputs: HD-SDI, DVI, HDMI, VGA (any equivalent complete set)

- Compliant with IEC 60601-1 medical display requirements

Mobile Workstation Cart

- Medical-grade powder-coated steel or equivalent chassis with antimicrobial finish

- Isolation transformer: minimum 500 VA, medical grade, IEC 60601-1 compliant

- Multiple hospital-grade outlets with individual protection

- Locking castors (minimum 2 with brakes)

- Integrated cable management, scope hanger, keyboard tray, and equipment shelves

Electrical & Environmental

- Power supply: 100—240 VAC, 50/60 Hz, single phase

- Power consumption: typical up to 600 VA

- Operating temperature: 10—40 °C; humidity 30—85% non-condensing

- Storage temperature: −20 to +60 °C

Applicable Standards

- IEC 60601-1 (General safety)

- IEC 60601-1-2 (EMC)

- IEC 60601-2-18 (Endoscopic equipment)

- IEC 62304 (Software lifecycle)

- IEC 62366 (Usability)

- ISO 13485 (QMS)

- ISO 14971 (Risk management)

- DICOM 3.0, HL7

- CE MDR and/or US FDA 510(k) clearance

## 1.3 Standard Accessories

a) Flexible video bronchoscope insertion tube with high-resolution CMOS/CCD tip

- Distal-tip CMOS/CCD sensor, HD or better

- Four-way or two-way angulation as specified

- Working channel minimum 2.0 mm (adult) or 1.2 mm (slim)

- Fully immersible, reprocessing-compatible

- Autoclave/ETO/high-level disinfection validated

b) Combined HD video processor and high-intensity LED/Xenon light source console

- Integrated or stacked configuration

- HD/Full HD image processing with advanced imaging mode

- Multiple video outputs, DICOM/HL7 support

- LED ≥ 20,000 h or Xenon ≥ 500 h lamp life

c) High-resolution medical-grade LCD/OLED display monitor on mobile workstation cart

- Minimum 24-inch, Full HD, medical grade, IEC 60601-1 compliant

- Mounted on lockable, height-adjustable arm on the cart

d) Reusable biopsy valves and suction channel cleaning brushes

- Autoclavable silicone/elastomer biopsy valves, minimum 10 pieces

- Double-ended nylon/stainless bristle cleaning brushes sized for working channel, minimum 20 pieces

- Reusable, high-level disinfection compatible

e) Watertight leakage tester with pressure gauge and adapter tubing

- Manual or automatic pressure pump, calibrated gauge

- Pressure range appropriate for scope integrity testing (typically 150—250 mmHg)

- Adapter compatible with supplied scopes

f) Endotracheal tube (ETT) bite blocks and mouthpiece guards

- Medical-grade polymer, autoclavable or single-use as specified

- Adult and pediatric sizes, minimum 10 pieces each

g) Water bottle assembly with connecting air tube

- Autoclavable polycarbonate bottle, minimum 250 mL

- Silicone/PTFE air/water tubing with quick-connect fitting

h) Rigid protective transport and sterilization storage case

- Impact-resistant polymer/aluminum construction

- Custom foam insert for scope, processor, and accessories

- Lockable, with humidity indicator

## 1.4 Installation Requirements

- Electrical: 100—240 VAC, 50/60 Hz, dedicated 15 A circuit with proper earthing (earth resistance \< 1 Ω)

- UPS compatibility: minimum 1 kVA online UPS recommended

- HVAC: ambient 18—26 °C, humidity 40—60%

- Network: RJ-45 Ethernet drop for PACS/HIS connectivity

- Space: minimum 2 m² floor area for cart mobility

- No medical gas or water supply requirement at unit level

## 1.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), Preventive Maintenance Schedule, Installation Report, Commissioning Report, DICOM Conformance Statement.

## 1.6 Training

Clinical training for pulmonologists/anesthetists and endoscopy nurses; biomedical engineering training on preventive maintenance and troubleshooting; reprocessing/infection control training; application training on advanced imaging modes and DICOM workflow.

## 1.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, spare parts, preventive maintenance (minimum two visits per year), breakdown maintenance, software updates, and technical support. Insertion tube warranty against manufacturing defects per international norms.

# 2. Upper GI Video Endoscope (Gastroscope)

Quantity: 1

## 2.1 Clinical & Functional Specifications

The system shall be a complete high-definition upper gastrointestinal video endoscopy platform for diagnostic and therapeutic examination of the esophagus, stomach, and duodenum in adult patients. Clinical applications shall include diagnostic endoscopy, biopsy, polypectomy, hemostasis, foreign body removal, dilation guidance, variceal band ligation, PEG placement, and advanced imaging for early neoplasia detection. The system shall consist of a flexible upper GI video scope with four-way angulation, a digital HD video image processor with digital enhancement imaging, a high-output light source, a large medical-grade display monitor, and a dedicated endoscopy cart. The system shall provide real-time HD imaging, still/video capture, patient data management, DICOM export, and HL7 worklist integration.

## 2.2 Engineering Technical Specifications

Gastroscope (Insertion Scope)

- HD CMOS/CCD distal-tip sensor, equivalent or better

- Field of view: minimum 140°

- Depth of field: approximately 3—100 mm

- Distal end outer diameter: approximately 9.0—9.9 mm

- Insertion tube outer diameter: approximately 9.0—9.9 mm

- Working channel: minimum 2.8 mm

- Working length: minimum 1050 mm

- Angulation: up ≥ 210°, down ≥ 90°, right ≥ 100°, left ≥ 100° (four-way)

- Fully immersible, autoclave-valve compatible where specified

Video Image Processor

- HD/Full HD 1920 × 1080 native processing, equivalent or better

- Digital enhancement imaging: structure, contrast, and spectral/narrow-band imaging or equivalent

- Minimum two selectable advanced imaging modes

- Video outputs: HD-SDI, DVI, HDMI, USB (equivalent complete set)

- Still capture: JPEG/TIFF/PNG; video: MPEG-4/H.264 or equivalent

- Internal storage: minimum 250 GB, plus USB/external media

- DICOM 3.0 output, HL7 worklist, Gigabit Ethernet

- Patient data via keyboard, touchscreen, or barcode

Light Source

- High-output Xenon (≥ 300 W equivalent) or high-intensity LED

- Lamp life: Xenon ≥ 500 h; LED ≥ 20,000 h

- Color temperature approximately 5,500—6,500 K

- Multi-level adjustable brightness with auto/manual control

- Integrated air pump for insufflation (adjustable, minimum 3 levels)

Display Monitor

- Medical-grade LCD/LED, 24-inch diagonal

- Native Full HD 1920 × 1080 or better

- Brightness ≥ 400 cd/m²; contrast ≥ 1000:1

- Viewing angle ≥ 170°; multiple inputs (HD-SDI, DVI, HDMI, VGA)

- IEC 60601-1 compliant

Endoscopy Cart

- Medical-grade construction with isolation transformer (minimum 1000 VA)

- Hospital-grade outlets, locking castors, cable management, scope hanger, keyboard tray

Electrical & Environmental

- Power: 100—240 VAC, 50/60 Hz

- Consumption: typical up to 800 VA

- Operating: 10—40 °C, 30—85% RH non-condensing

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-2-18

- IEC 62304, 62366; ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR / US FDA 510(k)

## 2.3 Standard Accessories

a) Flexible upper GI video scope with 4-way angulation controls

- HD CMOS/CCD tip; four-way angulation with locking controls

- Working channel ≥ 2.8 mm; working length ≥ 1050 mm; fully immersible

b) HD video image processor with digital enhancement imaging software

- Full HD processing, multiple advanced imaging modes, DICOM/HL7 compliant

c) High-output Xenon or LED medical light source console

- Xenon ≥ 300 W equivalent or LED with equivalent output; multi-level brightness

d) 24-inch medical-grade flat panel display monitor on endoscopy cart

- Full HD, medical-grade, IEC 60601-1 compliant, articulated cart-mount

e) Reusable autoclavable suction and air/water valve buttons

- Silicone/elastomer, color-coded, autoclavable, minimum 5 sets

f) Autoclavable biopsy channel port valves

- Silicone/elastomer, autoclavable, minimum 10 pieces

g) Water container bottle with stainless steel mounting bracket and air/water feed line

- Autoclavable polycarbonate bottle ≥ 250 mL, stainless bracket, silicone/PTFE tubing

h) Watertight immersion leakage tester

- Manual/automatic pressure pump, calibrated gauge, scope-compatible adapter

i) Reusable channel cleaning brushes and suction line cleaning adapters

- Double-ended nylon/stainless brushes sized for channel; minimum 20 brushes plus adapters

j) Reusable adult patient mouth guards/bite blocks

- Autoclavable polymer, adult size, minimum 10 pieces

## 2.4 Installation Requirements

Electrical 100—240 VAC, dedicated 15 A circuit, proper earthing; UPS ≥ 1.5 kVA recommended; HVAC 18—26 °C, 40—60% RH; RJ-45 network drop for PACS/HIS; minimum 3 m² footprint.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), PM Schedule, Installation Report, Commissioning Report, DICOM Conformance Statement.

## 2.6 Training

Clinical training for gastroenterologists and endoscopy nursing staff; biomedical training on PM and troubleshooting; reprocessing training; application training on advanced imaging and DICOM/HL7 integration.

## 2.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, spare parts, PM (two visits/year minimum), breakdown maintenance, software updates, and technical support.

# 3. Lower GI Video Endoscope (Colonoscope)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The system shall be a complete HD lower gastrointestinal video endoscopy platform for diagnostic and therapeutic examination of the colon and terminal ileum in adult patients. Clinical applications shall include screening and diagnostic colonoscopy, polypectomy, EMR, hemostasis, biopsy, dilation, tattooing, and advanced imaging for early neoplasia detection. The scope shall provide graduated stiffness (variable flexibility) for optimized insertion and patient comfort. The system shall consist of a variable-stiffness video colonoscope, HD digital video processor, high-intensity light source, HD monitor, and mobile workstation cart with full data management, DICOM export, and HL7 worklist support.

## 3.2 Engineering Technical Specifications

Colonoscope (Insertion Scope)

- HD CMOS/CCD distal-tip sensor

- Field of view: minimum 140—170°

- Depth of field: approximately 3—100 mm

- Distal end/insertion tube outer diameter: approximately 11.5—13.2 mm

- Working channel: minimum 3.2 mm (suitable for standard therapeutic accessories)

- Working length: minimum 1600—1700 mm

- Angulation: up ≥ 180°, down ≥ 180°, right ≥ 160°, left ≥ 160°

- Graduated/variable insertion tube stiffness, adjustable by user

- Fully immersible; auxiliary water-jet channel supported

Video Image Processor

- HD/Full HD 1920 × 1080 native processing

- Digital enhancement with minimum two selectable advanced imaging modes (spectral/narrow-band/structure/contrast or equivalent)

- Video outputs: HD-SDI, DVI, HDMI, USB

- Still/video capture with internal storage ≥ 250 GB

- DICOM 3.0, HL7 worklist, Gigabit Ethernet

Light Source

- Xenon ≥ 300 W equivalent or high-intensity LED

- Lamp life: Xenon ≥ 500 h; LED ≥ 20,000 h

- Color temperature approximately 5,500—6,500 K

- Multi-level adjustable brightness; integrated CO₂/air insufflation interface

Display Monitor

- 24-inch medical-grade Full HD, ≥ 400 cd/m², contrast ≥ 1000:1, IEC 60601-1 compliant

Mobile Workstation Cart

- Medical-grade construction, isolation transformer ≥ 1000 VA, hospital-grade outlets, locking castors, cable management

Electrical & Environmental

- 100—240 VAC, 50/60 Hz; typical up to 800 VA

- Operating 10—40 °C, 30—85% RH

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-2-18; IEC 62304, 62366; ISO 13485, 14971; DICOM 3.0; HL7; CE MDR / FDA 510(k)

## 3.3 Standard Accessories

a) Flexible video colonoscope with graduated stiffness and variable insertion tube flexibility

- HD tip, variable stiffness control, four-way angulation, working channel ≥ 3.2 mm, working length ≥ 1600 mm, fully immersible

b) HD digital video processor system

- Full HD, advanced imaging modes, DICOM/HL7 compliant

c) Medical-grade high-intensity light source console

- Xenon or LED with lamp life as specified, multi-level brightness

d) High-resolution diagnostic endoscopy monitor with mobile workstation cart

- 24-inch Full HD medical-grade monitor on cart with isolation transformer and cable management

e) Reusable air/water and suction valve kits

- Autoclavable silicone/elastomer, color-coded, minimum 5 sets

f) Removable biopsy channel inlet valves

- Autoclavable silicone/elastomer, minimum 10 pieces

g) High-capacity water bottle system with air/water supply tube

- Autoclavable polycarbonate bottle ≥ 250 mL with silicone/PTFE tubing

h) Handheld immersion leak tester with pressure gauge

- Manual/automatic pressure pump, calibrated gauge, scope-compatible

i) Specialized channel cleaning brushes and port adapters

- Double-ended brushes sized for 3.2 mm channel, port cleaning adapters, minimum 20 pieces

j) Auxiliary water injection pump interface connection cable

- OEM-compatible interface cable for auxiliary water pump connection, autoclavable/reprocessing compatible

## 3.4 Installation Requirements

Electrical 100—240 VAC, dedicated 15 A circuit, proper earthing; UPS ≥ 1.5 kVA recommended; HVAC 18—26 °C, 40—60% RH; network drop for PACS/HIS; provision for CO₂ insufflator connection if used; minimum 3 m² footprint.

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), PM Schedule, Installation Report, Commissioning Report, DICOM Conformance Statement.

## 3.6 Training

Clinical training for gastroenterologists/endoscopy nurses; biomedical PM and troubleshooting training; reprocessing training; application training on advanced imaging and DICOM/HL7 workflow.

## 3.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, parts, PM (two visits/year minimum), breakdown maintenance, software updates, and technical support.

# 4. Duodenoscope (Side-Viewing Therapeutic, for ERCP)

Quantity: 1

## 4.1 Clinical & Functional Specifications

The system shall be a complete side-viewing therapeutic video duodenoscope platform for endoscopic retrograde cholangiopancreatography (ERCP) and related biliary/pancreatic therapeutic interventions in adult patients. Clinical applications shall include cannulation of the ampulla of Vater, sphincterotomy, stone extraction, stricture dilation, biliary/pancreatic stenting, cytology brushing, cholangioscopy access, and tissue sampling. The scope shall include a distal elevator mechanism for precise accessory guidance and a detachable/removable distal cap or tip design to permit thorough elevator-channel reprocessing in line with current international infection-prevention recommendations. The complete platform shall include an HD video processor with ERCP-oriented recording capability, high-intensity light source, dual medical-grade monitors (procedure and reference/live-preview), a heavy-duty endoscopy cart, and full DICOM/HL7 integration.

## 4.2 Engineering Technical Specifications

Duodenoscope (Side-Viewing)

- HD CMOS/CCD distal side-viewing sensor

- Field of view: minimum 100—120° (side-viewing configuration)

- Depth of field: approximately 5—60 mm

- Distal end outer diameter: approximately 11.3—13.5 mm

- Working channel: minimum 4.2 mm (therapeutic; suitable for large-bore ERCP accessories)

- Working length: minimum 1240 mm

- Angulation: up ≥ 120°, down ≥ 90°, right ≥ 105°, left ≥ 90°

- Forceps elevator with precise control lever

- Detachable/removable distal cap or tip for elevator-channel access and reprocessing

- Fully immersible; validated for high-level disinfection and/or sterilization per manufacturer's approved cycle

Video Image Processor

- Full HD 1920 × 1080 native processing

- ERCP-oriented workflow with dedicated recording capability (still and video)

- Digital enhancement imaging with selectable advanced modes

- Video outputs: HD-SDI, DVI, HDMI, USB

- Internal storage ≥ 500 GB (to support long ERCP procedure recordings) plus external media

- DICOM 3.0, HL7 worklist, Gigabit Ethernet

- Compatible with fluoroscopy image overlay/side-by-side display where configured

Light Source

- High-intensity LED or Xenon

- Lamp life: LED ≥ 20,000 h; Xenon ≥ 500 h

- Color temperature approximately 5,500—6,500 K

- Multi-level adjustable brightness with auto exposure

Display Monitors (Dual)

- Two medical-grade LCD/LED monitors, each minimum 24-inch, Full HD

- Brightness ≥ 400 cd/m², contrast ≥ 1000:1, viewing angle ≥ 170°

- Multiple inputs (HD-SDI, DVI, HDMI, VGA)

- IEC 60601-1 compliant

Heavy-Duty Endoscopy Cart

- Reinforced medical-grade chassis rated for dual-monitor and processor/light-source load

- Isolation transformer ≥ 1500 VA

- Hospital-grade outlets, heavy-duty locking castors, dual-monitor articulating arms, cable management, scope hanger, keyboard tray

Electrical & Environmental

- Power: 100—240 VAC, 50/60 Hz

- Consumption: typical up to 1200 VA

- Operating: 10—40 °C, 30—85% RH non-condensing

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-2-18

- IEC 62304, 62366; ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR / US FDA 510(k)

- Compliance with current international duodenoscope reprocessing guidance (e.g., FDA/CDC recommendations on removable/disposable distal components)

## 4.3 Standard Accessories

a) Flexible side-viewing video duodenoscope with specialized elevator mechanism

- HD side-viewing tip, elevator control lever, therapeutic working channel ≥ 4.2 mm, working length ≥ 1240 mm, fully immersible

b) HD video processor console with ERCP recording capability

- Full HD processing, dedicated ERCP recording workflow, internal storage ≥ 500 GB, DICOM/HL7 compliant

c) High-intensity LED/Xenon endoscopic light source console

- Lamp life as specified, multi-level adjustable brightness, auto exposure

d) Dual medical-grade preview/live monitors on heavy-duty endoscopy cart

- Two 24-inch Full HD medical-grade monitors on articulating arms, IEC 60601-1 compliant

e) Reusable, detachable distal cap/tip for elevator mechanism cleaning access

- Removable/detachable distal cap or tip design permitting direct elevator-channel reprocessing

- Reusable and/or single-use variants as validated by manufacturer, minimum quantity per manufacturer's reprocessing protocol

- Compatible with hospital's high-level disinfection or sterilization workflow

f) Reusable autoclavable suction, air/water, and biopsy valves

- Silicone/elastomer, color-coded, autoclavable, minimum 5 complete sets

g) Water bottle assembly with dedicated air/water supply lines

- Autoclavable polycarbonate bottle ≥ 250 mL with silicone/PTFE supply lines and quick-connect fittings

h) Specialized elevator channel cleaning brush and main channel brush set

- Dedicated micro-brush for elevator channel plus double-ended brushes for main working channel; sized per scope specification; minimum 20 pieces

i) Automated watertight immersion leakage testing kit

- Automated pressure pump with calibrated gauge, scope-compatible adapter, audible/visual leak indication

j) Reusable adult bite blocks

- Autoclavable polymer, adult size, minimum 10 pieces

## 4.4 Installation Requirements

Electrical 100—240 VAC, dedicated 20 A circuit with proper earthing (earth resistance \< 1 Ω); UPS ≥ 2 kVA online recommended; HVAC 18—26 °C, 40—60% RH; RJ-45 network drop for PACS/HIS and integration with fluoroscopy/C-arm room; minimum 4 m² footprint within ERCP/fluoroscopy suite; provision for interface with fluoroscopy display where configured.

## 4.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), PM Schedule, Installation Report, Commissioning Report, DICOM Conformance Statement, and manufacturer's validated reprocessing instructions specific to the elevator mechanism and detachable distal component.

## 4.6 Training

Clinical training for ERCP endoscopists and endoscopy nursing/technical staff; biomedical engineering training on PM and troubleshooting; dedicated duodenoscope reprocessing training with special emphasis on elevator mechanism and detachable distal cap handling; application training on ERCP recording, DICOM/HL7 workflow, and fluoroscopy integration.

## 4.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, spare parts, PM (minimum two visits per year), breakdown maintenance, software updates, and technical support. Warranty shall include coverage of the elevator mechanism per international norms and supply of manufacturer's validated reprocessing consumables/detachable components as applicable during the warranty period.