TECHNICAL SPECIFICATIONS — GENERAL & SHARED HOSPITAL EQUIPMENT

Technical Tender Document

Table of Contents

1. MEDICAL GAS PIPELINE SYSTEM WITH OUTLETS	4

1.1 Clinical & Functional Specifications	4

1.2 Engineering Technical Specifications	4

1.3 Standard Accessories	5

1.4 Installation Requirements	6

1.5 Documentation	6

1.6 Training	6

1.7 Warranty	6

2. OPERATION THEATRE COMPLEX (CARDIAC OT)	8

2.1 Clinical & Functional Specifications	8

2.2 Engineering Technical Specifications	8

2.3 Standard Accessories	9

2.4 Installation Requirements	10

2.5 Documentation	10

2.6 Training	10

2.7 Warranty	11

3. DIAGNOSTIC SET (OTOSCOPE + OPHTHALMOSCOPE)	12

3.1 Clinical & Functional Specifications	12

3.2 Engineering Technical Specifications	12

3.3 Standard Accessories	12

3.4 Installation Requirements	13

3.5 Documentation	13

3.6 Training	13

3.7 Warranty	14

4. STETHOSCOPE (CARDIOLOGY GRADE)	15

4.1 Clinical & Functional Specifications	15

4.2 Engineering Technical Specifications	15

4.3 Standard Accessories	15

4.4 Installation Requirements	16

4.5 Documentation	16

4.6 Training	16

4.7 Warranty	16

5. BP APPARATUS	17

5.1 Clinical & Functional Specifications	17

5.2 Engineering Technical Specifications	17

5.3 Standard Accessories	17

5.4 Installation Requirements	18

5.5 Documentation	18

5.6 Training	18

5.7 Warranty	18

6. WEIGHT AND HEIGHT MACHINE	19

6.1 Clinical & Functional Specifications	19

6.2 Engineering Technical Specifications	19

6.3 Standard Accessories	19

6.4 Installation Requirements	20

6.5 Documentation	20

6.6 Training	20

6.7 Warranty	20

7. THERMOMETER (DIGITAL, INFRARED)	21

7.1 Clinical & Functional Specifications	21

7.2 Engineering Technical Specifications	21

7.3 Standard Accessories	21

7.4 Installation Requirements	22

7.5 Documentation	22

7.6 Training	22

7.7 Warranty	22

# 1. MEDICAL GAS PIPELINE SYSTEM WITH OUTLETS

Quantity: 2

## 1.1 Clinical & Functional Specifications

The Medical Gas Pipeline System (MGPS) shall be a centrally supplied, permanently installed distribution network delivering medical-grade gases (Oxygen, Nitrous Oxide, Medical Air 4-bar, Medical Vacuum, and Anaesthetic Gas Scavenging) to clinical points-of-use throughout the hospital including wards, ICUs, operation theatres, emergency, and recovery areas. The system shall be designed, manufactured, tested, installed, and commissioned in full compliance with HTM 02-01 / NFPA 99 / ISO 7396-1 to ensure uninterrupted, safe, and contamination-free delivery of gases at correct pressure, purity, and flow to every terminal outlet under all clinical loading conditions.

The system shall include gas source manifolds, pressure regulation, distribution pipeline, area/zone isolation, alarm monitoring, and terminal outlets. It shall provide continuous automatic changeover between duty, reserve, and emergency banks without pressure fluctuation, and shall permit isolation of any zone without disrupting the remainder of the network.

## 1.2 Engineering Technical Specifications

- Applicable Gases: Oxygen (O₂), Nitrous Oxide (N₂O), Medical Air 4-bar, Medical Vacuum, AGSS; provision for Carbon Dioxide and Instrument Air 7-bar where clinically indicated.

- Pipeline Material: Degreased seamless medical-grade copper conforming to EN 13348 / ASTM B819, minimum wall thickness suitable for working pressure with a safety factor of not less than 4.

- Jointing: Silver brazed under inert nitrogen purge; all joints traceable and pressure-tested.

- Working Pressure (Distribution): Oxygen / N₂O / Medical Air — nominal 4.0—4.5 bar (400—450 kPa); Vacuum — minimum −60 kPa (−450 mmHg or better).

- Manifold Design: Fully automatic duty/standby/reserve changeover; changeover shall not cause distribution pressure deviation exceeding ±10%.

- Pressure Regulation: Two-stage line regulators with lockable isolation; upstream and downstream pressure gauges; relief valves set to manufacturer/standard-specified values.

- Zone Valve Boxes: Wall-recessed, lockable, with break-glass access panel, gas-specific labelling, upstream and downstream pressure indication, and colour-coded ball valves.

- Area Alarm Panels: Microprocessor-based, minimum 4-gas capable, TFT/LED display, visual and audible alarms for high/low pressure deviation (typically ±20% of nominal), mute with automatic re-alarm, alarm history log, and volt-free contacts for BMS integration.

- Terminal Outlets: Gas-specific, non-interchangeable (pin-index or quick-connect type per purchaser-specified national standard — BS 5682 / DIN / NIST / AFNOR), self-sealing, minimum 100,000 insertion cycles, colour-coded and permanently labelled.

- Flow Capacity per Outlet (minimum): O₂ ≥ 100 L/min; Medical Air ≥ 200 L/min; Vacuum ≥ 40 L/min at rated pressure drop.

- Gas Purity: Meets European/US Pharmacopoeia requirements; particulate filtration ≤ 25 micron; moisture trap and inline coalescing filters at manifold.

- Alarm & Monitoring: Master, area, and local alarms as per HTM 02-01 / NFPA 99; networked monitoring with BMS/SCADA interface via RS-485/Modbus/Ethernet.

- Electrical Supply: 100—240 VAC, 50/60 Hz for alarm panels and manifold controls; UPS backup minimum 4 hours for alarms.

- Environmental: Operating temperature 5—40 °C; storage temperature −10 to 60 °C; relative humidity up to 90% non-condensing.

- Applicable Standards: ISO 7396-1, HTM 02-01, NFPA 99, EN 737 series, EN 13348, ASTM B819, IEC 60601-1 (for alarm panels), CE MDR / FDA clearance for terminal units.

## 1.3 Standard Accessories

a) Digital Area Alarm Panels

- Microprocessor-controlled, minimum 4.3-inch colour TFT/LED display.

- Simultaneous monitoring of minimum 4—7 gases, configurable.

- Audible alarm ≥ 65 dB at 1 m; visual LED indicators per gas.

- High/low pressure deviation alarms adjustable within ±10—30% of nominal.

- Non-volatile alarm event log (minimum 500 events).

- Battery backup minimum 1 hour on-board plus UPS input.

- Volt-free contacts and Ethernet/RS-485 for BMS integration.

b) Zone Valve Boxes

- Wall-recessed stainless steel or powder-coated MS enclosure.

- Break-glass or hinged access with tamper-evident feature.

- Full-bore quarter-turn ball valves, brass/DZR body, PTFE seats.

- Pre-installed pressure gauges upstream and downstream (0—10 bar for pressurised gases; −100 to 0 kPa for vacuum).

- Colour-coded and gas-specific labelling per national standard.

- Emergency shut-off lever accessible without tools once glass is broken.

c) High-Flow Pressure Regulator Valves

- Two-stage regulation for O₂, N₂O, and Medical Air manifolds.

- Inlet pressure up to 200 bar; outlet nominal 4.0—4.5 bar adjustable.

- Flow capacity minimum 3000 L/min at rated outlet pressure.

- Brass/chrome-plated body, oxygen-cleaned and degreased.

- Integral safety relief valve and pressure gauges.

d) Quick-Connect Terminal Outlet Blocks

- Gas-specific, non-interchangeable per purchaser's declared national standard.

- Self-sealing on disconnect; single-hand operation.

- Minimum 100,000 insertion/withdrawal cycles.

- First-fix and second-fix components with concealed rough-in.

- Colour-coded, permanently laser-etched gas identification.

e) Flexible High-Pressure Copper Pipeline Coils with Silver-Brazed Fittings

- Seamless copper to EN 13348 / ASTM B819, degreased and capped.

- Sizes 12 mm, 15 mm, 22 mm, 28 mm, 42 mm as required by flow calculations.

- Silver brazing alloy minimum 40% Ag; nitrogen purge during brazing.

f) Moisture Traps and Inline Particulate Filters

- Pre-installed at manifold outlet; filtration ≤ 25 micron particulate and ≤ 0.01 micron coalescing where required.

- Auto-drain moisture trap with sight glass.

- Serviceable elements with pressure differential indicator.

## 1.4 Installation Requirements

Structural provision for manifold room with adequate ventilation, fire rating, and separation per HTM 02-01/NFPA 99; dedicated electrical supply 100—240 VAC with UPS backup for alarm panels; earthing and equipotential bonding of all metallic pipeline components; concealed pipeline routing in slabs/walls with proper support brackets at maximum 2 m intervals; provision for expansion loops; test and purge points at every zone; identification labelling every 3 m and at all junctions; pressure testing at 1.5× working pressure minimum 24 hours; anti-confusion (cross-connection) testing prior to commissioning; particulate and purity testing per ISO 7396-1.

## 1.5 Documentation

User Manual, Installation & Service Manual, Parts Catalogue, As-built drawings (isometric and single-line), Pressure Test Certificates, Cross-Connection Test Report, Particulate and Purity Analysis Report, Electrical Safety Test Report, Commissioning Report, Preventive Maintenance Schedule, and Certificates of Conformity (ISO 7396-1 / HTM 02-01 / NFPA 99).

## 1.6 Training

On-site clinical user training for nursing and OT staff on outlet use and emergency shut-off; biomedical/facility engineering training on manifold changeover, alarm interpretation, filter and regulator servicing, leak detection, and pipeline expansion procedures; documented training records with attendance certificates.

## 1.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, spare parts, preventive maintenance visits (minimum 2 per year), breakdown maintenance, software/firmware updates for alarm panels, and 24×7 technical support. Availability of spare parts guaranteed for minimum 10 years.

# 2. OPERATION THEATRE COMPLEX (CARDIAC OT)

Quantity: 1

## 2.1 Clinical & Functional Specifications

The Cardiac Operation Theatre Complex shall be a fully integrated, turnkey ultra-clean surgical suite designed for open-heart, cardiothoracic, and complex cardiovascular procedures including CABG, valve replacement, congenital corrections, and hybrid cardiac interventions. The complex shall provide an ultra-clean (ISO Class 5 / Grade A under the surgical zone) environment with laminar airflow, integrated surgical audiovisual routing, environmental and equipment control from a central console, medical gas terminals, potential equalisation, and hermetic access — supporting continuous, safe, sterile, and ergonomically optimised surgical workflow for a full cardiac surgery team including surgeons, anaesthesiologists, perfusionists, and scrub personnel.

The complex shall support intraoperative imaging integration (C-arm, TEE, angiography), heart-lung machine, IABP, and ECMO operations, with all utility, gas, electrical, data, and video services delivered via pendants, ceiling supply units, and integrated wall panels without compromising sterility.

## 2.2 Engineering Technical Specifications

- Room Classification: ISO 14644-1 Class 5 (Grade A) under laminar flow zone; Class 7 (Grade B) periphery; ≥ 20 air changes per hour minimum; positive pressure ≥ 15 Pa relative to adjacent areas.

- Room Dimensions (typical): Minimum 55—65 m² floor area, minimum 3.0 m clear ceiling height.

- Wall & Ceiling Finish: Modular, hermetic, powder-coated GI or stainless steel panels with flush joints; antibacterial, non-porous, chemical-resistant finish; radius coving at all wall/floor/ceiling junctions.

- Flooring: Seamless conductive/anti-static homogeneous vinyl; resistance 25 kΩ—1 MΩ per IEC 61340; welded seams; radius covings.

- Laminar Airflow Enclosure: Ceiling-mounted, minimum 2.4 m × 2.4 m or 3.0 m × 3.0 m coverage; terminal HEPA filters H14 (≥ 99.995% at 0.3 μm); face velocity 0.20—0.30 m/s; integrated LED perimeter lighting.

- HVAC: 100% fresh air or minimum 20% fresh air with HEPA-filtered recirculation; temperature 18—24 °C adjustable; humidity 40—60% RH; independent AHU with N+1 redundancy.

- Surgical Video Routing Matrix: Minimum 8×8 (expandable to 16×16) 4K/UHD-capable video matrix over fibre-optic; supports HD-SDI, DVI, HDMI, DisplayPort, IP streaming; recording and live-broadcast capable.

- Control Console (Wall-Integrated): Colour touchscreen minimum 15-inch; controls for lighting, HVAC set-points, medical gas alarms, video routing, music, communication, and case timer/clock; user-selectable presets.

- Ambient LED Lighting: Multi-colour (RGB) tunable, dimmable 0—100%; wireless or wall-panel control.

- Medical Gas Outlets (per OT): Minimum — O₂ ×4, N₂O ×2, Medical Air ×2, Vacuum ×4, AGSS ×2, delivered via ceiling pendants and wall panels.

- Potential Equalisation: Central equipotential busbar with minimum 12 stud terminals; earth resistance ≤ 0.1 Ω between any two exposed conductive parts.

- Doors: Hermetically sealed, motorised sliding, minimum 1400 mm clear opening, with hands-free foot switch, elbow switch, and touchless sensor; interlocking with adjacent scrub/prep areas.

- Electrical: Isolated Power System (IPS) with Line Isolation Monitor per IEC 60364-7-710; minimum 8 kVA per OT; UPS backup for critical circuits ≥ 30 minutes; emergency generator changeover ≤ 10 seconds.

- Data & Communication: Structured cabling Cat 6A minimum; intercom to control room, ICU, and perfusion; DICOM/HL7 integration capability.

- Applicable Standards: ISO 14644-1/2, HTM 03-01, ASHRAE 170, DIN 1946-4, IEC 60364-7-710, IEC 60601-1, ISO 14971, CE MDR.

## 2.3 Standard Accessories

a) Laminar Airflow Ceiling Enclosure with HEPA Filtration

- Coverage area minimum 2.4 × 2.4 m; unidirectional downflow.

- Terminal HEPA H14 modules, DOP/PAO-testable, factory-certified ≥ 99.995% @ 0.3 μm.

- Face velocity 0.20—0.30 m/s adjustable.

- Integrated dimmable LED perimeter lighting, colour temperature 3500—5000 K.

- Pressure differential indicator across filters.

b) Surgical Matrix Video Routing Switcher with Fibre-Optic Connections

- Minimum 8 inputs × 8 outputs, expandable.

- Native 4K/UHD support; backward compatible HD-SDI/DVI/HDMI.

- Fibre-optic transport minimum OM3, distances up to 300 m without repeaters.

- Recording (H.264/H.265) to internal storage ≥ 1 TB and network storage.

- Live streaming to auditorium/telemedicine per user configuration.

c) Wall-Integrated Surgeon Control Console

- Minimum 15-inch capacitive colour touchscreen, resolution ≥ 1280 × 800.

- Controls: lighting, HVAC set-points, humidity, differential pressure display, medical gas alarm status, video routing, music, intercom, case clock, and elapsed timer.

- Multi-user profile with password protection.

- Hermetic flush-mount stainless steel bezel.

d) Peripheral LED Ambient Room Lights with Wireless Dimmer

- RGB tunable + white; CRI ≥ 80.

- 0—100% smooth dimming.

- Wireless remote and console-integrated control.

- IP54 minimum; suitable for OT cleaning protocols.

e) Medical Gas Outlet Panel Manifolds (Flush-Mounted)

- Gas-specific quick-connect outlets to national standard.

- Stainless steel frontplate, flush wall-mount.

- Self-sealing, colour-coded, minimum 100,000 cycles.

f) Potential Equalisation Grounding Busbars

- Copper busbar minimum 30 × 5 mm cross-section.

- Minimum 12 stud terminals, brass/copper.

- Continuous impedance ≤ 0.1 Ω to main earth bar.

- Labelled and accessible for periodic testing.

g) Hermetically Sealed Motorised Sliding Doors with Foot-Switch

- Clear opening minimum 1400 mm width × 2100 mm height.

- Automatic operation via foot switch, elbow switch, and touchless sensor.

- Manual override; anti-crush safety sensors.

- Powder-coated steel or stainless steel with viewing window (lead-lined optional).

- Hermetic gasket seal; sound insulation ≥ 30 dB.

## 2.4 Installation Requirements

Turnkey civil, mechanical, electrical, HVAC, medical gas, and IT works per approved layout; dedicated AHU room with redundancy; independent isolated power panel with LIM; UPS and generator backup; medical gas connections from central manifold; structured cabling backbone; fire detection and suppression per NFPA 99/HTM 05-02; radiation shielding (lead-lined walls/doors) if hybrid imaging installed; validation and certification of cleanroom class, laminar flow velocity, HEPA integrity (DOP/PAO), air-change rate, pressure cascade, temperature/humidity, and noise level prior to handover.

## 2.5 Documentation

Complete as-built architectural, MEP, HVAC, medical gas, and electrical drawings; O&M manuals for every subsystem; HEPA certification reports; cleanroom classification report (ISO 14644-1); IPS/LIM test report; earth continuity and equipotential test report; commissioning report; validation protocols (IQ/OQ/PQ); preventive maintenance schedules; and CE MDR / declarations of conformity for all integrated medical devices.

## 2.6 Training

Clinical training for surgical, anaesthesia, perfusion, and nursing teams on video routing, console operation, and workflow; biomedical/facility engineering training on HVAC, LAF, HEPA change-out, IPS/LIM, door mechanics, and integration systems; troubleshooting and emergency procedures; documented certification.

## 2.7 Warranty

Minimum two (2) years comprehensive warranty on all integrated subsystems including labour, spare parts, preventive maintenance (minimum 4 visits/year), breakdown maintenance, software/firmware updates, and 24×7 technical support with defined response and resolution SLAs. HEPA filter first replacement included. Spare parts guaranteed for minimum 10 years.

# 3. DIAGNOSTIC SET (OTOSCOPE + OPHTHALMOSCOPE)

Quantity: 10

## 3.1 Clinical & Functional Specifications

The Diagnostic Set shall be a portable, handheld, rechargeable combination instrument for routine physical examination of the ear (external auditory canal and tympanic membrane) and the eye (anterior segment, media, and fundus). The set shall support outpatient clinics, wards, emergency, and pediatric use, providing clear, high-intensity, colour-true illumination and optical magnification suitable for accurate visualisation and diagnosis by physicians, ENT/ophthalmology specialists, and trainees.

## 3.2 Engineering Technical Specifications

- Illumination Source: High-intensity white LED, colour temperature 3800—5500 K, CRI ≥ 90, rated life ≥ 20,000 hours.

- Illumination Intensity: Otoscope ≥ 20,000 lux at distal tip; Ophthalmoscope ≥ 15,000 lux at working distance.

- Otoscope Head: Fibre-optic illumination path (distal, non-obstructing); sealed pneumatic port for tympanic mobility testing; ≥ 3× magnifying viewing lens with anti-reflective coating; swing-away or removable lens.

- Ophthalmoscope Head: Minimum 5 apertures (large spot, small spot, micro, fixation/graticule, slit, red-free filter); minimum 18—20 corrective diopter lens wheel (typically −20 D to +20 D in single-diopter steps).

- Handle: Rechargeable lithium-ion battery, minimum 3.5 V DC output, knurled/textured anti-slip grip; interchangeable between otoscope and ophthalmoscope heads.

- Battery Runtime: Minimum 60 minutes continuous use on full charge.

- Charging Time: Maximum 4 hours to full charge.

- Charger: Desktop dual-slot intelligent charging dock, LED status indication, overcharge and short-circuit protection.

- Power Supply: 100—240 VAC, 50/60 Hz, medical-grade adapter.

- Weight (per head + handle): Approximately 200—400 g.

- Environmental: Operating 10—40 °C; storage −10 to 55 °C; humidity up to 90% non-condensing.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), ISO 10993 (biocompatibility of skin/patient-contact parts), CE MDR / FDA clearance.

## 3.3 Standard Accessories

a) Otoscope Optical Head with LED Fibre-Optic Illumination

- Distal fibre-optic illumination; no direct-view lamp obstruction.

- Sealed pneumatic port with rubber bulb attachment point.

- ≥ 3× magnification; anti-reflective viewing lens.

- Autoclavable or wipe-disinfect compatible per manufacturer.

b) Ophthalmoscope Optical Head with Multi-Aperture and Filters

- Minimum 5 apertures including fixation graticule.

- Red-free and cobalt/polarising filters.

- Corrective lens wheel ≥ ±20 D in 1 D steps.

- Anti-glare, anti-fog viewing window.

c) Reusable Polypropylene Ear Specula — Starter Kit

- Sizes 2.5 mm, 3 mm, 4 mm, 5 mm minimum.

- Autoclavable to 121 °C or chemical-disinfect compatible.

- Smooth, non-traumatic edges; colour-coded by size.

d) Rechargeable Lithium-Ion Battery Handle

- Nominal 3.5 V DC; capacity ≥ 1200 mAh.

- Knurled/textured anti-slip grip.

- Protection against overcharge, over-discharge, short circuit.

- Minimum 500 charge cycles to 80% capacity retention.

e) Desktop Twin-Slot Intelligent Charging Dock

- Simultaneous charge of two handles.

- Auto-detect battery state; trickle charge on completion.

- LED indication per slot (charging / full / fault).

- Medical-grade power adapter, 100—240 VAC input.

f) Rigid Protective Storage Case

- Form-fitted foam inserts for all components.

- Impact- and moisture-resistant ABS/polypropylene shell.

- Latching closure with carry handle.

## 3.4 Installation Requirements

Tabletop use; standard mains socket 100—240 VAC for charging dock; no special environmental controls beyond normal clinical use.

## 3.5 Documentation

User Manual, Quick-Reference Card, Cleaning & Disinfection Guide, Warranty Certificate, Declaration of Conformity (CE MDR / FDA), Electrical Safety Test Report.

## 3.6 Training

On-site clinical demonstration and hands-on training for medical officers, residents, and nursing staff on head interchange, aperture/filter selection, pneumatic otoscopy, battery care, cleaning, and troubleshooting.

## 3.7 Warranty

Minimum two (2) years comprehensive warranty including labour, spare parts (excluding consumables), battery replacement if capacity falls below 70% within warranty period, and technical support. Spare parts availability minimum 7 years.

# 4. STETHOSCOPE (CARDIOLOGY GRADE)

Quantity: 26

## 4.1 Clinical & Functional Specifications

The Cardiology-Grade Stethoscope shall be a high-acoustic-performance dual-head acoustic stethoscope designed for detailed auscultation of cardiac, pulmonary, and vascular sounds across the full clinically relevant frequency spectrum. It shall support use by cardiologists, internists, anaesthesiologists, and senior residents in wards, OPD, ICU, and cardiac care areas, delivering superior sound clarity and fidelity for detection of subtle murmurs, gallops, bruits, and adventitious breath sounds in adult and pediatric patients.

## 4.2 Engineering Technical Specifications

- Chestpiece: Stainless steel dual-sided (adult diaphragm + pediatric diaphragm or open bell), machined finish, non-chill rim.

- Diaphragm Technology: Tunable single-diaphragm allowing selection between low- and high-frequency sounds by contact pressure variation.

- Frequency Response Range: Minimum 20 Hz — 1000 Hz effective auscultation range.

- Tubing: Dual-lumen (two-in-one) heavy-walled PVC or thermoplastic, latex-free; internal dual sound channels to preserve stereo acoustic separation.

- Tubing Length: Approximately 68—74 cm (27—29 in).

- Headset: Anatomically angled (approximately 15°) stainless steel binaural, adjustable tension.

- Eartips: Soft-sealing snap-tight eartips in two sizes/types; latex-free hypoallergenic material.

- Chestpiece Weight: Approximately 150—200 g.

- Total Weight: Approximately 170—220 g.

- Non-Chill Rim: Silicone or equivalent, patient-comfort.

- Applicable Standards: ISO 10993 (biocompatibility), CE MDR / FDA clearance, RoHS.

## 4.3 Standard Accessories

a) Dual-Lumen Heavy-Walled Acoustic Tubing with Non-Chill Coating

- Two parallel internal channels for left/right acoustic isolation.

- Latex-free thermoplastic, resistant to skin oils and alcohol wipes.

- Length approximately 68—74 cm.

- Non-chill exterior coating.

b) Stainless Steel Dual-Head Chestpiece with Tunable Adult Diaphragm

- Machined stainless steel body.

- Tunable diaphragm — low-frequency (bell mode) via light contact; high-frequency (diaphragm mode) via firm contact.

- Adult side diameter approximately 45—48 mm.

c) Pediatric Tunable Diaphragm or Convertible Open Bell

- Pediatric side diameter approximately 30—35 mm.

- Interchangeable with open bell configuration.

- Non-chill rim included.

d) Soft-Sealing Ergonomic Eartips (1 pair)

- Snap-tight secure attachment.

- Soft silicone or equivalent, hypoallergenic, latex-free.

e) Firm Seating Replacement Eartips (1 pair)

- Firmer durometer for tight seal.

- Compatible with same binaural stems.

f) Microfibre Cleaning Cloth

- Lint-free, suitable for chestpiece and tubing.

## 4.4 Installation Requirements

None. Ready-to-use portable clinical instrument.

## 4.5 Documentation

User Manual, Cleaning & Disinfection Instructions, Warranty Certificate, Declaration of Conformity, Biocompatibility Statement.

## 4.6 Training

Brief clinical orientation on tunable diaphragm technique, eartip fitting, and care/cleaning. No technical training required.

## 4.7 Warranty

Minimum two (2) years comprehensive warranty against manufacturing defects, including free replacement of tubing, diaphragm, and eartips within warranty period. Spare parts availability minimum 5 years.

# 5. BP APPARATUS

Quantity: 23

## 5.1 Clinical & Functional Specifications

The Blood Pressure Apparatus shall be a mobile, mechanical (aneroid) sphygmomanometer designed for accurate non-invasive measurement of systolic and diastolic blood pressure by auscultatory technique in adult inpatients across wards, OPD, and emergency areas. The apparatus shall be robust, portable on a mobile stand, easy to read, calibration-stable, and suitable for continuous clinical use.

## 5.2 Engineering Technical Specifications

- Manometer Type: Aneroid dial, high-visibility clockface, minimum 145—150 mm scale diameter.

- Measurement Range: 0—300 mmHg.

- Graduation: 2 mmHg increments.

- Accuracy: Within ±3 mmHg across full range at 18—33 °C.

- Dial Cover: Shatter-resistant, anti-glare glass or polycarbonate.

- Movement: Copper-beryllium or equivalent long-life mechanism; minimum 10,000 cycle rated without recalibration drift beyond specification.

- Inflation System: Latex-free rubber bulb with metal thumbscrew air release valve; smooth, controlled deflation.

- Cuff Material: Reusable nylon or equivalent, latex-free, with hook-and-loop closure, marked artery position and range indicator.

- Cuff Sizes Supplied: Adult (standard) and Large Adult minimum.

- Bladder: Two-tube configuration, latex-free.

- Mobile Stand: Powder-coated steel or stainless steel; 5-caster anti-static base with minimum 2 lockable casters; storage basket for cuff/accessories.

- Height (Stand): Adjustable or fixed approximately 120—140 cm working height.

- Weight (Complete Unit): Approximately 5—8 kg.

- Environmental: Operating 10—40 °C; storage −20 to 70 °C; RH up to 85% non-condensing.

- Applicable Standards: ISO 81060-1 (mechanical sphygmomanometers), IEC 60601-1 (where applicable), ISO 10993, EN 1060 series, CE MDR / FDA clearance.

## 5.3 Standard Accessories

a) Heavy-Duty Mobile Pole Stand with Anti-Static Casters and Storage Basket

- 5-caster base, minimum 2 lockable; anti-static wheels.

- Powder-coated MS or stainless steel column.

- Wire storage basket for cuffs and bulb.

- Load capacity ≥ 5 kg at manometer mount.

b) Large High-Visibility Clockface Dial Manometer

- Scale diameter ≥ 145 mm; graduated 0—300 mmHg in 2 mmHg steps.

- Anti-glare shatter-resistant cover.

- Calibration accuracy ±3 mmHg.

c) Reusable Adult Nylon Cuff (Standard Medium)

- Latex-free nylon.

- Hook-and-loop closure, marked artery position.

- Range indicator, size marking.

- Fits arm circumference approximately 25—35 cm.

d) Reusable Large Adult Nylon Cuff (Large)

- Fits arm circumference approximately 34—50 cm.

- Same construction as standard cuff.

e) Heavy-Duty Latex-Free Rubber Inflation Bulb with Air-Release Screw Valve

- Latex-free, ergonomic squeeze bulb.

- Metal thumbscrew deflation valve — smooth, controllable release.

- Check-valve for one-way inflation.

f) Coiled Flexible Twin-Core Air Connecting Hoses

- Latex-free, coiled, kink-resistant.

- Quick-click metal couplers, compatible with cuff and manometer.

- Length approximately 60—90 cm extended.

## 5.4 Installation Requirements

None beyond assembly of mobile stand. Store in normal clinical environment.

## 5.5 Documentation

User Manual, Calibration Certificate (factory), Cleaning & Disinfection Instructions, Warranty Certificate, Declaration of Conformity (ISO 81060-1 / CE MDR).

## 5.6 Training

Clinical training for nursing staff on correct cuff selection, placement, auscultatory technique, deflation rate, and periodic calibration checks. Biomedical training on annual calibration verification.

## 5.7 Warranty

Minimum two (2) years comprehensive warranty covering manometer calibration, cuff, bulb, and hose, including annual calibration verification within warranty. Spare parts availability minimum 7 years.

# 6. WEIGHT AND HEIGHT MACHINE

Quantity: 11

## 6.1 Clinical & Functional Specifications

The Weight and Height Machine shall be a robust, floor-standing combined anthropometric measurement device providing accurate simultaneous or independent measurement of body weight and standing height for adult and pediatric ambulatory patients in OPD, wards, admission, and screening areas. The device shall support routine clinical assessment, BMI calculation, growth monitoring, and admission workflow with rapid, repeatable, and easily readable results.

## 6.2 Engineering Technical Specifications

- Weighing Technology: Digital electronic load cell (strain gauge), 4-point or equivalent.

- Weight Measurement Range: Minimum 2—200 kg; extended capacity 250 kg preferred.

- Weight Graduation: 100 g (0.1 kg) or better.

- Weight Accuracy: Within ±0.1% of reading or ±100 g, whichever is greater; conforms to Class III per OIML R76 / NAWI Directive.

- Height Measurement: Telescopic mechanical or digital rod, graduated in mm.

- Height Range: Minimum 60—200 cm; extended to 210 cm preferred.

- Height Accuracy: Within ±5 mm.

- Display: Digital LCD, minimum 25 mm digit height, backlit, mounted at eye level on column.

- Functions: Zero/tare, hold, unit selection (kg/lb), BMI calculation (with height input), auto power-off.

- Platform: Cast-iron or heavy-gauge steel, non-slip textured rubber mat, minimum 300 × 300 mm.

- Column: Powder-coated steel; adjustable/telescopic height rod.

- Power Supply: 100—240 VAC, 50/60 Hz medical-grade adapter; battery backup (rechargeable or standard cells) minimum 20 hours continuous use.

- Weight (Device): Approximately 15—25 kg.

- Environmental: Operating 5—40 °C; RH up to 85% non-condensing.

- Applicable Standards: OIML R76 Class III, EN 45501, IEC 60601-1, IEC 60601-1-2 (EMC), CE MDR / FDA clearance.

## 6.3 Standard Accessories

a) Heavy-Duty Cast-Iron Platform Base Deck with Non-Slip Rubber Mat

- Cast-iron or heavy-gauge steel platform.

- Textured non-slip rubber mat, easy-clean.

- Platform dimensions minimum 300 × 300 mm.

- Load capacity ≥ 250 kg.

b) Telescopic Height-Measuring Rod (Mechanical or Digital)

- Range 60—200 cm minimum.

- Graduated in mm, high-contrast markings.

- Locking head-piece with clear pointer.

- Foldable/retractable for storage.

c) High-Visibility Digital LCD Console

- Backlit LCD, digit height ≥ 25 mm.

- Eye-level column mounting.

- Displays weight, height, BMI, unit indicator.

d) Electronic Tare and Calibration Function (Software)

- Zero/tare function.

- Field calibration menu (password-protected).

- Auto-zero drift compensation.

- BMI computation from entered height.

e) AC Power Adapter with Medical-Grade Cord

- 100—240 VAC input, low-voltage DC output.

- Medical-grade IEC 60601-1 compliant.

- Detachable earthed power cord to national standard.

## 6.4 Installation Requirements

Level, hard floor surface; standard mains socket; adequate wall clearance for height rod extension; calibration verification prior to first use.

## 6.5 Documentation

User Manual, Calibration Certificate (factory), Verification Certificate (where required by local metrology law), Warranty Certificate, Declaration of Conformity (OIML / CE MDR), Electrical Safety Test Report.

## 6.6 Training

Clinical staff training on weighing procedure, height measurement, BMI use, unit switching, and daily zero-check. Biomedical training on calibration verification and battery replacement.

## 6.7 Warranty

Minimum two (2) years comprehensive warranty including load cell, display, power adapter, labour, and one annual calibration verification. Spare parts availability minimum 7 years.

# 7. THERMOMETER (DIGITAL, INFRARED)

Quantity: 46

## 7.1 Clinical & Functional Specifications

The Digital Infrared Thermometer shall be a handheld, non-contact clinical thermometer for rapid measurement of body temperature via forehead (temporal) infrared emission, suitable for screening and routine temperature monitoring in OPD, wards, emergency, pediatric, and infection-control areas. The device shall provide fast (≤ 1 second) measurement, high accuracy, audible fever alert, and hygienic non-contact operation to minimise cross-contamination risk.

## 7.2 Engineering Technical Specifications

- Measurement Method: Non-contact infrared, forehead/temporal.

- Measurement Range (Body Mode): 32.0 — 43.0 °C (89.6 — 109.4 °F).

- Measurement Range (Object/Surface Mode): 0 — 100 °C, selectable.

- Accuracy: Within ±0.2 °C (±0.4 °F) over 35.0—42.0 °C; ±0.3 °C elsewhere in range.

- Resolution: 0.1 °C / 0.1 °F.

- Response Time: ≤ 1 second.

- Measurement Distance: 3—8 cm (typical) from forehead.

- Display: Backlit LCD, minimum 3-colour (green/yellow/red) fever-indication backlight.

- Memory: Minimum 30 last readings recall.

- Fever Alert: Audible buzzer + colour-coded backlight above configurable threshold (typically ≥ 37.5 °C).

- Unit Selection: °C / °F user-selectable.

- Auto Power-Off: Within 30 seconds of inactivity.

- Battery: 2× AAA alkaline or rechargeable Li-ion; ≥ 3000 measurements per battery set.

- Environmental (Operating): 10—40 °C, RH ≤ 85% non-condensing.

- Environmental (Storage): −20 to 55 °C.

- Weight: Approximately 150—250 g including batteries.

- Applicable Standards: ASTM E1965, ISO 80601-2-56, IEC 60601-1, IEC 60601-1-2 (EMC), CE MDR / FDA clearance.

## 7.3 Standard Accessories

a) Ergonomic Pistol-Grip Housing with Backlit LCD

- ABS/polycarbonate housing, ergonomic grip, single-hand operation.

- Backlit LCD, minimum 3-colour fever backlight.

- Trigger-style measurement button.

b) Integrated Acoustic Buzzer for Fever Alert

- Audible ≥ 60 dB at 30 cm.

- Distinct fever-alert tone above configurable threshold.

- Mute option user-selectable.

c) Removable Protective Optical Lens Cover Caps

- Snap-fit or slide protective cover.

- Prevents lens contamination between uses.

d) Rechargeable Li-Ion or Alkaline AAA Battery Pack

- Standard 2× AAA alkaline or equivalent rechargeable set.

- Battery life ≥ 3000 measurements per full set.

e) Wrist Strap Lanyard and Fabric Storage Pouch

- Textile wrist strap for drop-prevention.

- Soft fabric slip pouch for transport and storage.

## 7.4 Installation Requirements

None. Ready-to-use portable handheld device.

## 7.5 Documentation

User Manual, Quick-Start Guide, Calibration Certificate (factory), Cleaning & Disinfection Instructions, Warranty Certificate, Declaration of Conformity (ISO 80601-2-56 / CE MDR).

## 7.6 Training

Brief clinical orientation on correct measurement distance, forehead technique, unit switching, cleaning, and battery replacement.

## 7.7 Warranty

Minimum two (2) years comprehensive warranty covering electronics, sensor, and display; free replacement of unit if accuracy drift exceeds specification during warranty period. Spare parts availability minimum 5 years.