TECHNICAL TENDER SPECIFICATIONS — IMAGING & RADIOLOGY EQUIPMENT

Technical Tender Document

Table of Contents

1. CT SCANNER — 256-SLICE MULTI-DETECTOR WITH WORKSTATION	6

1.1 Clinical & Functional Specifications	6

1.2 Engineering Technical Specifications	6

1.3 Standard Accessories	9

1.4 Installation Requirements	10

1.5 Documentation	10

1.6 Training	11

1.7 Warranty	11

2. 1.5T MRI MACHINE	12

2.1 Clinical & Functional Specifications	12

2.2 Engineering Technical Specifications	12

2.3 Standard Accessories	14

2.4 Installation Requirements	16

2.5 Documentation	16

2.6 Training	16

2.7 Warranty	17

3. MONO PLAN ANGIOGRAPHY SYSTEM (with radiation glasses, radiation sheets, storage rack, procedure stool)	18

3.1 Clinical & Functional Specifications	18

3.2 Engineering Technical Specifications	18

3.3 Standard Accessories	20

3.4 Installation Requirements	21

3.5 Documentation	22

3.6 Training	22

3.7 Warranty	22

4. FLUOROSCOPY UNIT (R/F SYSTEM WITH IMAGE INTENSIFIER)	23

4.1 Clinical & Functional Specifications	23

4.2 Engineering Technical Specifications	23

4.3 Standard Accessories	24

4.4 Installation Requirements	25

4.5 Documentation	26

4.6 Training	26

4.7 Warranty	26

5. DIGITAL FIXED X-RAY MACHINE (high-frequency, 500—650 mA, overhead tube suspension)	27

5.1 Clinical & Functional Specifications	27

5.2 Engineering Technical Specifications	27

5.3 Standard Accessories	29

5.4 Installation Requirements	29

5.5 Documentation	30

5.6 Training	30

5.7 Warranty	30

6. MOBILE / PORTABLE X-RAY MACHINE	31

6.1 Clinical & Functional Specifications	31

6.2 Engineering Technical Specifications	31

6.3 Standard Accessories	33

6.4 Installation Requirements	33

6.5 Documentation	34

6.6 Training	34

6.7 Warranty	34

7. FULL-FIELD DIGITAL MAMMOGRAPHY UNIT	35

7.1 Clinical & Functional Specifications	35

7.2 Engineering Technical Specifications	35

7.3 Standard Accessories	37

7.4 Installation Requirements	38

7.5 Documentation	38

7.6 Training	38

7.7 Warranty	38

8. DEXA SCANNER (Bone Mineral Densitometry)	40

8.1 Clinical & Functional Specifications	40

8.2 Engineering Technical Specifications	40

8.3 Standard Accessories	41

8.4 Installation Requirements	42

8.5 Documentation	42

8.6 Training	43

8.7 Warranty	43

9. C-ARM IMAGE INTENSIFIER WITH DUAL MONITORS	44

9.1 Clinical & Functional Specifications	44

9.2 Engineering Technical Specifications	44

9.3 Standard Accessories	46

9.4 Installation Requirements	46

9.5 Documentation	46

9.6 Training	47

9.7 Warranty	47

10. POWER CONTRAST INJECTOR (CT-COMPATIBLE)	48

10.1 Clinical & Functional Specifications	48

10.2 Engineering Technical Specifications	48

10.3 Standard Accessories	50

10.4 Installation Requirements	51

10.5 Documentation	51

10.6 Training	51

10.7 Warranty	51

11. COMPUTED / DIGITAL RADIOGRAPHY (CR / DR) SYSTEM	52

11.1 Clinical & Functional Specifications	52

11.2 Engineering Technical Specifications	52

11.3 Standard Accessories	53

11.4 Installation Requirements	54

11.5 Documentation	54

11.6 Training	54

11.7 Warranty	55

12. DICOM VIEWING / DIAGNOSTIC REPORTING WORKSTATION	56

12.1 Clinical & Functional Specifications	56

12.2 Engineering Technical Specifications	56

12.3 Standard Accessories	58

12.4 Installation Requirements	58

12.5 Documentation	59

12.6 Training	59

12.7 Warranty	59

13. FILM VIEWER BOX / DIGITAL VIEWING STATION	60

13.1 Clinical & Functional Specifications	60

13.2 Engineering Technical Specifications	60

13.3 Standard Accessories	61

13.4 Installation Requirements	62

13.5 Documentation	62

13.6 Training	62

13.7 Warranty	62

# 1. CT SCANNER — 256-SLICE MULTI-DETECTOR WITH WORKSTATION

Quantity: 1

## 1.1 Clinical & Functional Specifications

The system shall be a state-of-the-art, whole-body, multi-detector row Computed Tomography (MDCT) scanner capable of acquiring a minimum of 256 slices per gantry rotation, intended for high-resolution diagnostic imaging across all clinical applications including routine head, chest, abdomen, pelvis, musculoskeletal, cardiac, coronary CT angiography (CCTA), neuro-perfusion, whole-body CT angiography, oncology staging, pediatric imaging, trauma imaging, interventional guidance, and radiotherapy planning.

The system shall provide isotropic sub-millimetric resolution, high temporal resolution suitable for cardiac imaging without beta-blocker preparation in typical adult patients, and shall incorporate advanced iterative or deep-learning-based reconstruction technology to enable low-dose scanning in compliance with ALARA principles.

The equipment shall be delivered as a complete turn-key system comprising gantry, patient table, high-voltage generator, X-ray tube, detector array, operator console, advanced diagnostic post-processing workstation, isolation transformer, and all required accessories, cabling, and interconnects for full clinical operation.

The system shall support automated workflow including patient registration, protocol selection, automatic exposure control, automatic tube-current modulation, automatic organ-based dose modulation, and DICOM transmission to PACS/RIS/HIS with full connectivity.

## 1.2 Engineering Technical Specifications

Gantry

- Aperture (bore) diameter: minimum 78 cm

- Gantry tilt: minimum ±30° (motorized)

- Gantry rotation speed: minimum 0.30 seconds per full 360° rotation, or faster

- Scan field of view (SFOV): minimum 50 cm

- Display field of view (DFOV): selectable, minimum 5 cm to full SFOV

- Laser patient positioning markers: internal and external

Detector System

- Detector type: solid-state, high-efficiency ceramic scintillator with photodiode array or equivalent

- Number of simultaneously acquired slices per rotation: minimum 256

- Minimum slice thickness: 0.625 mm or thinner

- Detector coverage at isocenter (z-axis): minimum 80 mm per rotation

- Detector rows: configured to provide at least 256 simultaneous data acquisition channels

X-Ray Tube

- Anode heat storage capacity: minimum 8 MHU (equivalent effective cooling)

- Anode heat dissipation rate: minimum 900 kHU/min

- Focal spot sizes: minimum two selectable (small and large), typically ≤0.7 mm and ≤1.2 mm

- Tube voltage (kV): selectable, minimum range 70—140 kV in configurable steps

- Dual-energy / multi-energy scanning capability: shall be supported (via kV switching, dual-source, or spectral detector) — bidders shall declare method

- Tube current (mA): minimum range 10—800 mA

High-Voltage Generator

- Generator power output: minimum 72 kW

- Waveform: high-frequency, closed-loop regulated

Patient Table

- Type: motorized, horizontal, longitudinal and vertical movement

- Maximum patient weight capacity: minimum 220 kg

- Table material: low-attenuation carbon fiber tabletop

- Longitudinal scan range: minimum 170 cm

- Vertical travel: adjustable for patient loading

- Positioning accuracy: within ±0.25 mm

Image Acquisition and Reconstruction

- Reconstruction matrix: minimum 512 × 512; 1024 × 1024 configurable

- Reconstruction speed: minimum 40 images per second at 512 × 512

- Iterative and/or AI/deep-learning-based reconstruction: mandatory, with configurable strength levels

- Spatial resolution (high-contrast): minimum 18 lp/cm at 0% MTF

- Low-contrast resolution: minimum 3 mm at 0.3% contrast at standard dose

- Temporal resolution (cardiac): equivalent to or better than 175 ms in single-segment mode

Dose Management

- Automatic tube-current modulation (angular and longitudinal): required

- Organ-based dose modulation: required

- Automatic kV selection based on patient size: required

- CTDI and DLP display and reporting: required, DICOM structured dose report compliant

Operator Console

- Dual medical-grade color LCD/LED monitors, minimum 24-inch diagonal each

- Monitor resolution: minimum 1920 × 1200 or better

- CPU: multi-core 64-bit processor, minimum 16 GB RAM, minimum 2 TB HDD/SSD storage

- Raw data storage: minimum 500,000 images

- Operating system: current-generation medical-grade OS

- DICOM 3.0 compliant: Send, Receive, Print, Query/Retrieve, Storage Commitment, MPPS, Worklist

Advanced Post-Processing Diagnostic Workstation

- Separate independent workstation, 64-bit multi-core processor

- Minimum 32 GB RAM, minimum 4 TB storage

- Dual medical-grade diagnostic displays, minimum 24-inch, resolution minimum 1920 × 1200

- Advanced clinical software packages: MPR, MIP, 3D volume rendering, virtual endoscopy, vessel analysis, cardiac analysis with calcium scoring and CCTA, neuro-perfusion, colonography, dental, lung nodule analysis, oncology follow-up, dual-energy analysis (if applicable)

- Full DICOM connectivity

Cardiac Imaging Capability

- Prospective and retrospective ECG gating

- Prospective ECG-triggered axial cardiac acquisition

- Suitable for coronary CT angiography at elevated heart rates

Connectivity & Data Management

- DICOM 3.0 full compliance

- HL7 interface for RIS/HIS integration

- IHE profiles compliance (Scheduled Workflow, Consistent Presentation of Images)

- Ethernet minimum 1 Gbps

- USB 3.0 and DVD-RW media output for DICOM export

Electrical Specifications

- Input power supply: 380—415 VAC, 3-phase, 50/60 Hz (or as per site conditions)

- Isolation transformer included

- Total connected load: as declared by supplier; typically ≤150 kVA

- UPS support for console and workstation

Environmental Conditions

- Operating temperature: 18—24°C

- Operating relative humidity: 30—70%, non-condensing

- Heat dissipation: shall be declared by supplier for HVAC sizing

Applicable Standards

- IEC 60601-1 (General safety)

- IEC 60601-1-2 (EMC)

- IEC 60601-2-44 (Particular requirements for CT safety)

- IEC 60601-1-3 (Radiation protection)

- IEC 62304 (Medical device software)

- IEC 62366 (Usability)

- ISO 13485 (Manufacturer QMS)

- ISO 14971 (Risk management)

- DICOM 3.0

- HL7

- CE MDR or US FDA 510(k) clearance

## 1.3 Standard Accessories

Patient Examination Couch (motorized)

- Motorized vertical and longitudinal movement

- Carbon fiber, low-attenuation tabletop

- Minimum 220 kg patient weight capacity

- Emergency stop buttons on both sides

- Integrated with gantry and console controls

Head Holder and Carbon Fiber Table Extension

- Radiolucent carbon fiber construction

- Compatible with table quick-lock interface

- Adjustable head immobilization straps

- Cleaning/disinfection compatible

Patient Positioning Pad Set, Restraint Straps, and Cushions

- Radiolucent, low-attenuation foam material

- Wipe-clean, disinfectant-compatible surface

- Complete set covering head, arm, knee, and foot positioning

- Adjustable restraint straps with quick-release buckles

Internal Intercom System and Patient Communication Module

- Two-way voice communication between operator and gantry room

- Pre-recorded breath-hold instructions in multiple languages (configurable)

- Volume-adjustable, hands-free operation

Integrated Overhead Gantry Display Panel

- Displays patient position, scan status, and instructions

- Visible from patient position

- Multi-language capable

Operator Console with Dual Medical-Grade Monitors

- As per Section 2 specifications

- Ergonomic keyboard and dedicated CT function keyboard/pad

- Optical mouse

Advanced Post-Processing Diagnostic Workstation

- As per Section 2 specifications

- Full 3D/MPR/MIP software suite

Integrated Isolation Transformer and Power Distribution Unit

- Rated for total system connected load

- IEC 61558 or equivalent compliant

- Overload and short-circuit protection

Standard Phantom Kit for Daily QA and Calibration

- Water phantom (large and small heads/body)

- Uniformity, linearity, and slice-thickness phantoms

- Compatible with manufacturer QA protocol

- Storage case included

ECG Gating Cable and Patient Lead Set

- Minimum 3-lead patient cable

- Disposable and reusable electrode compatibility

- Shielded cable for artifact-free acquisition

- Cardiac synchronization interface with gantry

## 1.4 Installation Requirements

- Dedicated CT scanner room with structural load capacity for gantry and table

- Lead-shielded walls, doors, and control room viewing window per NCRP/IAEA/local AERB requirements

- Radiation warning light and interlock system at room entrance

- Dedicated 3-phase electrical supply, 380—415 VAC, 50/60 Hz

- Dedicated earthing with earth resistance ≤1 ohm

- HVAC providing temperature 18—24°C, humidity 30—70%, with heat load per manufacturer datasheet

- UPS backup for console and workstation (minimum 30 minutes)

- Network connectivity to PACS/RIS/HIS (Gigabit Ethernet)

- False flooring or cable trench for interconnect cables

- Patient and technologist emergency stop buttons

## 1.5 Documentation

Supplier shall provide, in English:

- User/Operator Manual (hard copy and soft copy)

- Service Manual with schematics

- Parts Catalogue

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation Report and Site Acceptance Test (SAT) Report

- Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 1.6 Training

- Clinical (radiographer/technologist) training: minimum 5 working days on-site

- Radiologist application training: minimum 3 working days on-site

- Biomedical engineering user maintenance training: minimum 3 working days

- Advanced service engineer training: for minimum 2 nominated engineers

- Refresher training at 6 months post-commissioning

## 1.7 Warranty

- Comprehensive warranty: minimum 2 years from date of commissioning

- Includes all parts (including X-ray tube), labor, preventive maintenance (minimum 4 visits/year), breakdown maintenance, software updates and upgrades within the warranty period, remote diagnostic support

- Uptime guarantee: minimum 95%

- Response time: within 24 hours; resolution within 72 hours

- Availability of spare parts and after-sales service: minimum 10 years post-installation

# 2. 1.5T MRI MACHINE

Quantity: 1

## 2.1 Clinical & Functional Specifications

The system shall be a whole-body, high-field, superconducting Magnetic Resonance Imaging (MRI) scanner with a nominal field strength of 1.5 Tesla, intended for advanced diagnostic imaging across all anatomical regions including neuro, spine, musculoskeletal, body (chest, abdomen, pelvis), vascular (MR angiography with and without contrast), breast, cardiac, pediatric, and oncology applications.

The system shall provide high-resolution anatomical and functional imaging, including diffusion-weighted imaging (DWI), diffusion tensor imaging (DTI), spectroscopy (MRS), functional MRI (fMRI), MR angiography (contrast-enhanced and non-contrast time-of-flight/phase contrast), dynamic contrast-enhanced imaging, cardiac cine and delayed enhancement studies, and advanced quantitative imaging.

The system shall support high patient throughput via advanced parallel imaging, compressed sensing / accelerated acquisition, and AI-assisted image reconstruction where applicable.

The equipment shall be supplied complete with magnet, gradient system, RF transmit/receive system, patient table, operator console, RF-shielded cage (Faraday enclosure) with penetration panel, cryogen management system, quench pipe and installation, coils package, and patient physiological monitoring.

## 2.2 Engineering Technical Specifications

Magnet

- Type: superconducting, actively shielded

- Nominal field strength: 1.5 Tesla

- Homogeneity: typical values ≤1 ppm over 40 cm DSV (as declared per VRMS/peak method)

- Bore diameter: minimum 60 cm (wide-bore configurations of ≥70 cm preferred where available)

- Patient tunnel length: as compact as commercially available, typically ≤170 cm

- Cryogen: liquid helium, closed-loop zero-boil-off or minimum-boil-off cryocooler technology

- 5-Gauss fringe field footprint: shall be declared by supplier and shall fit within allocated room

Gradient System

- Peak amplitude per axis: minimum 33 mT/m

- Slew rate per axis: minimum 120 T/m/s

- Effective simultaneous performance on all three axes

- Water-cooled gradient coils

RF System

- Multi-channel RF transmit and receive architecture

- Number of independent RF receive channels: minimum 16 (32 or more preferred)

- Digital RF technology with in-bore or near-magnet digitization

- Compatible with parallel imaging techniques (SENSE, GRAPPA, ASSET, or equivalent)

Patient Table

- Motorized horizontal and vertical movement

- Maximum patient weight capacity: minimum 200 kg

- Detachable / docking transport trolley for patient transfer

- Positioning accuracy: within ±1 mm

- Emergency release for rapid patient removal

Imaging Sequences and Applications

- Full package of standard 2D and 3D sequences: SE, TSE/FSE, GRE, IR, FLAIR, STIR, EPI, SSFP/bSSFP

- Diffusion imaging (DWI, DTI with minimum 20 directions)

- Perfusion imaging (DSC, DCE, ASL)

- MR Spectroscopy (single-voxel and multi-voxel/CSI)

- MR Angiography: TOF, PC, contrast-enhanced

- Cardiac package: cine, tagging, delayed enhancement, flow quantification

- Advanced neuro package including fMRI (BOLD)

- Body: breath-hold and free-breathing techniques

- Breast imaging package

- Silent or quiet scanning sequences (acoustic noise reduction): required

- Parallel imaging and compressed sensing acceleration: required

- AI/deep-learning image reconstruction: required or upgradeable

Operator Console

- Dual medical-grade high-resolution color LCD/LED monitors, minimum 24-inch, minimum 1920 × 1200 resolution

- Multi-core 64-bit CPU, minimum 32 GB RAM, minimum 4 TB storage

- Image storage capacity: minimum 500,000 images

- DICOM 3.0 full compliance

Reconstruction System

- Dedicated reconstruction hardware

- Reconstruction speed: minimum 20,000 images per second (2D FFT baseline) or vendor-declared equivalent

- Real-time image display during acquisition

Connectivity

- DICOM 3.0 (Send, Receive, Query/Retrieve, MPPS, Worklist, Structured Reports)

- HL7 interface

- Ethernet minimum 1 Gbps

- USB and DVD-RW for media export

Electrical Specifications

- Input supply: 380—415 VAC, 3-phase, 50/60 Hz

- Connected load: as declared; typically ≤100 kVA

- Isolation transformer and power distribution included

- UPS for console

Cryogen Management

- Helium compressor / cryocooler with continuous operation

- Zero or minimum boil-off design

- Quench pipe with external venting

- Cryogen level monitoring and remote alerting

Environmental Conditions

- Operating temperature: 18—22°C

- Humidity: 40—60%, non-condensing

- RF shielding: complete Faraday cage, attenuation minimum 100 dB at 64 MHz

Applicable Standards

- IEC 60601-1

- IEC 60601-1-2 (EMC)

- IEC 60601-2-33 (Particular requirements for MRI safety)

- IEC 62304

- IEC 62366

- ISO 13485

- ISO 14971

- DICOM 3.0

- HL7

- CE MDR or US FDA 510(k)

## 2.3 Standard Accessories

Patient Transportation Table with Detachable Trolley System

- MR-conditional/MR-safe construction

- Non-ferromagnetic materials verified

- Compatible with magnet docking interface

- Minimum 200 kg patient weight capacity

- Locking wheels with foot brake

Patient Comfort Kit (Pneumatic Headphones, Earplugs, Emergency Squeeze Bulb)

- MR-safe pneumatic headphones with acoustic noise attenuation minimum 25 dB

- Disposable and reusable earplug options

- Emergency squeeze bulb (patient alert) with fiber-optic or pneumatic signaling to console — reliable non-electrical operation

Head/Neck Matrix Coil

- Multi-element phased-array design, minimum 16 channels

- Integrated head and neck coverage

- Compatible with parallel imaging

- Ergonomic patient interface with mirror and immobilization

Spine Matrix Coil

- Integrated table-embedded or drop-in phased-array coil

- Full cervico-thoraco-lumbar spine coverage

- Minimum 16 channels

- Compatible with parallel imaging

Body Matrix Coil

- Flexible phased-array design

- Minimum 16 channels

- Adjustable for chest, abdomen, and pelvis

- Lightweight and ergonomic

Integrated RF Shielding Enclosure System

- Complete modular Faraday cage

- Attenuation minimum 100 dB at 64 MHz

- RF penetration panel with filters for all cables and gas lines

- RF-shielded viewing window and door with contact fingers

- Installation and certification included

Operator Console with Dual High-Resolution Medical Displays

- As per Section 2 specifications

Helium Compressor Unit and Cryocooler System

- Continuous-duty compressor

- Water-cooled or air-cooled per site infrastructure

- Remote monitoring capability

Patient Physiological Monitoring Sensors (ECG, Pulse, Respiratory)

- MR-safe/MR-conditional certified

- Wireless or fiber-optic transmission

- 3-lead MR-compatible ECG

- Peripheral pulse oximeter (finger)

- Respiratory belt / bellows

- Interface with scanner for gating and triggering

MRI-Safe Phantom Set for Daily Calibration and QA

- Full ACR-compatible or equivalent phantom set

- Uniformity, SNR, geometric distortion, and slice-position phantoms

- Storage case included

## 2.4 Installation Requirements

- Structural: floor load capacity per magnet weight (declared by supplier, typically 4—6 tons); vibration-isolated foundation

- Room dimensions per manufacturer siting drawings; 5-Gauss line contained within room

- RF-shielded cage installed by qualified subcontractor

- Quench pipe: routing to exterior with proper venting, sized per magnet requirements

- HVAC: temperature 18—22°C, humidity 40—60%, dedicated for magnet room, control room, and equipment room; heat load per manufacturer

- Chilled water supply for cryocooler (if water-cooled)

- Dedicated 3-phase electrical supply

- Dedicated earthing with resistance ≤1 ohm; single-point ground

- UPS for console and workstations

- Network connectivity (Gigabit Ethernet)

- Ferromagnetic detection system at magnet room entrance

- Magnet quench button and emergency ventilation

- Non-ferromagnetic fixtures throughout magnet room (lights, fire suppression, HVAC diffusers)

## 2.5 Documentation

- User/Operator Manual (hard and soft copies)

- Service Manual with schematics

- Parts Catalogue

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Homogeneity and Magnet Field Test Report

- Installation and Site Acceptance Test Report

- Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

- MR Safety Documentation and Signage Package

## 2.6 Training

- Clinical/technologist training: minimum 10 working days on-site

- Radiologist application training: minimum 5 working days on-site (neuro, body, cardiac, MSK)

- Advanced applications training (spectroscopy, fMRI, cardiac): minimum 3 additional days

- Biomedical engineering training: minimum 5 working days

- MR safety officer training: minimum 2 nominated staff

- Refresher training at 6 months post-commissioning

## 2.7 Warranty

- Comprehensive warranty: minimum 2 years from date of commissioning

- Includes all parts, labor, preventive maintenance (minimum 4 visits/year), breakdown maintenance, cryogen top-up (if required by design), software updates and upgrades, coil replacement in case of failure, remote diagnostic support

- Uptime guarantee: minimum 95%

- Response time: within 24 hours; resolution within 72 hours

- Availability of spare parts and after-sales service: minimum 10 years post-installation

- Helium/cryogen supply guarantee included during warranty

# 3. MONO PLAN ANGIOGRAPHY SYSTEM (with radiation glasses, radiation sheets, storage rack, procedure stool)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The system shall be a single-plane (mono-plane), floor-mounted or ceiling-mounted digital angiography system with high-resolution flat-panel detector, intended for diagnostic and interventional cardiovascular, neuro-interventional, peripheral vascular, and general interventional radiology procedures.

The system shall support diagnostic angiography, DSA (Digital Subtraction Angiography), road-mapping, rotational angiography, 3D reconstruction, cone-beam CT (CBCT), stent enhancement, and advanced interventional workflow for endovascular procedures.

The equipment shall be delivered as a complete turn-key system comprising motorized C-arm gantry, flat-panel digital detector, high-frequency X-ray generator with cooling, motorized carbon-fiber patient table, ceiling-suspended monitor display, control-room and table-side operator consoles, radiation protection accessories, and full physiological interface module.

## 3.2 Engineering Technical Specifications

C-Arm Gantry

- Configuration: floor-mounted or ceiling-mounted mono-plane

- Motorized isocentric positioning

- Rotational speed for rotational angiography: minimum 40°/second

- LAO/RAO angulation: minimum ±120°

- CRA/CAU angulation: minimum ±45°

- SID (Source-to-Image Distance): motorized, adjustable range minimum 90—120 cm

- Isocenter positioning accuracy: within ±1 mm

Flat-Panel Digital Detector

- Type: amorphous silicon or CMOS with cesium iodide (CsI) scintillator

- Active imaging area: minimum 20 × 20 cm (30 × 40 cm preferred for cardiovascular; bidders to declare)

- Pixel pitch: maximum 200 μm (154 μm or finer preferred)

- Matrix: minimum 1024 × 1024

- Dynamic range: minimum 14-bit

- DQE: minimum 70% at 1 lp/mm at 0 lp/mm

- Fluoroscopy and acquisition modes supported

X-Ray Generator

- Power output: minimum 100 kW

- kV range: minimum 40—125 kV

- mA range: minimum 10—1000 mA

- Pulse fluoroscopy: selectable pulse rates 3.75, 7.5, 15, 30 pps

- Grid-controlled or grid-switched pulsed fluoroscopy

- Automatic exposure and dose control (ABC/ADRC)

X-Ray Tube

- Rotating anode, dual focal spot

- Focal spots: minimum small ≤0.6 mm and large ≤1.0 mm

- Anode heat storage capacity: minimum 2.4 MHU

- Liquid-cooled housing

Patient Table

- Motorized, float-top, carbon-fiber tabletop

- Longitudinal travel: minimum 130 cm

- Lateral travel: minimum ±18 cm

- Height adjustment: minimum 80—105 cm

- Patient weight capacity: minimum 200 kg (CPR capability up to 250 kg)

Monitor Display System

- Ceiling-suspended monitor suspension arm with multi-axis movement

- Minimum one large-format medical-grade display (minimum 55-inch) or matrix of dual/quad medical monitors (each minimum 19-inch, 2MP)

- Live and reference image display capability

- Physiological waveforms and hemodynamic parameters display

Imaging Modes and Software

- Digital fluoroscopy with pulsed and continuous modes

- DSA (Digital Subtraction Angiography)

- Road-mapping and vessel tracing

- Rotational angiography and 3D-DSA reconstruction

- Cone-Beam CT (CBCT) acquisition and reconstruction

- Stent enhancement / stent boost

- Automatic image quality optimization

- Dose reporting per IEC 60601-2-43 with DICOM structured dose reports

Radiation Dose Management

- Air-kerma and dose-area-product (DAP) meter integrated

- Real-time dose display

- Dose reporting on DICOM structured report

- Copper filtration (spectral shaping) selectable

- Last-image hold and fluoroscopy-loop store

Operator Consoles

- Control room console with dual medical-grade displays, minimum 19-inch, minimum 1280 × 1024

- Table-side control modules with touchscreen and joystick

- Multi-functional foot switches (waterproof)

Connectivity

- DICOM 3.0 full compliance

- HL7 interface

- Ethernet minimum 1 Gbps

Electrical Specifications

- Input: 380—415 VAC, 3-phase, 50/60 Hz

- Connected load: as declared; typically ≤100 kVA

Applicable Standards

- IEC 60601-1

- IEC 60601-1-2 (EMC)

- IEC 60601-1-3 (Radiation protection)

- IEC 60601-2-43 (Particular requirements for interventional X-ray)

- IEC 60601-2-54

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k)

## 3.3 Standard Accessories

Ceiling or Floor-Mounted Motorized C-Arm Gantry

- As per Section 2 specifications

- Collision detection and automatic stop

- Emergency stop buttons

High-Resolution Flat-Panel Digital Detector

- As per Section 2 specifications

Motorized Float-Top Carbon-Fiber Catheterization Table

- As per Section 2 specifications

- Radiolucent tabletop

- Emergency CPR release

- Accessory rails on both sides

Ceiling-Suspended Live and Reference Monitor Display Matrix

- As per Section 2 specifications

- Multi-axis motorized arm

- Cable management integrated

Control Room Operator Console with Dual Medical Displays

- As per Section 2 specifications

- Ergonomic workstation with keyboard, mouse, and function panel

Examination Room Table-Side Control Modules

- Touchscreen and joystick control

- Sterile-drapable

- Duplicate control of gantry, table, and image acquisition

X-Ray Generator Cabinet and Integrated Cooling Unit

- As per Section 2 specifications

- Closed-loop water or oil cooling with heat exchanger

Table-Mounted Lower Body Lead Curtains

- Lead equivalence minimum 0.5 mm Pb

- Sterilizable outer cover

- Table-rail mounted

Ceiling-Suspended Lead Glass Radiation Shield with Articulated Arm

- Lead equivalence minimum 0.5 mm Pb

- Optical clarity, minimum 40 × 60 cm viewing area

- Multi-axis articulated positioning

Continuous Physiological Monitoring Interface Module

- ECG, invasive/non-invasive blood pressure, SpO₂, respiration

- Integrated display on suspension monitor

- Recording and playback capability

Radiation Protection Accessories (as listed in tender)

- Radiation-protective glasses: minimum 0.75 mm Pb equivalent, wraparound design, minimum 2 pairs

- Radiation-protective aprons/sheets: minimum 0.5 mm Pb equivalent, front-and-back protection, various sizes, minimum 4 units

- Storage rack: wall-mounted or freestanding, powder-coated steel, capacity for minimum 6 aprons

- Procedure stool: adjustable height, lockable casters, foot-ring, wipe-clean upholstery, minimum 2 units

## 3.4 Installation Requirements

- Dedicated cath-lab room with structural load capacity for ceiling/floor mount

- Lead-shielded walls, doors, and viewing window per local radiation protection code

- Radiation warning light and door interlock

- Dedicated 3-phase electrical supply

- Dedicated earthing, resistance ≤1 ohm

- HVAC: temperature 20—22°C, humidity 40—60%, laminar flow if procedures require

- Chilled water for generator cooling (if applicable)

- UPS backup for console and life-safety systems (minimum 30 minutes)

- Network connectivity (Gigabit Ethernet)

- Medical gas outlets (O₂, vacuum, medical air)

- False ceiling with reinforcement for ceiling-mounted components

## 3.5 Documentation

- User Manual, Service Manual, Parts Catalogue (English)

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation and SAT Report

- Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 3.6 Training

- Clinical training for interventionalists and technologists: minimum 5 working days

- Application training including advanced 3D and CBCT: minimum 3 working days

- Biomedical engineering training: minimum 3 working days

- Refresher training at 6 months post-commissioning

## 3.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube and flat panel detector), labor, preventive maintenance (minimum 4 visits/year), breakdown maintenance, software updates, remote diagnostic support

- Uptime guarantee: minimum 95%

- Response time: within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 4. FLUOROSCOPY UNIT (R/F SYSTEM WITH IMAGE INTENSIFIER)

Quantity: 1

## 4.1 Clinical & Functional Specifications

The system shall be a motorized remote-controlled or table-side controlled Radiography/Fluoroscopy (R/F) system with image intensifier or flat-panel fluoroscopy detector, intended for gastrointestinal contrast studies, urological studies, myelography, arthrography, ERCP, hysterosalpingography, orthopedic reduction procedures, general fluoroscopic-guided procedures, and general radiography.

The system shall provide dynamic real-time fluoroscopy and digital spot radiography with dose-optimization features.

## 4.2 Engineering Technical Specifications

Tilting Table

- Motorized tilting: minimum +90° to −45° (−90° preferred where available)

- Longitudinal tabletop travel: minimum 100 cm

- Transverse tabletop travel: minimum ±12 cm

- Patient weight capacity: minimum 200 kg

- Radiolucent carbon-fiber tabletop

- SID: adjustable, minimum 100—150 cm

Image Detector

- Image intensifier: minimum 30 cm (12-inch) input diameter with multi-mode magnification (minimum 3 modes), OR

- Dynamic flat-panel detector: active area minimum 30 × 30 cm, pixel pitch ≤200 μm, matrix minimum 1024 × 1024, dynamic range minimum 14-bit

- Bidder to declare which option offered

X-Ray Generator

- High-frequency, microprocessor-controlled

- Power output: minimum 65 kW

- kV range: minimum 40—125 kV

- mA range: minimum 10—800 mA

- Pulsed fluoroscopy: selectable pulse rates including 1, 3, 7.5, 15, 30 pps

- Automatic Brightness Control (ABC) and Automatic Dose Rate Control (ADRC)

X-Ray Tube

- Rotating anode, dual focal spot (≤0.6 mm and ≤1.2 mm)

- Anode heat storage: minimum 300 kHU

- Liquid or oil cooling

Digital Acquisition Workstation

- Dual medical-grade preview monitors, minimum 19-inch, minimum 1280 × 1024

- Multi-core 64-bit CPU, minimum 16 GB RAM, minimum 1 TB SSD/HDD

- Image storage minimum 100,000 images

- Real-time image processing (edge enhancement, noise reduction)

- DSA and road-map (optional, bidder to declare)

- Last-image hold and fluoroscopy-loop store

Operating Modes

- Continuous and pulsed fluoroscopy

- Digital spot radiography

- Serial radiography (up to 6 fps minimum)

- Tomography (optional)

Radiation Dose Management

- DAP meter integrated

- Real-time dose display

- Copper spectral filtration (selectable, e.g., 0.1/0.2/0.3/0.6 mm Cu)

- DICOM dose structured report

Connectivity

- DICOM 3.0 full compliance (Send, Receive, MWL, MPPS, Storage Commit)

- HL7 interface

- Ethernet 1 Gbps

Electrical Specifications

- Input: 380—415 VAC 3-phase, 50/60 Hz

- Connected load: as declared

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-54 (radiography and radioscopy)

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k)

## 4.3 Standard Accessories

Motorized Tilting R/F Table (+90° to −45° or −90°)

- As per Section 2 specifications

- Foot-rest and shoulder-rest attachments for vertical positioning

- Compression cone/paddle

- Patient hand grips

High-Resolution Image Intensifier or Flat-Panel Fluoroscopy Detector

- As per Section 2 specifications

Dynamic Acquisition and Review Workstation

- As per Section 2 specifications

Ceiling-Mounted Twin Monitor Suspension System

- Multi-axis articulated arm

- Cable management

- Vibration-damped

Dual Medical-Grade Preview Monitors

- Minimum 19-inch, minimum 1280 × 1024 resolution

- Medical-grade LCD/LED

- DICOM GSDF calibrated

Multi-Functional Foot Switch

- Waterproof (IPX8)

- Multi-pedal for fluoroscopy and exposure

- Ergonomic non-slip design

Table-Mounted Lead Protective Drapes

- Lead equivalence minimum 0.5 mm Pb

- Sterilizable outer cover

- Quick-attach to table rail

Grid for Scatter Radiation Reduction

- Anti-scatter grid, ratio minimum 10:1

- Removable and interchangeable

- Focused grid, appropriate focal distance

## 4.4 Installation Requirements

- Dedicated R/F room with lead shielding per local code

- Radiation warning light at entrance

- Dedicated 3-phase electrical supply

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS for console

- Network connectivity (Gigabit Ethernet)

- Lead-glass viewing window from control area

## 4.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 4.6 Training

- Clinical and technologist training: minimum 3 working days

- Application training: minimum 2 working days

- Biomedical engineering training: minimum 2 working days

- Refresher training at 6 months

## 4.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube), labor, PM (minimum 4 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 5. DIGITAL FIXED X-RAY MACHINE (high-frequency, 500—650 mA, overhead tube suspension)

Quantity: 2

## 5.1 Clinical & Functional Specifications

The system shall be a high-frequency, digital, fixed radiography system with overhead tube suspension and 3D ceiling rails, intended for general radiographic examinations including chest, abdomen, spine, skull, musculoskeletal, and pediatric imaging in both supine (table) and erect (wall Bucky) positions.

The system shall use flat-panel digital detectors (fixed and/or wireless tethered) for immediate image preview, review, and DICOM transmission.

## 5.2 Engineering Technical Specifications

X-Ray Generator

- High-frequency, microprocessor-controlled

- Power output: minimum 65 kW (80 kW preferred)

- Maximum tube current: minimum 500 mA, up to 650 mA

- kV range: minimum 40—150 kV

- mAs range: minimum 0.5—500 mAs

- Exposure time: minimum 1 ms to 6.3 s

- Automatic Exposure Control (AEC) with minimum 3 ionization chambers

X-Ray Tube

- Rotating anode, dual focal spot (small ≤0.6 mm, large ≤1.2 mm)

- Anode heat storage capacity: minimum 300 kHU

- Anode rotation speed: minimum 9,000 rpm (high-speed)

- Housing heat dissipation with oil cooling

Overhead Tube Suspension (OTS)

- Motorized 3D ceiling rails

- Longitudinal travel: minimum 300 cm

- Transverse travel: minimum 200 cm

- Vertical travel: minimum 150 cm

- Tube rotation: minimum ±180°

- Tube angulation: minimum ±120°

- Auto-positioning and auto-tracking with table and wall Bucky

- Integrated touchscreen display on tube head showing kV, mAs, SID, angle, patient data

Radiography Table

- Motorized 4-way floating tabletop

- Longitudinal travel: minimum ±45 cm

- Transverse travel: minimum ±12 cm

- Height adjustment (motorized): minimum 55—85 cm

- Patient weight capacity: minimum 250 kg

- Radiolucent carbon-fiber tabletop

- Integrated Bucky with reciprocating grid

Vertical Wall Bucky

- Manual or motorized vertical travel: minimum 40—180 cm

- Rotatable/tiltable Bucky assembly

- Integrated reciprocating grid

- AEC chambers

Flat-Panel Digital Detectors

- Fixed detectors in table and wall Bucky: active area minimum 43 × 43 cm

- Pixel pitch: maximum 150 μm

- Matrix: minimum 2800 × 2800

- Dynamic range: minimum 16-bit

- DQE: minimum 65% at 0 lp/mm

- Optional wireless tethered/wireless portable detector, minimum 35 × 43 cm, IP54 or better

Acquisition Workstation

- Medical-grade preview monitor: minimum 19-inch, minimum 1280 × 1024

- Multi-core 64-bit CPU, minimum 16 GB RAM, minimum 1 TB storage

- Image storage minimum 50,000 images

- Image processing: automatic exposure indicator, grayscale optimization, edge enhancement, tissue equalization

- DICOM 3.0 full compliance

Connectivity

- DICOM 3.0 (Send, Receive, MWL, MPPS, Storage Commit, Print)

- HL7 interface

- Ethernet 1 Gbps

Electrical Specifications

- Input: 380—415 VAC 3-phase, 50/60 Hz

- Connected load: as declared

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-54

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k)

## 5.3 Standard Accessories

Overhead Tube Suspension System with 3D Ceiling Rails

- As per Section 2 specifications

- Collision detection and emergency stop

- Cable management chain

Motorized 4-Way Floating-Top Radiography Table

- As per Section 2 specifications

- Foot switch for tabletop movement

- Handrails and accessory rails

Vertical Bucky Wall Stand

- As per Section 2 specifications

- Chin/head rest, hand grips

- Emergency lock

Flat-Panel Digital Detectors (Fixed or Tethered)

- As per Section 2 specifications

- Charging cradle for portable/wireless detectors (if supplied)

High-Frequency X-Ray Generator Cabinet

- As per Section 2 specifications

- Wall or floor mounted

- Ventilated enclosure

Acquisition Workstation with Medical-Grade Preview Monitor

- As per Section 2 specifications

Handheld Exposure Switch with Coiled Cable

- Two-stage exposure button (prep + expose)

- Coiled cable minimum 4 m extended

- Ergonomic grip

Grid Set for Table and Wall Bucky (Removable)

- Anti-scatter grid, ratio minimum 10:1 (12:1 optional)

- Focused grids appropriate to SID

- Removable and interchangeable

## 5.4 Installation Requirements

- Radiography room with lead shielding per local code

- Radiation warning light at entrance

- Reinforced ceiling for overhead tube suspension rails

- Dedicated 3-phase electrical supply

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS for workstation

- Network connectivity (Gigabit Ethernet)

- Lead-glass viewing window from control area

## 5.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 5.6 Training

- Clinical/technologist training: minimum 3 working days per unit

- Biomedical engineering training: minimum 2 working days

- Refresher training at 6 months

## 5.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube and flat-panel detectors), labor, PM (minimum 4 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 6. MOBILE / PORTABLE X-RAY MACHINE

Quantity: 2

## 6.1 Clinical & Functional Specifications

The system shall be a motorized, battery-powered mobile digital radiography unit intended for bedside imaging in intensive care units, emergency departments, operating theatres, wards, and neonatal units, providing high-quality digital radiographs with wireless flat-panel detector.

The system shall support fully mobile operation including motorized drive, on-board image acquisition, review, and DICOM wireless transmission to PACS.

## 6.2 Engineering Technical Specifications

Generator

- High-frequency, microprocessor-controlled

- Power output: minimum 30 kW (32—40 kW preferred)

- kV range: minimum 40—125 kV, selectable in configurable steps

- mAs range: minimum 0.1—320 mAs

- Exposure time: minimum 1 ms to 6.3 s

- Battery-powered exposure independent of mains

X-Ray Tube

- Rotating anode, dual focal spot (small ≤0.6 mm, large ≤1.2 mm)

- Anode heat storage: minimum 140 kHU

- Oil cooling

Mobility System

- Integrated motorized drive with forward/reverse, variable speed

- Rechargeable battery pack for drive and exposures

- Battery capacity: minimum 300 exposures at typical technique per full charge

- Drive range: minimum 20 km per full charge

- Total unit weight: manageable for hospital mobility (typically ≤600 kg)

- Anti-collision sensors and emergency brake

- Automatic parking brake

Tube Arm

- Telescopic or articulated

- Vertical travel: minimum 130 cm

- Horizontal reach: minimum 130 cm from column

- Rotation: minimum ±270° at column

- Counterbalanced with self-locking positioning

Collimator

- Manual multi-leaf collimator with light-field indicator

- Rotatable

- Filter tray

Wireless Flat-Panel Detector

- Active area: minimum 35 × 43 cm

- Pixel pitch: maximum 150 μm

- Matrix: minimum 2300 × 2800

- Dynamic range: minimum 16-bit

- Wireless connectivity: Wi-Fi (802.11 a/b/g/n/ac dual-band)

- Battery-powered detector: minimum 8 hours or 300 exposures per charge

- IP rating: minimum IP54

- Impact-resistant construction

- Protective storage bin on machine

Acquisition Workstation (On-Board)

- Medical-grade touchscreen preview monitor: minimum 17-inch, minimum 1280 × 1024

- Multi-core CPU, minimum 8 GB RAM, minimum 500 GB storage

- Image storage minimum 20,000 images

- Image processing and preview within seconds of exposure

- DICOM 3.0 (Send, MWL, MPPS, Storage Commit)

Connectivity

- Wi-Fi and wired Ethernet

- DICOM 3.0 full compliance

- HL7 interface

Electrical Specifications

- Charging input: 100—240 VAC, 50/60 Hz, single-phase

- Charging time: maximum 8 hours from fully depleted

- Battery type: sealed maintenance-free or lithium-ion

Radiation Dose Management

- DAP meter (optional/preferred)

- Dose reporting on DICOM structured report

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-54

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k)

## 6.3 Standard Accessories

Integrated Motorized Drive System with Rechargeable Battery Pack

- As per Section 2 specifications

- Battery status display

- Battery charging cable and connector

Telescopic or Articulated Tube Arm with Manual Collimator

- As per Section 2 specifications

- Multi-leaf collimator with light-field indicator

- Rotatable collimator

Wireless Flat-Panel Digital Detector with Protective Storage Bin

- As per Section 2 specifications

- Detector bag with cleaning-compatible cover

- Storage bin with cushioned lining and secure lock

Integrated Battery Charger and Power Cord

- Charging via standard hospital outlet

- Cord length minimum 4 m

- Auto-shutoff at full charge

Handheld Exposure Switch with Extendable Cable

- Two-stage exposure button (prep + expose)

- Coiled cable minimum 4 m extended

- Ergonomic non-slip grip

Integrated Barcode Scanner for Patient Data Entry

- 1D and 2D barcode compatibility

- Wired or wireless

- Integrated with worklist and patient data population

## 6.4 Installation Requirements

- No fixed installation required

- Charging outlet: standard 100—240 VAC single-phase, dedicated 15 A circuit preferred

- Wi-Fi coverage in operating areas (ICU, OT, wards, ED)

- Network integration with PACS/RIS

- Storage/parking area with secured charging point

- Radiation warning signage for use areas

- Lead aprons and portable lead screens (procured separately)

## 6.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 6.6 Training

- Clinical/technologist training: minimum 2 working days per unit

- Biomedical engineering training: minimum 2 working days

- Refresher training at 6 months

## 6.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube, flat-panel detector, and batteries), labor, PM (minimum 4 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 7. FULL-FIELD DIGITAL MAMMOGRAPHY UNIT

Quantity: 1

## 7.1 Clinical & Functional Specifications

The system shall be a full-field digital mammography (FFDM) unit intended for high-resolution diagnostic and screening mammography of the female breast, including standard 2D imaging, magnification views, spot compression views, and shall be upgradeable or supplied with digital breast tomosynthesis (DBT) capability where specified.

The system shall provide low-dose, high-contrast, high-spatial-resolution imaging with automatic exposure control optimized for breast composition.

## 7.2 Engineering Technical Specifications

Gantry / C-Arm

- Motorized C-arm with isocentric rotation

- Rotation range: minimum −180° to +180°

- Motorized vertical travel: minimum 70—150 cm to accommodate patients seated or standing, pediatric to tall adult

- Automatic and manual angulation

X-Ray Tube

- Rotating anode

- Target material: dual-track (e.g., Mo/Rh, W/Rh, or W/Ag) — bidder to declare configuration

- Focal spot: minimum small ≤0.1 mm and large ≤0.3 mm

- Anode heat storage: minimum 300 kHU

- Filter selection: automatic based on breast composition

X-Ray Generator

- High-frequency, microprocessor-controlled

- Power output: minimum 5 kW

- kV range: minimum 22—39 kV in 1 kV steps

- mAs range: minimum 2—500 mAs

- Automatic Exposure Control (AEC) with breast-composition analysis

Digital Detector

- Type: amorphous selenium (a-Se) direct conversion or CsI-based indirect conversion

- Active area: minimum 24 × 30 cm

- Pixel pitch: maximum 100 μm (85 μm or finer preferred)

- Matrix: minimum 2400 × 3000

- Dynamic range: minimum 14-bit

- DQE: minimum 65% at 0 lp/mm

Compression System

- Motorized compression with manual override

- Compression force: adjustable, minimum 30 N to 200 N

- Compression paddle sizes: standard 24 × 30 cm and 18 × 24 cm

- Automatic compression release after exposure

- Dual-pedal foot switch for hands-free compression

Acquisition Workstation

- Medical-grade high-resolution preview monitor

- Multi-core 64-bit CPU, minimum 16 GB RAM, minimum 2 TB storage

- Image storage minimum 20,000 studies

- Image processing tools including CLAHE, contrast enhancement, and image reversal

- DICOM 3.0 full compliance

Diagnostic Review Monitor

- 5 MP medical-grade grayscale mammography monitor, minimum 21-inch diagonal

- Resolution: minimum 2048 × 2560

- DICOM GSDF calibrated

- Luminance: minimum 1000 cd/m² calibrated

- Backlight stabilization sensor

Imaging Modes

- Standard 2D FFDM

- Magnification view with dedicated magnification stand

- Spot compression view

- Digital Breast Tomosynthesis (DBT) — supplied or upgradeable (bidder to declare)

- Synthesized 2D from tomosynthesis (if DBT supplied)

Radiation Dose Management

- Automatic Exposure Control optimized per breast thickness and composition

- Dose display per exposure

- Average glandular dose (AGD) calculation and DICOM structured dose reporting

Connectivity

- DICOM 3.0 (Send, Receive, MWL, MPPS, Storage Commit, Print, Mammography CAD SR)

- HL7 interface

- Ethernet 1 Gbps

Electrical Specifications

- Input: 200—240 VAC, single- or three-phase, 50/60 Hz (as per configuration)

- Connected load: as declared

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-45 (Particular requirements for mammographic X-ray equipment)

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- MQSA-equivalent quality standards

- CE MDR or US FDA 510(k)

## 7.3 Standard Accessories

Motorized C-Arm Compression Device

- As per Section 2 specifications

- Emergency compression release

- Compression force display

High-Resolution Digital Mammography Detector

- As per Section 2 specifications

Face Shield and Protective Patient Guards

- Transparent, radiolucent face shield

- Removable and autoclavable/wipe-clean

- Ergonomic patient positioning

Standard Compression Plates (24 × 30 cm and 18 × 24 cm)

- Polycarbonate or equivalent radiolucent material

- Autoclave or high-level disinfection compatible

- Curved and flat variants for optimal compression

Magnification Kit (Magnification Plate and Spot Compression Paddles)

- Magnification stand with 1.5×—2.0× factor

- Small spot compression paddles (multiple sizes)

- Compatible with C-arm quick-lock interface

Dual-Pedal Foot Switch for Hands-Free Compression Control

- Waterproof (IPX8)

- Two-pedal design (compress/release)

- Ergonomic, non-slip

Dedicated Acquisition Workstation with 5 MP Medical-Grade Diagnostic Monitor

- As per Section 2 specifications

- Ergonomic keyboard and optical mouse

- Dictation microphone (optional)

Transparent Radiation Protective Screen for the Operator

- Lead-acrylic or lead-glass, minimum 0.5 mm Pb equivalent

- Freestanding on lockable casters or wall-mounted

- Optical clarity, minimum 60 × 90 cm

## 7.4 Installation Requirements

- Dedicated mammography room with lead shielding per local code

- Radiation warning light at entrance

- Dedicated single- or three-phase electrical supply

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS for workstation

- Network connectivity (Gigabit Ethernet)

- Privacy considerations (curtained/private examination area)

## 7.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate (per IEC 61223-3-2)

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

- Mammography-specific QA/QC manual

## 7.6 Training

- Clinical/technologist training: minimum 3 working days

- Radiologist application training: minimum 2 working days (including DBT if supplied)

- Biomedical engineering training: minimum 2 working days

- Refresher training at 6 months

## 7.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube and digital detector), labor, PM (minimum 4 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 8. DEXA SCANNER (Bone Mineral Densitometry)

Quantity: 1

## 8.1 Clinical & Functional Specifications

The system shall be a central Dual-Energy X-ray Absorptiometry (DEXA/DXA) scanner intended for accurate measurement of bone mineral density (BMD) at the lumbar spine, proximal femur (hip), forearm, and whole-body sites for diagnosis and monitoring of osteoporosis, osteopenia, and body-composition analysis (fat mass, lean mass, and bone mass).

The system shall provide low-dose imaging, high measurement precision and accuracy suitable for clinical and research use, and shall support pediatric and adult reference databases.

## 8.2 Engineering Technical Specifications

Scanning Technology

- Dual-energy X-ray absorptiometry with narrow-fan-beam or pencil-beam technology

- Simultaneous or switched dual-energy acquisition

X-Ray Source

- kV: dual-energy switched or filtered

- mA: low-current, low-dose design

- Radiation dose per lumbar spine scan: typical ≤10 μSv

Detector

- Multi-element solid-state detector array

- High-count-rate photon counting or integrating

Scanning Table

- Motorized central DXA table with C-arm containing X-ray tube and detector

- Table dimensions accommodating patients up to minimum 200 kg

- Scan area coverage: full body plus dedicated hip, spine, and forearm modes

- Motorized scanning arm

Measurement Precision and Accuracy

- BMD precision (in-vivo CV): typical ≤1% at lumbar spine

- BMD accuracy: within ±1% versus reference phantom

- Long-term stability: within ±0.5% over 1 year with proper QA

Scan Modes

- AP lumbar spine

- Proximal femur (hip) — dual hip

- Forearm

- Whole-body composition (fat, lean, bone)

- Lateral spine (VFA — Vertebral Fracture Assessment) — preferred

- Pediatric mode with pediatric reference database

- Small animal / research mode (optional)

Analysis Software

- T-score and Z-score calculation per WHO criteria

- Reference databases: NHANES, manufacturer's normative database, adjustable ethnicity

- FRAX or equivalent fracture risk calculation

- Trend analysis and serial comparison

- Body-composition analysis (VAT, android/gynoid regions)

- Automatic ROI placement with manual override

Acquisition Workstation

- Medical-grade color LCD/LED monitor: minimum 22-inch, minimum 1920 × 1080

- Multi-core 64-bit CPU, minimum 8 GB RAM, minimum 1 TB storage

- DICOM 3.0 full compliance

Connectivity

- DICOM 3.0 (Send, Receive, MWL, MPPS, Structured Report — Bone Densitometry)

- HL7 interface

- Ethernet 1 Gbps

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Connected load: typically ≤3 kVA

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-54

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0 (with BMD SR)

- HL7

- CE MDR or US FDA 510(k)

## 8.3 Standard Accessories

Central DXA Scanning Table with Motorized Arm

- As per Section 2 specifications

- Patient weight capacity minimum 200 kg

- Radiolucent scanning surface

- Emergency stop and patient safety features

Patient Positioning Pillows, Leg Blocks, and Foam Supports

- Radiolucent foam material

- Wipe-clean cover

- Full set for spine, hip, forearm, and whole-body positioning

- Hip positioner with foot rotation angle

Dedicated Acquisition Computer System with Medical-Grade Monitor

- As per Section 2 specifications

- Ergonomic keyboard and optical mouse

QA Calibration Phantom Block

- Manufacturer-specific daily QA phantom

- Traceable calibration certificate

- Storage case

Spine and Hip Reference Standards

- Anthropomorphic spine and hip phantoms for periodic verification

- Traceable to national/international reference standards

- Storage case

## 8.4 Installation Requirements

- Room dimensions per manufacturer siting requirements (typically 3 × 4 m minimum)

- Lead shielding: minimal per local code (DXA is very low dose but shielding assessment required)

- Radiation warning signage at entrance

- Dedicated single-phase electrical supply

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS for workstation

- Network connectivity (Gigabit Ethernet)

## 8.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Radiation Output and BMD Calibration Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 8.6 Training

- Clinical/technologist training: minimum 3 working days

- Physician interpretation training: minimum 2 working days

- Biomedical engineering training: minimum 1 working day

- Refresher training at 6 months

## 8.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts, labor, PM (minimum 2 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 9. C-ARM IMAGE INTENSIFIER WITH DUAL MONITORS

Quantity: 1

## 9.1 Clinical & Functional Specifications

The system shall be a mobile surgical C-arm image intensifier system intended for intraoperative fluoroscopic guidance in orthopedic, general surgical, urological, pain-management, vascular, and cardiac procedures.

The system shall provide high-resolution real-time fluoroscopy, digital spot imaging, image processing, and DICOM connectivity to PACS.

## 9.2 Engineering Technical Specifications

C-Arm Mechanical

- Mobile C-arm with manual steering and mechanical locks on all axes

- C-arm depth (throat): minimum 65 cm

- C-arm free space: minimum 75 cm

- Orbital rotation: minimum 130° (e.g., +90°/−40°)

- Angular (roll) rotation: minimum ±225°

- Horizontal travel: minimum 20 cm

- Vertical (lift) travel: minimum 40 cm (motorized preferred)

- Swivel: ±12° minimum

- Wig-wag: ±12° minimum

- 5 or 6 axes of motion

X-Ray Tube

- Rotating or stationary anode

- Focal spot: minimum small ≤0.6 mm, large ≤1.5 mm

- Anode heat storage: minimum 45 kHU

- Cooling: passive or forced air

X-Ray Generator

- High-frequency, microprocessor-controlled

- Power output: minimum 3 kW (continuous) / 15 kW (boost) or equivalent

- kV range: minimum 40—120 kV

- mA range: minimum 0.1—15 mA (fluoroscopy), up to 100 mA (digital spot)

- Pulsed fluoroscopy: selectable pulse rates 1, 4, 8, 12.5, 25 pps

- ABC/ADRC with dose optimization

Image Detector

- Image Intensifier: minimum 23 cm (9-inch) input diameter, multi-mode magnification (minimum 3), OR

- Flat-panel detector: active area minimum 20 × 20 cm, pixel pitch ≤200 μm, matrix ≥1024 × 1024, dynamic range ≥14-bit

- Bidder to declare option offered

Image Processing / Workstation

- Dual medical-grade high-resolution monitors, minimum 19-inch, minimum 1280 × 1024

- Multi-core CPU, minimum 8 GB RAM, minimum 500 GB storage

- Image storage minimum 20,000 images

- Real-time image processing: noise reduction, edge enhancement, contrast optimization

- Last-image hold, fluoroscopy loop store, cine review

- DICOM 3.0 (Send, MWL, MPPS)

Mobile Viewing Cart

- Integrated storage drawers

- Ergonomic monitor arm

- Cable management

- Locking wheels

Radiation Dose Management

- DAP meter

- Cumulative dose display

- Pulsed fluoroscopy default

- Copper filtration selectable (optional)

Connectivity

- DICOM 3.0

- Wi-Fi and Ethernet

- USB and DVD-RW for export

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Connected load: typically ≤2 kVA (fluoro), ≤6 kVA (boost)

Applicable Standards

- IEC 60601-1, 60601-1-2, 60601-1-3

- IEC 60601-2-43 or 60601-2-54 as applicable

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0

- CE MDR or US FDA 510(k)

## 9.3 Standard Accessories

Mobile C-Arm Gantry with Manual Steering and Mechanical Locks

- As per Section 2 specifications

- Steerable rear wheels

- Central braking system

High-Resolution Image Intensifier or Flat-Panel Detector

- As per Section 2 specifications

Mobile Viewing Cart with Integrated Storage Drawers

- As per Section 2 specifications

- Cable management

- Ergonomic monitor arm

Dual Medical-Grade High-Resolution Display Monitors

- As per Section 2 specifications

- DICOM GSDF calibrated

Multi-Functional Waterproof Foot Switch

- IPX8 rated

- Multi-pedal (fluoro low/high, digital acquisition)

- Cable length minimum 4 m

Laser Positioning Pointer Device

- Class 1 or Class 2 laser

- Cross-hair positioning

- Integrated with tube head

Sterile Drape Holders and Clips

- Autoclavable stainless steel or equivalent

- Compatible with standard surgical drapes

- Sterile-field compatible

## 9.4 Installation Requirements

- No fixed installation

- Storage/parking area with mains outlet for standby power

- Radiation shielding of user areas per local code (OT is typically shielded)

- Network connectivity in OT (wired or Wi-Fi)

- Lead aprons and thyroid shields (procured separately)

## 9.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Radiation Output Calibration Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 9.6 Training

- Clinical/OT staff training: minimum 2 working days

- Biomedical engineering training: minimum 2 working days

- Refresher training at 6 months

## 9.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including X-ray tube and detector), labor, PM (minimum 4 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 10. POWER CONTRAST INJECTOR (CT-COMPATIBLE)

Quantity: 1

## 10.1 Clinical & Functional Specifications

The system shall be a dual-head power contrast injector intended for controlled intravenous administration of iodinated contrast media and saline flush during CT imaging examinations, including routine contrast-enhanced CT, CT angiography (CTA), and dynamic multi-phase studies.

The injector shall provide precise, programmable flow-rate, volume, and pressure control, and shall synchronize with the CT scanner via a dedicated interface for automated bolus timing and multi-phase protocols.

## 10.2 Engineering Technical Specifications

Injector Configuration

- Dual-head (dual-syringe) design: one for contrast, one for saline flush

- Motorized syringe drive with precise servo control

Flow Rate

- Programmable range: minimum 0.1 to 10.0 mL/second

- Flow-rate resolution: 0.1 mL/second

- Flow-rate accuracy: within ±5% of set value

Volume

- Programmable volume per phase: minimum 1 to 200 mL

- Volume resolution: 1 mL

- Volume accuracy: within ±3% of set value

Pressure

- Programmable pressure limit: adjustable, minimum 50 to 325 psi

- Pressure resolution: configurable

- Real-time pressure display and pressure-limit alarm

Multi-Phase Programming

- Minimum 6 programmable phases per protocol

- Configurable flow rate, volume, pressure, and delay per phase

- Pause and hold function between phases

- Test injection capability

Syringe System

- Compatible with standard 100 mL, 150 mL, and 200 mL disposable syringes (or equivalent per manufacturer's system)

- Automatic syringe recognition and priming

- Air-detection sensor with automatic stop

- Anti-siphon feature

User Interface

- Color touchscreen display on injector head: minimum 7-inch

- Color touchscreen on remote console: minimum 10-inch

- Multi-language support

- Protocol library with user-configurable presets (minimum 50 protocols)

- Injection history logging (minimum 500 cases)

Remote Control Console

- Located in control room

- Touchscreen operation

- Duplicate control and monitoring

- Emergency stop

Safety Features

- Air detection and auto-stop

- Pressure-limit auto-stop

- Extravasation detection sensor (optional/preferred)

- Audible and visual alarms

- Manual emergency stop

- Automatic drive-back on completion

CT Synchronization Interface

- Hardwired trigger interface to CT scanner

- Bolus-tracking integration

- Simultaneous start of scan and injection

Connectivity

- DICOM 3.0 MPPS and Structured Reports (Injection Record)

- Ethernet interface

- USB for protocol import/export

Mobile Pedestal Stand

- Lockable casters (all four)

- Height-adjustable

- Cable management

Syringe Heater Jacket

- Maintains contrast at body temperature (typically 37°C)

- Temperature accuracy within ±2°C

- Automatic temperature control with over-temperature protection

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Battery backup: minimum 30 minutes of full operation

- Connected load: typically ≤500 VA

Applicable Standards

- IEC 60601-1

- IEC 60601-1-2 (EMC)

- IEC 60601-2-24 (Particular requirements for infusion pumps and controllers)

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0

- CE MDR or US FDA 510(k)

## 10.3 Standard Accessories

Dual-Head Injector Arm Assembly

- As per Section 2 specifications

- Ergonomic arm with multi-axis positioning

- Sterilizable / wipe-clean surfaces

Mobile Pedestal Stand with Lockable Casters

- As per Section 2 specifications

- Powder-coated corrosion-resistant frame

Remote Control Console for Control Room Installation

- As per Section 2 specifications

- Cable length or wireless connection to injector head sufficient for control room installation

Syringe Heater Jacket

- As per Section 2 specifications

- Compatible with standard syringe sizes

Interface Cable for CT Scanner Synchronization

- Compatible with common CT scanner trigger interfaces

- Adequate length for typical CT room layout (minimum 8 m)

- Shielded for EMC compliance

Main Power Supply Unit and Power Cords

- Hospital-grade power cord

- Cord length minimum 4 m

- IEC 60601-1 compliant

## 10.4 Installation Requirements

- Installation within CT scanner room and control room

- Standard single-phase electrical supply

- Dedicated earthing ≤1 ohm

- Interface cabling routed with CT system

- Network connectivity for DICOM MPPS to RIS/PACS

## 10.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Flow-rate and volume accuracy verification certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 10.6 Training

- Clinical/technologist training: minimum 2 working days

- Biomedical engineering training: minimum 1 working day

- Refresher training at 6 months

## 10.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts, labor, PM (minimum 2 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 11. COMPUTED / DIGITAL RADIOGRAPHY (CR / DR) SYSTEM

Quantity: 1

## 11.1 Clinical & Functional Specifications

The system shall be a Computed Radiography (CR) or Digital Radiography (DR) image acquisition, processing, and management system intended for the digitization or direct acquisition, processing, review, and DICOM transmission of general radiographic images, converting analog X-ray workflow into a digital PACS-compatible workflow.

For CR configuration: the system shall accept standard photostimulable phosphor (PSP) imaging plates in multiple cassette sizes and produce high-resolution digital radiographs.

For DR configuration: the system shall accept wireless or tethered flat-panel detectors for direct digital acquisition.

Bidder shall declare the configuration offered (CR reader with cassettes, or DR retrofit with wireless flat-panel detectors) and provide full technical compliance accordingly.

## 11.2 Engineering Technical Specifications

CR Reader (if CR configuration offered)

- Cassette sizes accepted: minimum 18 × 24 cm, 24 × 30 cm, 35 × 35 cm, and 35 × 43 cm

- Throughput: minimum 60 plates per hour

- Pixel pitch: maximum 100 μm

- Spatial resolution: minimum 5 lp/mm

- Dynamic range: minimum 12-bit (16-bit preferred)

- Automatic cassette handling and plate erasure

DR Detector (if DR configuration offered)

- Wireless flat-panel detector

- Active area: minimum 35 × 43 cm

- Pixel pitch: maximum 150 μm

- Matrix: minimum 2300 × 2800

- Dynamic range: minimum 16-bit

- DQE: minimum 65% at 0 lp/mm

- Wireless connectivity: Wi-Fi 802.11 a/b/g/n/ac dual-band

- Battery: minimum 8 hours or 300 exposures per charge

- IP rating: minimum IP54

- Impact-resistant construction

Central Image Acquisition and Processing Workstation

- Multi-core 64-bit CPU

- Minimum 16 GB RAM

- Minimum 2 TB storage

- Medical-grade preview monitor: minimum 19-inch, minimum 1280 × 1024

- Image processing tools: automatic exposure indicator, grayscale optimization, tissue equalization, edge enhancement, noise reduction, image reversal

- Batch processing and worklist management

- Image storage minimum 50,000 images

- DICOM 3.0 full compliance

Connectivity

- DICOM 3.0 (Send, Receive, MWL, MPPS, Storage Commit, Print)

- HL7 interface

- Ethernet 1 Gbps

- Wi-Fi for wireless detectors (if DR)

User Interface

- Multi-language support

- Configurable examination protocols

- Worklist integration

- Barcode scanning integration

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Connected load: as declared

Applicable Standards

- IEC 60601-1, 60601-1-2

- IEC 60601-2-54 (if DR)

- IEC 62304, 62366

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k)

## 11.3 Standard Accessories

Central Image Acquisition and Processing Workstation

- As per Section 2 specifications

- Ergonomic keyboard and optical mouse

- Uninterruptible Power Supply (UPS) with surge protection

High-Resolution Medical-Grade Preview Monitor

- Minimum 19-inch diagonal

- Minimum 1280 × 1024 resolution

- Medical-grade LCD/LED

- DICOM GSDF calibrated

Barcode Reader for Cassette/Patient Identification

- 1D and 2D barcode compatibility

- Wired USB or wireless

- Integrated with worklist and patient data population

DICOM Connectivity Software Licenses

- Full DICOM 3.0 services (Send, Receive, MWL, MPPS, Storage Commit, Print)

- HL7 interface license

- Minimum 3 concurrent DICOM connections supported

Manufacturer-Standard Quality Control Phantom

- Compatible with the offered CR/DR configuration

- Compliant with IEC 62220-1 or equivalent QA standard

- Storage case included

- Calibration certificate traceable to national/international standards

## 11.4 Installation Requirements

- Dedicated workstation area with adequate space

- Dedicated single-phase electrical supply

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS backup for workstation (minimum 30 minutes)

- Network connectivity to PACS/RIS/HIS (Gigabit Ethernet)

- Wi-Fi coverage in radiography areas (for DR configuration)

## 11.5 Documentation

- User/Service Manual, Parts Catalogue

- Factory Test Certificate, Electrical Safety Test Report (IEC 62353)

- Image Quality Verification Certificate

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 11.6 Training

- Clinical/technologist training: minimum 2 working days

- Biomedical engineering training: minimum 1 working day

- Refresher training at 6 months

## 11.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (including CR plates or DR detectors), labor, PM (minimum 2 visits/year), breakdown maintenance, software updates

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts availability: minimum 10 years post-installation

# 12. DICOM VIEWING / DIAGNOSTIC REPORTING WORKSTATION

Quantity: 8

## 12.1 Clinical & Functional Specifications

The system shall be a high-performance diagnostic reporting workstation intended for radiologists' primary interpretation of medical images across all modalities including CR/DR, CT, MRI, mammography, fluoroscopy, angiography, ultrasound, and nuclear medicine.

Each workstation shall provide DICOM-compliant image viewing, advanced 2D/3D post-processing, multi-modality comparison, structured reporting, voice dictation integration, and full connectivity to PACS/RIS/HIS.

The workstations shall be configured for high-throughput clinical reporting with ergonomic user interface and diagnostic-grade image display.

## 12.2 Engineering Technical Specifications

Workstation Hardware (CPU Tower)

- Multi-core 64-bit processor (minimum 8 cores)

- RAM: minimum 32 GB (upgradeable to 64 GB)

- Storage: minimum 1 TB SSD (system) + 4 TB HDD (data cache)

- Dedicated professional graphics card with minimum 8 GB VRAM, medical-imaging optimized

- Network: dual Gigabit Ethernet

- USB 3.0, DVD-RW, and optional Blu-ray

- Operating system: current-generation, medical-grade / hospital-IT compatible

Diagnostic Monitors

- Dual medical-grade grayscale or color diagnostic monitors

- Resolution: minimum 3 MP (2048 × 1536) — 5 MP for mammography workstations (bidder to declare which workstations are configured for mammography)

- Diagonal: minimum 21-inch

- Luminance: minimum 500 cd/m² calibrated (1000 cd/m² for mammography)

- DICOM GSDF (Part 14) calibrated with automatic calibration sensor

- Uniformity correction

- Backlight stabilization sensor

Optional Third Monitor (Worklist / Navigator)

- Standard color monitor: minimum 24-inch, minimum 1920 × 1080

- For worklist, RIS, and reporting interface

DICOM Viewing Software

- DICOM 3.0 full compliance (Send, Receive, Query/Retrieve, Storage Commit, MPPS, Structured Reports)

- Multi-modality viewing: CR, DR, CT, MRI, US, MG, NM, PET, XA, RF

- 2D tools: window/level presets, zoom, pan, measurements, annotations, MPR, MIP

- 3D tools: volume rendering, MIP, MinIP, curved MPR, virtual endoscopy

- Cine mode for dynamic studies

- Multi-planar reconstruction (MPR)

- Fusion imaging (CT/MRI/PET)

- Advanced clinical packages (optional, bidder to declare): cardiac analysis, vessel analysis, oncology follow-up, lung nodule analysis

- Hanging protocols with user customization

- Prior study comparison and side-by-side display

Reporting

- Integrated structured reporting

- Voice dictation microphone integration (compatible with common speech-recognition systems)

- Report templates library

- HL7 export to RIS/HIS

Cybersecurity and User Management

- Role-based user access control

- Audit logging

- Encryption of data at rest and in transit (AES-256 or equivalent)

- Automatic screen lock and session timeout

- Compliance with IEC 80001, HIPAA (or equivalent local data protection standards)

Connectivity

- DICOM 3.0

- HL7

- Ethernet 1 Gbps (dual)

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Connected load per workstation: typically ≤500 VA

- UPS with surge protection: minimum 15 minutes runtime per workstation

Applicable Standards

- IEC 60601-1 (for medical-grade monitors)

- IEC 62304 (software)

- IEC 62366 (usability)

- IEC 80001 (network security)

- ISO 13485, 14971

- DICOM 3.0, HL7

- CE MDR or US FDA 510(k) for diagnostic software

## 12.3 Standard Accessories

High-Performance CPU Workstation Tower

- As per Section 2 specifications

- Tool-less serviceable chassis

- Vibration-damped storage bays

Dual 3 MP or 5 MP Medical-Grade Diagnostic Monitors

- As per Section 2 specifications

- Dual-monitor adjustable ergonomic stand

- DICOM GSDF calibration with certificate

- Automatic calibration sensor integrated

Ergonomic Alphanumeric Keyboard and Optical Mouse

- Ergonomic design, wired or wireless

- Wipe-clean, disinfectant-compatible surfaces (silicone-covered or sealed keyboard preferred)

- Programmable function keys (optional)

- USB or wireless (2.4 GHz) connectivity

Dedicated Multi-Key Dictation Microphone

- Handheld dictation microphone with minimum 6 programmable function buttons

- Noise-cancelling directional microphone

- USB connectivity

- Compatible with common speech-recognition and reporting systems

- Ergonomic grip, wipe-clean surface

Uninterruptible Power Supply (UPS) with Surge Protection

- Online double-conversion or line-interactive UPS

- Rated capacity: minimum 1 kVA per workstation

- Runtime: minimum 15 minutes at full load

- Surge protection (input) and output filtering

- Audible and visual alarms

- Communication interface (USB/network) for automatic shutdown

- Sealed maintenance-free battery, minimum 3-year life

## 12.4 Installation Requirements

- Dedicated reporting rooms with controlled ambient lighting (dimmable, indirect)

- Ergonomic workstation furniture (procured separately)

- Dedicated single-phase electrical supply per workstation

- Dedicated earthing ≤1 ohm

- HVAC: 20—24°C, 40—60% humidity

- UPS backup per workstation (included as accessory)

- Network connectivity to PACS/RIS/HIS (Gigabit Ethernet, dual)

- Sound-attenuated environment for dictation

## 12.5 Documentation

- User/Service Manual, Parts Catalogue (English)

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Monitor Calibration Certificate (DICOM GSDF)

- Installation and SAT Report, Commissioning Report

- Preventive Maintenance Schedule

- DICOM Conformance Statement

- Cybersecurity/MDS2 Statement

## 12.6 Training

- Radiologist user training: minimum 3 working days (across nominated users)

- Advanced applications training: minimum 2 working days

- Biomedical engineering / IT training: minimum 2 working days

- Refresher training at 6 months

## 12.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (workstation, monitors, UPS batteries, accessories), labor, PM (minimum 2 visits/year), breakdown maintenance, software updates, monitor recalibration

- Uptime guarantee: minimum 95%

- Response within 24 hours; resolution within 72 hours

- Spare parts and software support: minimum 7 years post-installation

# 13. FILM VIEWER BOX / DIGITAL VIEWING STATION

Quantity: 5

## 13.1 Clinical & Functional Specifications

The system shall be a high-luminance LED-based multi-panel radiographic film viewer intended for reviewing conventional radiographic films of all sizes, and shall be suitable for use in wards, outpatient clinics, operating theatres, and clinical consultation areas.

Where specified, the system may alternatively be supplied as a digital viewing station with LCD/LED display for reviewing digital images (bidder to declare).

For this specification, the primary requirement is the conventional multi-panel LED film viewer.

## 13.2 Engineering Technical Specifications

Viewing Panel

- Illumination source: LED (Light-Emitting Diode) — long-life, low heat, energy efficient

- Configuration: multi-panel, minimum 4-panel (4-bank) design accommodating up to four 35 × 43 cm films side-by-side

- Total viewing area: minimum 140 cm × 43 cm (for 4-panel configuration)

- Luminance: minimum 3000 cd/m² (adjustable)

- Luminance uniformity: within ±15% across viewing surface

- Color temperature: 6500 K (D65) daylight simulation

- Diffuser: opal acrylic or equivalent for uniform illumination

- Life expectancy of LED source: minimum 50,000 hours

Film Gripping / Retention

- Integrated spring-loaded film rails (top and/or bottom)

- Compatible with standard film sizes: 18 × 24 cm, 24 × 30 cm, 35 × 35 cm, 35 × 43 cm

- Quick-load and quick-release design

- Film retention clips (minimum 8 per panel)

Dimming Controller

- Continuous dimming control (0—100%)

- Integrated switch on front panel or side-mounted

- Independent dimming per panel (preferred) or global dimming (acceptable)

Housing / Frame

- Powder-coated aluminum or steel frame

- Anti-glare bezel

- Wipe-clean surfaces, disinfectant-compatible

Mounting

- Wall-mounted with adjustable mounting hardware, OR

- Desktop base stand configuration

- Both mounting options shall be supplied

Electrical Specifications

- Input: 100—240 VAC, single-phase, 50/60 Hz

- Power consumption: typically ≤100 W (LED-based)

- IEC 60601-1 compliant power supply

Environmental Conditions

- Operating temperature: 10—40°C

- Operating humidity: 20—80%, non-condensing

- Passive cooling (no fans required for LED design)

Safety and Compliance

- IEC 60601-1 (general safety, if used as medical device)

- IEC 60601-1-2 (EMC)

- CE marking (LVD, EMC, RoHS)

## 13.3 Standard Accessories

High-Luminance LED Viewing Panels (Multi-Panel Configuration)

- As per Section 2 specifications

- Minimum 4-panel configuration

- LED illumination with 50,000-hour life expectancy

- Adjustable luminance up to minimum 3000 cd/m²

Integrated Film Gripping Rails and Clips

- Spring-loaded film retention rails

- Compatible with all standard radiographic film sizes

- Minimum 8 clips per panel

- Corrosion-resistant material

Continuous Dimming Controller Switch

- Front-panel or side-mounted controller

- Continuous 0—100% adjustment

- Ergonomic, wipe-clean surface

Desktop Base Stand and Wall-Mounting Hardware Kit

- Desktop base stand: powder-coated steel, stable non-tip design, non-slip feet

- Wall-mounting hardware kit: mounting bracket, expansion bolts, mounting template, adjustable spacers

- Both options supplied with the unit

## 13.4 Installation Requirements

- Wall or desktop location in ward, clinic, OT, or consultation area

- Adequate wall structural support for wall-mounted installation

- Standard single-phase electrical outlet (100—240 VAC)

- Ambient lighting control (dimmable or shaded) for optimal viewing contrast

- No dedicated HVAC or network connectivity required

## 13.5 Documentation

- User Manual (English)

- Installation Guide

- Factory Test Certificate

- Electrical Safety Test Report (IEC 62353)

- Certificate of Compliance (CE / equivalent)

- Preventive Maintenance Schedule (minimal — LED-based, low-maintenance)

## 13.6 Training

- User familiarization training: minimum 1 working day (grouped, covering all supplied units)

- Biomedical engineering familiarization: minimum 0.5 working day

## 13.7 Warranty

- Comprehensive warranty: minimum 2 years from commissioning

- Includes all parts (LED panels, dimming controller, mounting hardware), labor, breakdown maintenance

- LED panel replacement warranty: minimum 5 years or 50,000 operating hours (whichever earlier)

- Response within 48 hours; resolution within 5 working days

- Spare parts availability: minimum 7 years post-installation