Technical Specifications for Radiation Protection Equipment

Technical Tender Document

Table of Contents

1. Lead Apron / Thyroid Shield / Gonadal Shield	3

1.1 Clinical & Functional Specifications	3

1.2 Engineering Technical Specifications	3

1.3 Standard Accessories	5

1.4 Installation Requirements	6

1.5 Documentation	6

1.6 Training	6

1.7 Warranty	6

2. Radiation Survey Meter / Personal Dosimeter	8

2.1 Clinical & Functional Specifications	8

2.2 Engineering Technical Specifications	8

2.3 Standard Accessories	10

2.4 Installation Requirements	11

2.5 Documentation	12

2.6 Training	12

2.7 Warranty	12

# 1. Lead Apron / Thyroid Shield / Gonadal Shield

Quantity: 18

## 1.1 Clinical & Functional Specifications

The radiation protection garments shall be designed for personal protection of clinical, surgical, and technical staff against scattered ionizing radiation during diagnostic radiology, fluoroscopy, interventional cardiology, interventional radiology, cardiac catheterization, orthopaedic C-arm procedures, operating theatre imaging, and dental radiography. The set shall comprise a full-body lead apron, a thyroid collar shield, and a gonadal shield, each providing effective attenuation of scattered X-ray radiation within the diagnostic energy range typically encountered in hospital imaging environments (approximately 50—150 kVp).

The garments shall be lightweight, ergonomically designed, and comfortable for prolonged wear during long clinical procedures. They shall provide unrestricted freedom of movement, appropriate anatomical coverage, and even weight distribution to reduce operator fatigue and musculoskeletal strain. All shielding materials shall be non-toxic, flexible, tear-resistant, and suitable for repeated clinical use in a hospital environment. The outer covering shall be antimicrobial, fluid-resistant, and easily cleanable using standard hospital disinfectants.

## 1.2 Engineering Technical Specifications

Shielding Material

- Lead, lead-composite, or lead-free composite attenuation material of equivalent or better shielding performance

- Flexible, non-cracking, tear-resistant core material

- Uniform distribution of attenuating material without voids, gaps, or hot spots

- Free from asbestos and other prohibited substances

Lead Equivalence (Attenuation)

- Front protection: Minimum 0.50 mm Pb equivalent

- Back protection: Minimum 0.25 mm Pb equivalent

- Overlap area (front): Minimum 0.75—1.00 mm Pb equivalent (where applicable in wrap-around style)

- Thyroid shield: Minimum 0.50 mm Pb equivalent

- Gonadal shield: Minimum 0.50 mm Pb equivalent

- Attenuation performance tested across the diagnostic energy range of 50—150 kVp

Design & Configuration

- Apron style: Front-protection, wrap-around, vest-and-skirt, or coat-style — as selectable/configurable per user requirement

- Two-piece (vest + skirt) design preferred for even weight distribution on lumbar spine, where applicable

- Sizes: Selectable range including Small, Medium, Large, and Extra-Large to accommodate all clinical staff

- Adjustable shoulder, waist, and back-support systems for individualized fit

- Ergonomic weight-distribution design to reduce spinal loading

Outer Cover Material

- Durable, tear-resistant, fluid-repellent textile or vinyl/polyurethane-coated fabric

- Antimicrobial and antibacterial treatment

- Resistant to hospital-grade disinfectants including alcohol-based, chlorine-based, and quaternary ammonium cleaners

- Available in multiple selectable colors

- Non-fading, non-cracking, and non-peeling under repeated cleaning

Seams & Construction

- Reinforced double-stitched seams

- Sealed or heat-welded edges to prevent fluid ingress

- Reinforced closure points and fastening zones

Closures & Fasteners

- Hook-and-loop (Velcro-type), buckle, or magnetic closure systems

- Quick-release and adjustable

- Corrosion-resistant hardware

Weight (Generalized)

- Full apron: Approximately 3—7 kg depending on size, style, and lead equivalence

- Thyroid shield: Approximately 0.1—0.3 kg

- Gonadal shield: Approximately 0.2—0.5 kg

Coverage (Generalized)

- Apron: From shoulders to below knee level, appropriate to user height

- Thyroid shield: Complete coverage of thyroid, larynx, and upper sternum region

- Gonadal shield: Adequate anatomical coverage for gonadal region

Applicable Standards

- IEC 61331-1: Protective devices against diagnostic medical X-radiation — Determination of attenuation properties of materials

- IEC 61331-3: Protective devices against diagnostic medical X-radiation — Protective clothing, eyewear and shielding devices for patients

- ISO 13485: Medical devices — Quality management systems

- ISO 14971: Medical devices — Application of risk management

- CE marking / FDA clearance or equivalent international regulatory approval

- REACH / RoHS compliance for material safety

## 1.3 Standard Accessories

a) Heavy-Duty Wall-Mounted Hanger Storage Racks with Multi-Arm Swivel Bars

- Wall-mounted, heavy-duty steel or stainless-steel construction

- Powder-coated or corrosion-resistant finish suitable for hospital environment

- Multi-arm swivel/rotating configuration with minimum 4—6 hanging arms per unit

- Each arm capable of supporting minimum 10 kg load without deformation

- Ergonomic hanging profile designed to preserve apron shape and prevent creasing or cracking of shielding material

- Wall-mounting hardware and fixings included

- Suitable for storage of aprons, thyroid shields, and accessories together

- Quantity sufficient to accommodate all supplied aprons

b) Integrated Thyroid Shield Attachment Collar Straps with Hook-and-Loop Fasteners

- Compatible and interchangeable with all supplied aprons

- Adjustable neck circumference with hook-and-loop (Velcro-type) fasteners

- Soft, non-abrasive inner lining for comfort against neck skin

- Outer covering matching apron material specifications (fluid-resistant, antimicrobial, disinfectant-compatible)

- Minimum 0.50 mm Pb equivalent shielding

- Reusable, cleanable, and durable for repeated clinical use

c) Adjustable Cross-Over Back-Suspension Belts with Quick-Release Matrix Buckles

- Ergonomically designed cross-over/X-back configuration for optimal weight distribution

- Adjustable strap length to accommodate all user sizes

- Quick-release buckle system for rapid donning and doffing

- High-strength, corrosion-resistant buckle hardware

- Padded shoulder and back sections for wearer comfort

- Compatible with supplied aprons

d) Removable Protective Textile Exterior Storage Covers

- Breathable, dust-resistant textile fabric

- Full-length cover with closure mechanism (zipper, hook-and-loop, or drawstring)

- Machine-washable or wipeable

- Designed for fabric preservation during storage and transport

- One cover per supplied apron

- Labeling/identification pocket or panel where applicable

## 1.4 Installation Requirements

- No electrical, plumbing, or gas connections required

- Adequate wall space required for mounting of storage racks

- Wall structure to be verified for load-bearing capacity prior to installation

- Storage location shall be dry, well-ventilated, and away from direct sunlight, sharp objects, and heat sources

- Ambient temperature: Approximately 10—35 °C

- Relative humidity: 20—80 % non-condensing

- Supplier shall coordinate rack installation with hospital biomedical/facilities department

## 1.5 Documentation

The supplier shall provide, for each apron and accessory, the following documentation:

- User Manual (English)

- Care, cleaning, and disinfection instructions

- Lead equivalence certificate for each apron, thyroid shield, and gonadal shield

- Attenuation test report per IEC 61331 across specified kVp range

- Material safety and composition certificate

- Manufacturer's Certificate of Conformity

- CE / FDA / equivalent regulatory certificate

- ISO 13485 certificate of the manufacturer

- Warranty certificate

- Recommended inspection and integrity-testing schedule

- Installation report for storage racks

## 1.6 Training

The supplier shall provide the following on-site training at the hospital:

- Clinical user training on correct donning, doffing, and wearing of aprons and shields

- Training on inspection procedures for detecting cracks, tears, or shielding defects

- Training on cleaning, disinfection, and storage protocols

- Biomedical engineering training on periodic fluoroscopic/radiographic integrity testing of aprons

## 1.7 Warranty

- Minimum comprehensive warranty of two (2) years from the date of successful commissioning, or as specified by the purchaser

- Warranty shall cover:

-   Defects in material, workmanship, and shielding integrity

-   Cracking, tearing, or loss of attenuation performance under normal clinical use

-   Failure of closures, fasteners, and stitching

-   Replacement of defective items

-   Free replacement of any apron failing the integrity test within the warranty period

-   Technical support and after-sales service

# 2. Radiation Survey Meter / Personal Dosimeter

Quantity: 14

## 2.1 Clinical & Functional Specifications

The radiation survey meter / personal dosimeter shall be a portable, handheld, battery-operated instrument designed for the measurement, monitoring, and survey of ionizing radiation levels in hospital environments including diagnostic radiology, fluoroscopy, nuclear medicine, radiotherapy planning areas, interventional imaging suites, radioactive material storage areas, radiation protection surveys, contamination monitoring, and area monitoring.

The instrument shall accurately measure ambient dose equivalent rate and cumulative dose from X-ray, gamma, and (where configured) beta radiation. It shall provide real-time readout of radiation levels, audible and visual alarm indications when preset thresholds are exceeded, and the capability to log, store, and export measurement data for radiation safety records and regulatory reporting.

The device shall be intuitive to operate, robust for field and clinical use, and suitable for use by radiation safety officers, medical physicists, biomedical engineers, and trained clinical personnel.

## 2.2 Engineering Technical Specifications

Detector Type

- Geiger-Müller (GM) tube, ion chamber, scintillation detector, or equivalent or better detection technology

- Detector suitable for measurement of X-ray, gamma, and (where applicable) beta radiation

- Energy-compensated detector for accurate dose-equivalent measurement across diagnostic and therapeutic energy ranges

Measurement Parameters

- Ambient dose equivalent rate H\*(10)

- Accumulated/integrated dose

- Count rate (cps/cpm) where applicable

- Personal dose equivalent Hp(10) for personal dosimeter function

Dose Rate Measurement Range

- Minimum range: 0.1 µSv/h to at least 100 mSv/h (or equivalent broader range)

- Auto-ranging capability

Accumulated Dose Range

- Minimum: 0.01 µSv to at least 10 Sv

Energy Response Range

- Minimum photon energy range: 30 keV to 1.3 MeV or better

- Energy-compensated response within ±30 % across the specified range

Accuracy

- Within ±10 % to ±15 % under reference conditions

- Traceable to national/international radiation standards (e.g., PTB, NIST, or equivalent)

Response Time

- Fast response, typically less than 5 seconds at high dose rates

- Configurable/adjustable time constant

Display

- Backlit LCD, LED, or equivalent digital display

- Minimum display area suitable for clear readout of dose rate, accumulated dose, and unit

- Alphanumeric display with configurable units (Sv/h, R/h, cps, cpm)

- Bar graph or analogue-style indicator for rate visualization

User Interface

- Multi-function keypad or membrane keys

- Menu-driven operation with selectable languages (English mandatory)

- Configurable alarm thresholds for dose rate and accumulated dose

- Data logging with time and date stamp

Alarms

- Audible alarm: Minimum 85 dB at 30 cm

- Visual alarm: Flashing LED or display indicator

- Vibration alarm (where applicable)

- Configurable dose-rate and dose alarm thresholds

Data Memory & Logging

- Internal memory for minimum 1000 measurement records with time/date stamp

- Data export via USB, Bluetooth, or equivalent digital interface

- PC software for data download, analysis, and reporting (included)

Power Supply

- Rechargeable battery (Li-ion, NiMH, or equivalent), and/or standard replaceable batteries (AA/AAA)

- Battery runtime: Minimum 100 hours of continuous operation under typical conditions

- Charging time: Maximum 8 hours for full charge

- Low-battery indication

- Mains adapter/charger: 100—240 VAC, 50/60 Hz

Environmental Conditions

- Operating temperature: −10 °C to +50 °C

- Storage temperature: −20 °C to +60 °C

- Relative humidity: 10—90 % non-condensing

- IP rating: Minimum IP54 or equivalent for dust and splash resistance

Dimensions & Weight (Generalized)

- Handheld, portable configuration

- Weight: Approximately 0.3—1.5 kg including battery

- Ergonomic single-hand operation

Connectivity

- USB, RS-232, Bluetooth, or equivalent digital data interface

- Compatibility with standard PC operating systems (Windows)

Calibration

- Factory calibrated with traceable certificate

- Recommended annual recalibration

- Calibration adjustment accessible through service menu

Applicable Standards

- IEC 60846: Ambient and/or directional dose equivalent (rate) meters for X, gamma and beta radiation

- IEC 61526: Measurement of personal dose equivalents Hp(10) and Hp(0.07)

- IEC 61010-1: Safety requirements for electrical equipment for measurement, control, and laboratory use

- IEC 61000-4 series: EMC requirements

- ISO 4037: Reference X and gamma radiation for calibrating dosimeters

- ISO 13485 (manufacturer)

- CE marking / FDA clearance or equivalent regulatory approval

## 2.3 Standard Accessories

a) Geiger-Müller or Ion-Chamber Detection Probe Wand with Protective Cap

- External detection probe compatible with the base survey meter

- Geiger-Müller tube, ion chamber, or equivalent detection technology

- Energy-compensated for diagnostic energy range

- Ergonomic wand-style handle for surface and area scanning

- Protective cap for detector window

- Robust construction suitable for repeated clinical/field use

- Compatible connector for interface to main chassis

b) Coiled Stretch Interface Cable

- Coiled/spiral construction with minimum extended length of approximately 1.5—2 meters

- Shielded signal cable to minimize electromagnetic interference

- Robust strain-relief connectors at both ends

- Compatible with probe and main chassis

- Flexible and durable insulation suitable for repeated use

c) Rigid Plastic Transport and Containment Carrying Case with Foam Inserts

- Rigid, impact-resistant plastic (ABS, polypropylene, or equivalent) construction

- Custom-cut foam inserts to accommodate main unit, probe, cable, charger, and accessories

- Lockable latches and carry handle

- Dust and splash resistant when closed

- Suitable for transport and field storage

d) Electronic Calibration Verification Certificate Block / Standard Reference Tool

- Calibration check source or electronic reference tool for functional/response verification

- Traceable to national or international radiation standards

- Documented reference activity/response value with certificate

- Suitable for routine constancy check by the user

- Where a check source is provided, activity shall be below regulatory exemption limits and supplied with all required safety documentation

e) Heavy-Duty Protective Rubber Boot Enclosure with Wrist Strap Lanyard

- Protective silicone or rubber sleeve/boot fitted to the main instrument

- Shock- and drop-resistant to protect against accidental impact

- Non-slip texture for secure grip

- Integrated or attached wrist strap lanyard for drop prevention

- Chemical-resistant and cleanable with hospital disinfectants

- Does not obstruct display, keypad, or detector window

f) Rechargeable Battery Power Set with Desktop Charging Cradle Station

- Rechargeable Li-ion, NiMH, or equivalent battery pack matched to the instrument

- Desktop charging cradle with mains adapter (100—240 VAC, 50/60 Hz)

- LED charging status indicators

- Overcharge, over-discharge, and short-circuit protection

- Minimum one spare battery pack included

- Charging time: Maximum 8 hours for full charge

## 2.4 Installation Requirements

- Standard mains power outlet: 100—240 VAC, 50/60 Hz for charging cradle

- Proper earthing per IEC 60364 or equivalent

- No specific HVAC, water, drainage, or medical gas requirements

- Ambient operating environment as specified above

- Adequate workspace for storage of unit, accessories, and calibration source

- Radiation safety zone for calibration verification, where applicable

## 2.5 Documentation

The supplier shall provide:

- User Manual (English)

- Service Manual

- Parts Catalogue

- Factory Calibration Certificate traceable to national/international standards

- Factory Test Certificate

- Electrical Safety Test Report (IEC 61010)

- EMC Test Report

- Type Test Certificate per IEC 60846 / IEC 61526

- Preventive Maintenance Schedule

- Installation and Commissioning Report

- PC software installation media and license

- CE / FDA / equivalent regulatory certificate

- ISO 13485 certificate of the manufacturer

- Warranty certificate

## 2.6 Training

The supplier shall provide on-site training including:

- Clinical/user training on operation, measurement techniques, and interpretation of readings

- Application training on radiation survey and personal monitoring procedures

- Biomedical engineering training on configuration, alarm setup, and data management

- Preventive maintenance training including battery care and cleaning

- Basic troubleshooting and error-code interpretation

- Training on constancy check using the supplied reference source/tool

- Data download, reporting, and software operation training

## 2.7 Warranty

- Minimum comprehensive warranty of two (2) years from the date of successful commissioning, or as specified by the purchaser

- Warranty shall include:

-   Labour

-   Spare Parts (excluding consumable batteries after first year, unless specified otherwise)

-   Preventive Maintenance visits (minimum two per year)

-   Breakdown Maintenance with defined response time

-   Software updates and upgrades

-   Free annual recalibration during warranty period, with issuance of calibration certificate

-   Technical support (telephone/email)

-   Replacement of defective probe, cable, or accessories