HOSPITAL TECHNICAL TENDER SPECIFICATIONS

Department: Obstetrics, Gynaecology & Delivery Suite

Table of Contents

1. DELIVERY / LITHOTOMY BED — ELECTRIC, MULTI-POSITION	5

1.1 Clinical & Functional Specifications	5

1.2 Engineering Technical Specifications	5

1.3 Standard Accessories	6

1.4 Installation Requirements	7

1.5 Documentation	7

1.6 Training	7

1.7 Warranty	7

2. CARDIOTOCOGRAPH (CTG)	8

2.1 Clinical & Functional Specifications	8

2.2 Engineering Technical Specifications	8

2.3 Standard Accessories	9

2.4 Installation Requirements	9

2.5 Documentation	10

2.6 Training	10

2.7 Warranty	10

3. MATERNAL-FOETAL MONITOR (COMBINED)	11

3.1 Clinical & Functional Specifications	11

3.2 Engineering Technical Specifications	11

3.3 Standard Accessories	12

3.4 Installation Requirements	12

3.5 Documentation	13

3.6 Training	13

3.7 Warranty	13

4. GYNAECOLOGY OPERATING TABLE	14

4.1 Clinical & Functional Specifications	14

4.2 Engineering Technical Specifications	14

4.3 Standard Accessories	15

4.4 Installation Requirements	15

4.5 Documentation	15

4.6 Training	15

4.7 Warranty	16

5. OPD DIAGNOSTIC HYSTEROSCOPE	17

5.1 Clinical & Functional Specifications	17

5.2 Engineering Technical Specifications	17

5.3 Standard Accessories	17

5.4 Installation Requirements	18

5.5 Documentation	18

5.6 Training	18

5.7 Warranty	18

6. OPERATIVE HYSTEROSCOPE WITH ACCESSORIES	19

6.1 Clinical & Functional Specifications	19

6.2 Engineering Technical Specifications	19

6.3 Standard Accessories	19

6.4 Installation Requirements	20

6.5 Documentation	20

6.6 Training	20

6.7 Warranty	20

7. GYNAECOLOGICAL LAPAROSCOPIC TOWER (COMPLETE SET)	21

7.1 Clinical & Functional Specifications	21

7.2 Engineering Technical Specifications	21

7.3 Standard Accessories	22

7.4 Installation Requirements	22

7.5 Documentation	23

7.6 Training	23

7.7 Warranty	23

8. NEONATAL RESUSCITATION TROLLEY WITH RADIANT WARMER, SUCTION, O₂	24

8.1 Clinical & Functional Specifications	24

8.2 Engineering Technical Specifications	24

8.3 Standard Accessories	25

8.4 Installation Requirements	25

8.5 Documentation	26

8.6 Training	26

8.7 Warranty	26

9. LEEP UNIT	27

9.1 Clinical & Functional Specifications	27

9.2 Engineering Technical Specifications	27

9.3 Standard Accessories	28

9.4 Installation Requirements	28

9.5 Documentation	28

9.6 Training	28

9.7 Warranty	28

10. VACUUM EXTRACTOR WITH SET OF CUPS	29

10.1 Clinical & Functional Specifications	29

10.2 Engineering Technical Specifications	29

10.3 Standard Accessories	29

10.4 Installation Requirements	30

10.5 Documentation	30

10.6 Training	30

10.7 Warranty	30

11. VIDEO COLPOSCOPE WITH LED DISPLAY MONITOR	31

11.1 Clinical & Functional Specifications	31

11.2 Engineering Technical Specifications	31

11.3 Standard Accessories	32

11.4 Installation Requirements	32

11.5 Documentation	32

11.6 Training	32

11.7 Warranty	32

# 1. DELIVERY / LITHOTOMY BED — ELECTRIC, MULTI-POSITION

Quantity: 2

## 1.1 Clinical & Functional Specifications

The equipment shall be a fully electrically operated, multi-position obstetric delivery and lithotomy bed intended for use in labour rooms, delivery suites, and obstetric operating rooms of a tertiary-care teaching hospital. It shall be suitable for antenatal examination, labour, normal vaginal delivery, assisted delivery, minor obstetric/gynaecological procedures, post-delivery recovery, and emergency caesarean preparation. The bed shall permit smooth motorised transition between examination, labour, delivery, lithotomy, Trendelenburg, reverse Trendelenburg, and cardiac (semi-Fowler) positions. It shall allow rapid detachment/retraction of the leg section to convert from the labour position to the delivery position without patient repositioning. It shall support obese/bariatric patients within the specified safe working load and shall permit quick emergency CPR flat-back release.

## 1.2 Engineering Technical Specifications

General Construction

- Frame: High-strength epoxy powder-coated steel or equivalent corrosion-resistant metal structure with antibacterial finish

- Base cover: Seamless single-piece ABS/high-impact polymer, easy-to-clean, radius-edged design

- Safe working load: Minimum 250 kg (patient + accessories)

- Bed sections: Minimum three-section articulating platform (backrest, seat, retractable/detachable leg section)

Dimensions (Generalized)

- Overall length: Approximately 2000—2200 mm (with leg section)

- Overall width: Approximately 700—900 mm (excluding side rails)

- Height adjustment range: Minimum 600—950 mm (electrically adjustable)

- Backrest tilt: 0° to at least +75°

- Trendelenburg / reverse Trendelenburg: Minimum ±12°

- Leg section: Retractable, foldable, or fully detachable

Electrical Drive System

- Motorised height, backrest, tilt (Trendelenburg), and seat/thigh section adjustment via low-noise linear actuators

- Actuator protection rating: Minimum IPX4 (individual actuators IPX6 preferred)

- Motor input: 24 V DC low-voltage system fed from internal isolated PSU

Control System

- Handheld wired pendant remote with membrane keypad and hang-hook

- Integrated side-rail mounted control panel with lock-out function

- Emergency CPR quick-release: Dual mechanical + gas-spring assisted, single-action operation

- Position memory / auto-return function (where applicable)

Mattress System

- High-density polyurethane foam (minimum density 45 kg/m³) with waterproof, seamless, antistatic, antibacterial, flame-retardant cover

- Multi-section design matching bed articulation, radiofrequency-welded seams

Castor & Braking System

- Four heavy-duty twin-wheel casters, minimum 125 mm diameter, antistatic, non-marking

- Central braking pedal system operable from both sides with directional/steering caster

- Locking action: Total lock / directional / free

Power & Battery

- Input: 100—240 V AC, 50/60 Hz

- Internal rechargeable sealed lead-acid or Li-ion backup battery, minimum 20 full position cycles on full charge

- Battery low indicator and charging status LED

- Class I, Type B applied part per IEC 60601-1

Safety & Standards

- Compliant with IEC 60601-1 (general safety), IEC 60601-1-2 (EMC), IEC 60601-2-52 (medical beds)

- CE MDR marked / FDA cleared / equivalent regulatory approval

- Manufactured under ISO 13485 quality management

## 1.3 Standard Accessories

- High-density seamless waterproof antistatic mattress pad set — Multi-section polyurethane foam, minimum 80 mm thickness, density ≥45 kg/m³, PU-coated waterproof cover, antibacterial, antistatic, flame-retardant, wipeable with hospital disinfectants

- Detachable leg-section mattress pad — Matching material and thickness, quick-release, autoclavable cover where declared

- Padded Goepel-type lithotomy leg crutches (1 pair) — Stainless steel frame with high-density foam padding, heavy-duty multi-axis mounting clamps, height/angle adjustable, load capacity minimum 25 kg per crutch

- Ergonomic stainless steel pull handles/handgrips (1 pair) — AISI 304 stainless steel, ergonomic contour, autoclavable, secure locking mount to bed frame

- Sliding stainless steel fluid collection basin — AISI 304 stainless steel, seamless, minimum 3-litre capacity, integrated drainage pipe with silicone tubing

- Handheld wired remote control unit — IPX4 rated, membrane keypad, coiled cable, illuminated icons, integrated side-rail control switches with lock-out

- Dual mechanical/gas-spring emergency CPR quick-release levers — Single-action, both sides accessible, tested for minimum 10,000 cycles

- Central braking and steering pedal system — Acts on all four casters, minimum three positions (brake/neutral/steer)

- Integrated rechargeable backup battery system with power cord — Sealed maintenance-free, minimum 20 cycle operation, hospital-grade IEC power cord, minimum 3 m length

## 1.4 Installation Requirements

- Electrical supply: 100—240 V AC, 50/60 Hz, single-phase, grounded 3-pin hospital-grade socket

- Room space: Minimum 3.5 m × 3.5 m clearance

- Flooring: Level, load-bearing minimum 400 kg

- Ambient conditions: Temperature 10—40 °C, humidity 30—75 % RH (non-condensing)

- Proper earthing per IEC 60364-7-710 (medical location Group 1)

## 1.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Preventive Maintenance Schedule, Installation & Commissioning Reports, CE/FDA/regulatory certificate, ISO 13485 certificate

## 1.6 Training

Clinical training for nursing/midwifery staff, biomedical engineering training, preventive maintenance training, basic troubleshooting, on-site application training — minimum 2 sessions

## 1.7 Warranty

Comprehensive minimum 2 years warranty including labour, all spare parts, preventive maintenance visits (minimum 2 per year), breakdown maintenance, software updates, and technical support.

# 2. CARDIOTOCOGRAPH (CTG)

Quantity: 3

## 2.1 Clinical & Functional Specifications

The equipment shall be a microprocessor-based antepartum and intrapartum fetal monitor for continuous non-invasive monitoring of fetal heart rate (FHR), uterine activity (TOCO), and fetal movement in singleton and twin pregnancies. It shall provide real-time graphical and numerical display, on-screen and printed CTG trace, automated event marking, configurable audible/visual alarms, and shall support NST (non-stress test) auto-interpretation where available. It shall be suitable for antenatal clinics, high-risk pregnancy units, and labour rooms.

## 2.2 Engineering Technical Specifications

Display

- Color TFT LCD or equivalent, minimum 7-inch diagonal, foldable/tiltable

- Minimum resolution 800 × 480 pixels

- Simultaneous display of FHR1, FHR2 (twin), TOCO, fetal movement, and numeric parameters

FHR (Fetal Heart Rate) Channel

- Method: Pulsed Doppler ultrasound

- Ultrasound frequency: Approximately 1 MHz

- Measurement range: Minimum 50—240 bpm

- Accuracy: ±2 bpm or better

- Resolution: 1 bpm

- Twin monitoring: Independent FHR1 and FHR2 with anti-cross-channel algorithm

TOCO (Uterine Activity) Channel

- External TOCO transducer, non-invasive

- Measurement range: 0—100 relative units

- Baseline auto-zero function

Fetal Movement

- Automatic detection (AFM) via Doppler signal, plus manual patient-controlled event marker

Alarms

- Configurable high/low FHR limits, audible + visual, minimum three priority levels per IEC 60601-1-8

Recorder / Printer

- Integrated thermal matrix / thermal array printer

- Paper width: Standard 150 mm Z-fold

- Selectable speeds: 1, 2, 3 cm/min

Data / Connectivity

- Internal memory: Minimum 24 hours of trace storage

- USB, LAN/Ethernet; HL7 support for hospital information system integration

- Central monitoring system compatibility

Power

- Input: 100—240 V AC, 50/60 Hz

- Internal Li-ion rechargeable battery, minimum 2 hours continuous operation, charging time maximum 8 hours

- Class I, Type BF applied parts

Environmental

- Operating temperature: 10—40 °C

- Humidity: 30—85 % RH non-condensing

Dimensions / Weight

- Compact tabletop / trolley configuration

- Weight: Approximately 4—8 kg

Standards

- IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-1-8 (alarms), IEC 60601-2-37 (ultrasound), IEC 62304 (software), ISO 13485, ISO 14971, CE MDR / FDA / equivalent

## 2.3 Standard Accessories

- High-sensitivity 1 MHz waterproof ultrasound transducers (2 pieces) — Multi-crystal Doppler, IPX7 or better, reusable, cable strain-relief, minimum 2.5 m cable, hospital-disinfectant compatible

- High-sensitivity external TOCO transducer — Waterproof (IPX7), reusable, matching connector, minimum 2.5 m cable

- Patient-controlled fetal movement event marker — Ergonomic push-button, coiled cable, matching connector

- Adjustable elastic transducer abdominal belts (set of 3) — Latex-free, washable, elastic knit, adjustable length up to 1.2 m

- High-resolution color LCD folding screen — See main specs

- Integrated thermal matrix slot printer — 150 mm paper width, high-resolution thermal array

- Z-fold thermal chart paper (2 packs) — Standard 150 mm CTG grid, obstetric-grade thermal paper

- Mains power adapter and battery backup pack — Medical-grade adapter, hospital-grade cord, matching Li-ion battery

## 2.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, hospital-grade earthed socket

- Tabletop or optional mobile trolley

- Network point for HL7/central station (where required)

## 2.5 Documentation

Full documentation package as per standard requirements.

## 2.6 Training

Clinical + biomedical + PM + application training.

## 2.7 Warranty

Minimum 2 years comprehensive warranty.

# 3. MATERNAL-FOETAL MONITOR (COMBINED)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The equipment shall be a combined maternal and foetal patient monitor providing simultaneous continuous monitoring of maternal vital signs (ECG, NIBP, SpO₂, temperature, respiration) together with foetal heart rate (single or twin) and uterine activity. It shall be used in high-risk labour rooms, obstetric HDU, and obstetric operating rooms. It shall provide unified numeric and waveform display, arrhythmia analysis, ST segment analysis (where available), configurable multi-level alarms, trend storage, and HL7 connectivity.

## 3.2 Engineering Technical Specifications

Display

- Color TFT LCD, minimum 12-inch diagonal, touchscreen preferred

- Minimum resolution 1280 × 800 pixels

- Simultaneous display of minimum 8 waveforms

Maternal ECG

- Lead configuration: Selectable 3-lead / 5-lead

- Bandwidth: Minimum 0.05—150 Hz (diagnostic mode)

- Sampling frequency: Minimum 500 samples/s/channel

- HR range: 15—300 bpm, accuracy ±1 % or ±1 bpm

- Arrhythmia detection with lethal arrhythmia alarm

NIBP (Non-Invasive Blood Pressure)

- Method: Oscillometric

- Modes: Manual, auto (selectable interval), STAT, continuous

- Range: Systolic 30—260 mmHg, Diastolic 10—220 mmHg, MAP 20—235 mmHg

- Accuracy: ±3 mmHg or ±2 % (whichever greater)

- Overpressure protection

SpO₂

- Range: 0—100 %; Accuracy: ±2 % (70—100 %)

- Pulse rate: 25—250 bpm

Temperature

- Range: 0—50 °C, accuracy ±0.1 °C

- Dual-channel capability

Respiration

- Impedance method via ECG leads, range 0—150 rpm

Fetal Channels

- FHR1, FHR2 (twin), TOCO, fetal movement (AFM + manual marker)

- FHR method: 1 MHz pulsed Doppler; range 50—240 bpm; accuracy ±2 bpm

Alarms

- Three-level priority per IEC 60601-1-8, configurable, audible + visual

Data & Connectivity

- Trend storage: Minimum 96 hours full disclosure

- Interfaces: Ethernet, USB, wireless (optional), HL7, integrated recorder

- Central monitoring system compatibility

Recorder

- Integrated thermal strip chart, minimum 50 mm width

Power

- 100—240 V AC, 50/60 Hz

- Rechargeable Li-ion battery, minimum 2 hours continuous operation

Standards

- IEC 60601-1, -1-2, -1-8, -2-27 (ECG), -2-30 (NIBP), -2-49 (multi-function), -2-37 (US), -2-27, IEC 62304, ISO 13485, CE MDR / FDA

## 3.3 Standard Accessories

- 3/5-lead ECG patient cable and reusable lead wire set — Trunk cable + reusable lead wires, snap/clip type, defibrillation-proof, matching connector, minimum 3 m

- Reusable adult SpO₂ finger sensor with extension cable — Silicone or hard-shell, adult finger, reusable, minimum 2 m cable, matching connector

- Reusable adult NIBP cuff with air hose — Nylon/PU, latex-free, size adult regular, matching double-tube hose, minimum 2 m

- Reusable skin temperature probe — YSI-400 series compatible or equivalent, range 0—50 °C, minimum 2 m cable

- High-sensitivity waterproof ultrasound transducers (2 pieces) — 1 MHz Doppler, IPX7 or better, matching connector

- High-sensitivity external TOCO transducer — Waterproof, matching connector

- Patient event marker button — Coiled cable, matching connector

- Adjustable elastic transducer belts (set of 4) — Latex-free, washable

- Integrated thermal strip chart recorder with paper roll — Minimum 50 mm width

- High-capacity rechargeable Li-ion battery pack — Medical grade, minimum 2 hours runtime

## 3.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, hospital-grade socket

- UPS backup recommended

- LAN point for HL7/central monitoring

## 3.5 Documentation

Full documentation package.

## 3.6 Training

Clinical + biomedical + PM + application.

## 3.7 Warranty

Minimum 2 years comprehensive.

# 4. GYNAECOLOGY OPERATING TABLE

Quantity: 2

## 4.1 Clinical & Functional Specifications

The equipment shall be a fully electrohydraulic multi-purpose operating table specifically configured for gynaecological, obstetric, urogynaecological, and endoscopic procedures. It shall support open gynaecology surgery, laparoscopy, hysteroscopy, colposcopy, LEEP, dilatation & curettage, and other perineal/lithotomy procedures. The table shall provide motorised height, tilt (Trendelenburg / reverse), lateral tilt, backrest, and seat/thigh section adjustments with lithotomy configuration and retractable/detachable leg section. Radiolucent tabletop suitable for C-arm/mobile X-ray imaging shall be provided.

## 4.2 Engineering Technical Specifications

Tabletop Construction

- Multi-section (minimum 4 sections: head, back, seat, leg) radiolucent tabletop

- Material: Carbon-fibre reinforced polymer / high-pressure laminate or equivalent radiolucent material

- X-ray translucent full length for C-arm imaging

Safe Working Load

- Minimum 250 kg (normal use), 135 kg fully articulated

Dimensions & Movements

- Overall length: Approximately 2000—2200 mm

- Width: Approximately 500—600 mm (excluding side rails)

- Height range: Minimum 600—1050 mm

- Backrest: −40° to +80°

- Trendelenburg / Reverse Trendelenburg: Minimum ±25°

- Lateral tilt: Minimum ±20°

- Leg section: Retractable and detachable; adjustable ±90°

- Longitudinal slide: Minimum 250 mm for C-arm positioning

Drive System

- Electro-hydraulic actuators, low noise, IPX4 or better

- 24 V DC low-voltage system with isolated PSU

- Emergency mechanical override for all critical movements

Control

- Handheld wired remote pendant, membrane keypad, IPX4

- Auxiliary column-mounted override panel

- Memory positions (minimum 3 programmable)

Base & Casters

- Retractable or fixed casters with central hydraulic/electric brake lock

- Antistatic non-marking casters, minimum 125 mm

Power

- Input 100—240 V AC, 50/60 Hz

- Rechargeable sealed battery, minimum 40 full cycles on full charge

Standards

- IEC 60601-1, IEC 60601-2-46 (operating tables), IEC 60601-1-2, ISO 13485, CE MDR / FDA

## 4.3 Standard Accessories

- High-density multi-section antistatic memory-foam mattress pad set — Minimum 80 mm thickness, density ≥45 kg/m³, radiolucent, seamless PU cover, antibacterial, flame-retardant

- Handheld wired remote control unit + auxiliary override panel — IPX4, membrane keypad

- Padded Goepel-type lithotomy leg holders (1 pair) — Multi-axis heavy-duty mounting clamps, height/angle adjustable, load capacity minimum 25 kg each

- Adjustable swivel armboards with fixation straps (1 pair) — Radiolucent, padded, ±180° swivel, clamp-mounted

- Sliding stainless steel waste collection drawer/basin with drain tube — AISI 304, seamless, minimum 3 L, silicone drain tube

- Padded lateral shoulder support braces (1 pair) — Adjustable, clamp-mounted, autoclavable pads

- Adjustable anaesthesia screen frame with clamp — Stainless steel, foldable

- Reusable body restraint strap with metal buckle — Wipeable, adjustable up to 2 m

- Integrated rechargeable backup battery + hospital-grade power cord — Sealed maintenance-free, minimum 3 m cord

## 4.4 Installation Requirements

- 100—240 V AC, 50/60 Hz hospital-grade socket, dedicated circuit

- Level OT floor, load-bearing minimum 500 kg

- Minimum room clearance 4 m × 4 m

- Ambient temp 15—35 °C

## 4.5 Documentation

Full documentation package.

## 4.6 Training

Clinical + biomedical + PM + application.

## 4.7 Warranty

Minimum 2 years comprehensive.

# 5. OPD DIAGNOSTIC HYSTEROSCOPE

Quantity: 1

## 5.1 Clinical & Functional Specifications

The set shall be a complete rigid diagnostic hysteroscopy system for office/OPD use, permitting direct visualisation of the endocervical canal and uterine cavity for diagnostic evaluation of abnormal uterine bleeding, infertility, intrauterine pathology, IUD localisation, and endometrial assessment. The set shall support continuous-flow saline distension and be compatible with standard fibre-optic cold light sources and imaging systems.

## 5.2 Engineering Technical Specifications

Optical System

- Rigid rod-lens (Hopkins-type) or equivalent high-resolution optics

- Field of view: Minimum 80°

- Direction of view: 0° and 30° (both provided)

- Working length: Minimum 250 mm (typical 302 mm)

- Optics diameter: Approximately 2.7—3.0 mm

Sheath

- Continuous-flow diagnostic sheath, outer diameter 3.0—4.0 mm

- Autoclavable stainless steel construction

- Inflow/outflow stopcocks with Luer-lock connectors

Light Transmission

- Standard fibre-optic light post with interchangeable adapters (Storz / Olympus / Wolf / ACMI universal compatibility)

Compatibility

- Compatible with standard cold light sources (150 W halogen / xenon / LED)

- Compatible with standard endoscopic camera systems

Sterilization

- Fully autoclavable at 134 °C, ETO compatible, immersion cleaning compatible

Standards

- ISO 8600 series (endoscopes), IEC 60601-1, ISO 13485, CE MDR / FDA

## 5.3 Standard Accessories

- Rigid 0° diagnostic hysteroscope sheath (OD 3—4 mm) — Stainless steel, autoclavable, continuous-flow, high-resolution optics

- Rigid 30° diagnostic hysteroscope sheath — Same construction, 30° oblique view

- Integrated fibre-optic light guide cable adapter plugs — Universal adapters (Storz/Olympus/Wolf/ACMI)

- Autoclavable diagnostic continuous-flow sheath assembly — Dual-channel inflow/outflow, Luer-lock stopcocks

- Heavy-duty fibre-optic light guide cable (2.5—3 m) — Minimum 3.5 mm active fibre bundle, autoclavable, kink-resistant

- Watertight protective cleaning cap set — Autoclavable, colour-coded

- Rigid perforated sterilization and storage container tray — Stainless steel or high-temp polymer, autoclavable at 134 °C, protective silicone holders

## 5.4 Installation Requirements

- Compatible with existing OT/OPD cold light source and camera system

- No dedicated electrical/environmental requirements

## 5.5 Documentation

Full documentation package.

## 5.6 Training

Clinical + biomedical + reprocessing/sterilization training.

## 5.7 Warranty

Minimum 2 years comprehensive including optics.

# 6. OPERATIVE HYSTEROSCOPE WITH ACCESSORIES

Quantity: 1

## 6.1 Clinical & Functional Specifications

The set shall be a complete rigid operative hysteroscopy system permitting simultaneous visualisation and intrauterine surgical intervention. It shall support endometrial biopsy, polypectomy, myomectomy, adhesiolysis, IUD removal, septum resection, and endometrial ablation using semi-rigid mechanical instruments and monopolar/bipolar high-frequency electrodes. Continuous-flow irrigation shall be supported for clear operative field.

## 6.2 Engineering Technical Specifications

Optical System

- Rigid rod-lens optics, 30° direction of view

- Field of view: Minimum 80°

- Working length: Approximately 300 mm

- Optics diameter: Approximately 3.0—4.0 mm

Operative Sheath

- Continuous-flow inner/outer sheath assembly

- Outer diameter: Approximately 7—9 mm

- Working channel: Minimum 5 Fr (1.6 mm)

- Autoclavable stainless steel

Instruments (Semi-Rigid, 5 Fr)

- Biopsy forceps, grasping forceps, scissors — autoclavable stainless steel

HF Electrodes

- Monopolar/bipolar compatibility

- Standard HF connector compatible with common electrosurgical generators

- Autoclavable/reusable

Light & Compatibility

- Universal fibre-optic light post

- Compatible with standard cold light sources and camera systems

Standards

- ISO 8600, IEC 60601-1, IEC 60601-2-2 (HF surgery, for electrodes), ISO 13485, CE MDR / FDA

## 6.3 Standard Accessories

- Rigid 30° operative hysteroscope optics tube — High-resolution rod-lens, autoclavable

- Continuous-flow operative inner/outer sheath assembly with 5 Fr working channel — Stainless steel, autoclavable

- Semi-rigid operative biopsy forceps (5 Fr) — Stainless steel, autoclavable, minimum 350 mm length

- Semi-rigid operative grasping forceps (5 Fr) — Same construction

- Semi-rigid operative scissors (5 Fr) — Same construction

- HF monopolar/bipolar resection electrodes (set of 2) — Loop/ball, insulated, compatible with standard HF generators

- Heavy-duty fibre-optic light guide cable — 2.5—3 m, minimum 3.5 mm active fibre bundle

- Direct-connect fluid inflow/outflow stopcocks — Luer-lock, autoclavable

- Watertight leakage test adapter — Compatible with reprocessing station

- Heavy-duty perforated stainless steel sterilization tray — Autoclavable at 134 °C, silicone instrument holders

## 6.4 Installation Requirements

- Compatible with existing OT cold light source, camera, HF generator, and irrigation system

## 6.5 Documentation

Full documentation package.

## 6.6 Training

Clinical + biomedical + sterilization training.

## 6.7 Warranty

Minimum 2 years comprehensive including optics.

# 7. GYNAECOLOGICAL LAPAROSCOPIC TOWER (COMPLETE SET)

Quantity: 1

## 7.1 Clinical & Functional Specifications

The equipment shall be a complete high-definition gynaecological laparoscopic imaging and insufflation system for diagnostic and operative laparoscopy including tubal ligation, ectopic surgery, ovarian cystectomy, myomectomy, hysterectomy, and endometriosis surgery. The tower shall integrate HD/4K imaging chain, high-intensity LED illumination, high-flow CO₂ insufflation, combined irrigation/suction, and standard reusable laparoscopic instruments on a mobile cart with isolation transformer.

## 7.2 Engineering Technical Specifications

Display Monitor

- Medical-grade LCD/OLED, minimum 32-inch diagonal (34-inch or larger where specified)

- Resolution: Full HD 1920 × 1080 minimum (4K UHD 3840 × 2160 preferred)

- Contrast ratio: Minimum 1000:1

- Wide viewing angle ≥170°

- Multiple inputs: SDI, DVI, HDMI, VGA

- Compliant with IEC 60601-1

Camera Control Unit (CCU) & Camera Head

- 3-CCD/3-CMOS full HD or 4K

- Resolution: Full HD minimum, 4K preferred

- Autoclavable or soakable camera head with optical zoom coupler (focal length \~15—30 mm)

- Integrated white balance, gain control, digital zoom

- Programmable head buttons

Light Source

- High-intensity medical LED

- Colour temperature: Approximately 5500—6500 K

- Output: Minimum 200 W equivalent (LED), life ≥30,000 hours

- Adjustable intensity, integrated in processor console (as per configuration)

CO₂ Insufflator

- Electronic high-flow automated

- Flow rate: Minimum 30 L/min, up to 45 L/min

- Pressure range: Selectable 3—30 mmHg

- Automatic pressure regulation with overpressure alarm

- Gas heating (optional/available)

Irrigation / Suction Pump

- Combined module

- Irrigation flow rate: Minimum 1.0 L/min adjustable

- Suction vacuum: Minimum −0.7 bar

- Foot-switch controlled

Mobile Cart

- Heavy-duty stainless steel / powder-coated steel

- Minimum 4 shelves, lockable antistatic casters

- Integrated medical-grade isolation transformer (minimum 1000 VA, per IEC 60601-1)

- Equipotential earth bonding

Reusable Instruments

- Autoclavable stainless steel: Maryland dissector, atraumatic graspers, scissors, needle holders (standard 5 mm × 330 mm)

- Autoclavable 5 mm and 10 mm trocar-cannula sets with insufflation valves

Standards

- IEC 60601-1, -1-2, -2-2 (HF where applicable), -2-18 (endoscopic), ISO 8600 series, ISO 13485, CE MDR / FDA

## 7.3 Standard Accessories

- Medical-grade LCD/OLED monitor, minimum 32—34 inch, Full HD/4K — See main specs

- 3-Chip HD/4K camera control head with optical zoom couplers — See main specs

- Combined HD video processor + medical LED light source console — Integrated unit, single-power, touchscreen interface

- High-flow automated electronic CO₂ insufflator (30—45 L/min) — See main specs

- High-pressure CO₂ inlet regulator with connecting hose — Standard medical yoke/pin-index, 2-stage regulator

- Combined laparoscopic irrigation/suction pump module — Foot-switch controlled

- Reusable fibre-optic light guide cable — Minimum 4.8 mm active fibre bundle, 2.5—3 m, autoclavable

- Heavy-duty mobile equipment tower cart with lockable casters + isolation transformer — Minimum 4 shelves, ≥1000 VA transformer

- Reusable laparoscopic hand instrument set — 5 mm × 330 mm, autoclavable stainless steel, includes Maryland dissectors, atraumatic graspers, scissors, needle holders

- Autoclavable trocar-cannula set (5 mm & 10 mm) — Reusable, insufflation valves, obturator included

## 7.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, minimum 20 A dedicated hospital-grade circuit

- UPS recommended (minimum 3 kVA)

- CO₂ cylinder / pipeline supply, medical-grade

- OT space minimum 6 m² for tower positioning

- Ambient temp 15—30 °C, humidity 30—75 %

## 7.5 Documentation

Full documentation package.

## 7.6 Training

Clinical + biomedical + PM + application (minimum 5 days on-site).

## 7.7 Warranty

Minimum 2 years comprehensive including camera head and light source LED module.

# 8. NEONATAL RESUSCITATION TROLLEY WITH RADIANT WARMER, SUCTION, O₂

Quantity: 2

## 8.1 Clinical & Functional Specifications

The equipment shall be a mobile open-care neonatal resuscitation platform with overhead radiant warmer for immediate stabilisation and resuscitation of newborns in delivery rooms and operating theatres. It shall integrate radiant heating, servo-controlled temperature, examination lamp, high-vacuum suction, oxygen delivery, T-piece resuscitator, weigh scale, and Apgar timer.

## 8.2 Engineering Technical Specifications

Radiant Heater

- Ceramic or quartz infrared, microprocessor-controlled

- Output: Minimum 500 W

- Modes: Manual (0—100 % power) and servo (skin temperature controlled)

- Servo range: 34.0—37.5 °C, resolution 0.1 °C

- Heater height above mattress: Approximately 700—900 mm

- Pre-warm mode

Skin Temperature Sensor

- Range: 25—45 °C, accuracy ±0.2 °C

- Alarms: Sensor failure, high/low temperature, deviation

Baby Platform

- Radiolucent examination surface with X-ray tray

- Dimensions: Approximately 600 × 700 mm

- Foldable transparent acrylic side shields (minimum 3 sides), height ≥150 mm

- Tilt: ±10° Trendelenburg/reverse

Suction Module

- Motorised high-vacuum diaphragm/rotary pump

- Vacuum range: 0 to −0.85 bar (0 to −650 mmHg), adjustable

- Autoclavable collection jar, minimum 500 mL, with overflow protection

Oxygen Delivery

- Integrated flowmeter: 0—15 L/min, calibrated

- Pipeline hose (medical O₂ probe) with NIST/DISS connector as per hospital standard

T-Piece Resuscitator (Neopuff-type)

- Adjustable PIP: 0—80 cmH₂O

- Adjustable PEEP: 0—20 cmH₂O

- Manometer indication

Weigh Scale

- Electronic, integrated

- Range: 0—10 kg, resolution 5—10 g

Examination Lamp

- LED spotlight, ≥30,000 lux at 500 mm, colour temperature \~4000 K, flexible arm

Apgar Timer

- Integrated digital, minimum 10 minutes

Power

- 100—240 V AC, 50/60 Hz, hospital-grade

- Class I, Type BF

Mobility

- Heavy-duty antistatic casters (≥125 mm), central brake, non-marking

- Two lockable drawers

Standards

- IEC 60601-1, -1-2, -2-19 (infant incubators/warmers), -2-21 (radiant warmers), ISO 13485, CE MDR / FDA

## 8.3 Standard Accessories

- Ceramic/quartz microprocessor-controlled IR overhead radiant heating module — See main specs

- Fixed examination mattress platform with folding transparent acrylic side shields — Radiolucent, X-ray tray, ≥150 mm shield height

- Integrated electronic baby weigh scale module — 0—10 kg, resolution ≤10 g

- Reusable skin temperature probe with attachment patches — Range 25—45 °C, accuracy ±0.2 °C, minimum 2 m cable

- Integrated motorised high-vacuum suction module with autoclavable collection jar — ≥−0.85 bar, ≥500 mL jar

- Integrated oxygen flowmeter with pipeline hose — 0—15 L/min, medical-grade connector

- T-piece infant resuscitator (Neopuff-type) with pressure gauge — PIP 0—80, PEEP 0—20 cmH₂O

- Adjustable IV pole + dual lockable storage drawers — Stainless steel IV pole, height-adjustable

- LED examination spotlight procedure lamp — ≥30,000 lux @ 500 mm, flexible arm

## 8.4 Installation Requirements

- 100—240 V AC hospital-grade socket

- Medical O₂ pipeline outlet

- Minimum 2 m × 2 m clear floor space

- Ambient temp 20—30 °C

## 8.5 Documentation

Full documentation package.

## 8.6 Training

Clinical + biomedical + PM + application.

## 8.7 Warranty

Minimum 2 years comprehensive.

# 9. LEEP UNIT

Quantity: 1

## 9.1 Clinical & Functional Specifications

The equipment shall be a high-frequency electrosurgical loop electrosurgical excision procedure (LEEP/LLETZ) system for treatment of cervical intraepithelial neoplasia (CIN), cervical polyps, and other cervical/vaginal lesions. It shall provide precise cutting and coagulation with minimal thermal artefact, integrated smoke evacuation, and full range of reusable electrodes.

## 9.2 Engineering Technical Specifications

HF Generator

- Microprocessor-controlled, isolated output

- Operating frequency: Approximately 350—500 kHz

- Modes: Pure cut, blended cut (multiple levels), coagulation (soft/forced/spray), bipolar

- Output power (maximum):

-   Cut: Minimum 80 W into 500 Ω

-   Coagulation: Minimum 60 W

-   Bipolar: Minimum 50 W

- Power adjustment: 1 W increments

- Return electrode monitoring (REM/CQM) with contact quality

- Digital display of mode and power

- Audible activation tone

Smoke Evacuator

- Integrated or console-mounted

- Flow rate: Minimum 1000 L/min at nozzle

- Filtration: HEPA (≥99.97 % @ 0.3 µm) + ULPA + activated carbon

- Auto-activation with HF generator

Safety

- Isolated output per IEC 60601-2-2

- Multiple safety alarms (open circuit, REM, over-temperature)

Power

- 100—240 V AC, 50/60 Hz

- Class I, Type CF applied parts

Standards

- IEC 60601-1, -1-2, -2-2 (HF surgery), ISO 13485, CE MDR / FDA

## 9.3 Standard Accessories

- HF electrosurgical LEEP/LLETZ generator unit — See main specs

- High-volume smoke evacuator with HEPA/ULPA filter — ≥1000 L/min, ≥99.97 % filtration

- Reusable LEEP loop electrodes (10×10 mm, 15×10 mm, 20×15 mm) — Tungsten wire loops, autoclavable insulated shafts

- Reusable ball/coagulation electrodes (3 mm, 5 mm) — Stainless steel, autoclavable

- Reusable insulated speculum with smoke evacuation channel — Stainless steel, insulated coating, integrated evacuation port

- Reusable silicone patient return pad with cable — REM-compatible split-pad, minimum 3 m cable

- Waterproof double-pedal foot switch — IPX8, cut/coag pedals

- Mobile utility cart with lockable casters and accessory storage — Powder-coated steel, minimum 2 shelves

## 9.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, hospital-grade dedicated socket

- Ventilated procedure room

- Equipotential earthing per IEC 60364-7-710

## 9.5 Documentation

Full documentation package.

## 9.6 Training

Clinical + biomedical + PM + application.

## 9.7 Warranty

Minimum 2 years comprehensive.

# 10. VACUUM EXTRACTOR WITH SET OF CUPS

Quantity: 1

## 10.1 Clinical & Functional Specifications

The equipment shall be an electric obstetric vacuum extractor system for assisted vaginal delivery. It shall provide controlled, accurate negative pressure via reusable metal (Malmström-type) and silicone cups with continuous pressure display, overflow protection, and foot-switch control.

## 10.2 Engineering Technical Specifications

Vacuum Pump

- Oil-free electric diaphragm/rotary pump

- Maximum vacuum: Minimum −0.85 bar (−650 mmHg / −85 kPa)

- Adjustable, low-noise operation (\<55 dB)

- Duty cycle: Continuous

Pressure Gauge / Display

- Dual-scale analog gauge (kPa/mmHg) with obstetric colour zones, or digital display

- Accuracy: ±5 % of full scale

- Precise low-vacuum control for obstetric application (0 to −0.8 bar)

Chassis

- Antistatic caster wheels, mobile trolley, powder-coated steel

- Central braking

Safety

- Overflow float valve (mechanical/electronic)

- Bacterial/hydrophobic filter in vacuum line

Power

- 100—240 V AC, 50/60 Hz

- Class I

Standards

- IEC 60601-1, -1-2, ISO 10079-1 (medical suction), ISO 13485, CE MDR / FDA

## 10.3 Standard Accessories

- Oil-free electric suction pump chassis with antistatic casters — See main specs

- Dual-scale analog/digital vacuum gauge — kPa/mmHg with obstetric zones

- Reusable autoclavable stainless steel Malmström vacuum cups (40, 50, 60 mm) — AISI 316L, autoclavable at 134 °C, complete with chain and traction handle

- Reusable flexible medical-grade silicone vacuum cups (50, 60 mm) — Autoclavable, atraumatic, colour-coded

- Reusable autoclavable polycarbonate collection jar (2 L) — Graduated, autoclavable, with sealing lid

- Overflow protection safety float valve — Mechanical/electronic, protects pump from fluid ingress

- Flexible heavy-duty silicone suction hose with traction handle — Autoclavable, minimum 1.5 m, kink-resistant

- Waterproof pneumatic foot switch — IPX8, ergonomic

## 10.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, hospital-grade socket

- Delivery room / OT space with proper earthing

## 10.5 Documentation

Full documentation package.

## 10.6 Training

Clinical + biomedical + PM.

## 10.7 Warranty

Minimum 2 years comprehensive.

# 11. VIDEO COLPOSCOPE WITH LED DISPLAY MONITOR

Quantity: 2

## 11.1 Clinical & Functional Specifications

The equipment shall be a digital video colposcope for magnified visual examination of the cervix, vagina, and vulva for detection of pre-invasive and invasive lesions. It shall provide high-resolution live imaging with optical zoom, green filter for vascular pattern enhancement, still and video capture, database management, and integration with LEEP/biopsy documentation.

## 11.2 Engineering Technical Specifications

Camera

- Digital CCD/CMOS sensor

- Resolution: Full HD 1920 × 1080 minimum

- Optical zoom: Minimum 10× (up to 30×), motorised or manual

- Digital zoom: Additional

- Auto/manual focus, auto white balance

Illumination

- High-intensity white LED, colour temperature \~5500—6500 K

- Coaxial illumination

- Integrated green filter (electronic or optical) for vascular enhancement

- LED life: Minimum 30,000 hours

- Adjustable intensity

Working Distance

- Approximately 250—350 mm

Monitor

- Medical-grade LED/LCD, minimum 21.5-inch diagonal

- Resolution: Full HD 1920 × 1080 minimum

- Mounted to swing-arm stand

Stand

- Floor-standing, heavy-duty swing arm, vertical and horizontal adjustment

- Lockable antistatic casters (minimum 4)

- Vertical height adjustment range: Minimum 900—1300 mm

Image Capture & Software

- Still image and video capture (MP4/AVI, JPEG/PNG)

- Foot-switch controlled freeze/capture

- Patient database with search, comparison, annotation, and report generation

- DICOM output support

- USB and LAN connectivity; HL7 support

Power

- 100—240 V AC, 50/60 Hz

- Integrated or dedicated UPS module (minimum 15 min backup for full system)

- Class I

Standards

- IEC 60601-1, -1-2, IEC 62304 (software), IEC 62366 (usability), DICOM, HL7, ISO 13485, CE MDR / FDA

## 11.3 Standard Accessories

- Digital CCD/CMOS video colposcope camera head with optical zoom — Full HD minimum, ≥10× optical zoom

- Coaxial green-filter LED illumination module — ≥30,000 h life, adjustable intensity

- Floor swing-arm stand with lockable rolling casters — Vertical/horizontal adjustment, heavy-duty

- 21.5-inch or larger medical-grade LED/LCD monitor stand-mounted — Full HD, medical-grade

- Handheld remote + integrated panel control switches — Ergonomic, membrane

- Waterproof foot switch for hands-free image capture — IPX8

- Image acquisition + patient database report software license — Full HD capture, DICOM, HL7, patient reporting, unrestricted concurrent-user license as per configuration

- UPS module for full-system electronics protection — Online/line-interactive, minimum 15 min runtime, sine-wave output

## 11.4 Installation Requirements

- 100—240 V AC, 50/60 Hz, hospital-grade earthed socket

- Network point for DICOM/HL7 integration

- Consultation/OPD room minimum 3 m × 3 m

- Ambient temperature 15—30 °C, humidity 30—75 %

## 11.5 Documentation

Full documentation package including software license certificate.

## 11.6 Training

Clinical + biomedical + software/database + PM + application.

## 11.7 Warranty

Minimum 2 years comprehensive including camera, LED module, monitor, and software updates.