TECHNICAL TENDER SPECIFICATIONS — OPHTHALMOLOGY DEPARTMENT

Technical Tender Document

Table of Contents

1. PHACOEMULSIFICATION UNIT	14

1.1 Clinical & Functional Specifications	14

1.2 Engineering Technical Specifications	14

1.3 Standard Accessories	15

1.4 Installation Requirements	15

1.5 Documentation	16

1.6 Training	16

1.7 Warranty	16

2. OPHTHALMIC OPERATING MICROSCOPE (WITH BIOM)	17

2.1 Clinical & Functional Specifications	17

2.2 Engineering Technical Specifications	17

2.3 Standard Accessories	17

2.4 Installation Requirements	18

2.5 Documentation	18

2.6 Training	18

2.7 Warranty	18

3. YAG LASER (POSTERIOR CAPSULOTOMY)	19

3.1 Clinical & Functional Specifications	19

3.2 Engineering Technical Specifications	19

3.3 Standard Accessories	19

3.4 Installation Requirements	20

3.5 Documentation	20

3.6 Training	20

3.7 Warranty	20

4. ARGON / MULTI-WAVELENGTH PHOTOCOAGULATION LASER	21

4.1 Clinical & Functional Specifications	21

4.2 Engineering Technical Specifications	21

4.3 Standard Accessories	21

4.4 Installation Requirements	22

4.5 Documentation	22

4.6 Training	22

4.7 Warranty	22

5. POSTERIOR VITRECTOMY UNIT	23

5.1 Clinical & Functional Specifications	23

5.2 Engineering Technical Specifications	23

5.3 Standard Accessories	23

5.4 Installation Requirements	24

5.5 Documentation	24

5.6 Training	24

5.7 Warranty	24

6. MICROPULSE TRANS-SCLERAL CYCLO-PHOTOCOAGULATION LASER	25

6.1 Clinical & Functional Specifications	25

6.2 Engineering Technical Specifications	25

6.3 Standard Accessories	25

6.4 Installation Requirements	25

6.5 Documentation	25

6.6 Training	26

6.7 Warranty	26

7. OPHTHALMIC CRYOTHERAPY UNIT	27

7.1 Clinical & Functional Specifications	27

7.2 Engineering Technical Specifications	27

7.3 Standard Accessories	27

7.4 Installation Requirements	27

7.5 Documentation	27

7.6 Training	28

7.7 Warranty	28

8. OPTICAL COHERENCE TOMOGRAPHY (OCT)	30

8.1 Clinical & Functional Specifications	30

8.2 Engineering Technical Specifications	30

8.3 Standard Accessories	30

8.4 Installation Requirements	31

8.5 Documentation	31

8.6 Training	31

8.7 Warranty	31

9. FUNDUS CAMERA	32

9.1 Clinical & Functional Specifications	32

9.2 Engineering Technical Specifications	32

9.3 Standard Accessories	32

9.4 Installation Requirements	33

9.5 Documentation	33

9.6 Training	33

9.7 Warranty	33

10. FUNDUS FLUORESCEIN ANGIOGRAPHY (FFA) UNIT	34

10.1 Clinical & Functional Specifications	34

10.2 Engineering Technical Specifications	34

10.3 Standard Accessories	34

10.4 Installation Requirements	35

10.5 Documentation	35

10.6 Training	35

10.7 Warranty	35

11. OPTICAL BIOMETER (IOL MASTER TYPE)	36

11.1 Clinical & Functional Specifications	36

11.2 Engineering Technical Specifications	36

11.3 Standard Accessories	36

11.4 Installation Requirements	37

11.5 Documentation	37

11.6 Training	37

11.7 Warranty	37

12. CORNEAL TOPOGRAPHY / TOMOGRAPHY UNIT	38

12.1 Clinical & Functional Specifications	38

12.2 Engineering Technical Specifications	38

12.3 Standard Accessories	38

12.4 Installation Requirements	39

12.5 Documentation	39

12.6 Training	39

12.7 Warranty	39

13. HUMPHREY VISUAL FIELD ANALYZER	40

13.1 Clinical & Functional Specifications	40

13.2 Engineering Technical Specifications	40

13.3 Standard Accessories	40

13.4 Installation Requirements	41

13.5 Documentation	41

13.6 Training	41

13.7 Warranty	41

14. AUTO-REFRACTOMETER / AUTO-KERATOMETER	42

14.1 Clinical & Functional Specifications	42

14.2 Engineering Technical Specifications	42

14.3 Standard Accessories	42

14.4 Installation Requirements	43

14.5 Documentation	43

14.6 Training	43

14.7 Warranty	43

15. NON-CONTACT (AIR-PUFF) TONOMETER	44

15.1 Clinical & Functional Specifications	44

15.2 Engineering Technical Specifications	44

15.3 Standard Accessories	44

15.4 Installation Requirements	44

15.5 Documentation	44

15.6 Training	45

15.7 Warranty	45

16. LENS METER — DIGITAL / AUTOMATIC	46

16.1 Clinical & Functional Specifications	46

16.2 Engineering Technical Specifications	46

16.3 Standard Accessories	46

16.4 Installation Requirements	46

16.5 Documentation	47

16.6 Training	47

16.7 Warranty	47

17. LENS METER — MANUAL	48

17.1 Clinical & Functional Specifications	48

17.2 Engineering Technical Specifications	48

17.3 Standard Accessories	48

17.4 Installation Requirements	48

17.5 Documentation	48

17.6 Training	49

17.7 Warranty	49

18. SLIT LAMP BIOMICROSCOPE WITH GOLDMANN TONOMETER	51

18.1 Clinical & Functional Specifications	51

18.2 Engineering Technical Specifications	51

18.3 Standard Accessories	51

18.4 Installation Requirements	52

18.5 Documentation	52

18.6 Training	52

18.7 Warranty	52

19. DIRECT OPHTHALMOSCOPE WITH RECHARGEABLE HANDLE	53

19.1 Clinical & Functional Specifications	53

19.2 Engineering Technical Specifications	53

19.3 Standard Accessories	53

19.4 Installation Requirements	53

19.5 Documentation	54

19.6 Training	54

19.7 Warranty	54

20. BINOCULAR INDIRECT OPHTHALMOSCOPE	55

20.1 Clinical & Functional Specifications	55

20.2 Engineering Technical Specifications	55

20.3 Standard Accessories	55

20.4 Installation Requirements	56

20.5 Documentation	56

20.6 Training	56

20.7 Warranty	56

21. STREAK RETINOSCOPE WITH RECHARGEABLE HANDLE	57

21.1 Clinical & Functional Specifications	57

21.2 Engineering Technical Specifications	57

21.3 Standard Accessories	57

21.4 Installation Requirements	57

21.5 Documentation	58

21.6 Training	58

21.7 Warranty	58

22. SCHIØTZ INDENTATION TONOMETER	59

22.1 Clinical & Functional Specifications	59

22.2 Engineering Technical Specifications	59

22.3 Standard Accessories	59

22.4 Installation Requirements	59

22.5 Documentation	59

22.6 Training	59

22.7 Warranty	60

23. HAND-HELD TONOMETER	61

23.1 Clinical & Functional Specifications	61

23.2 Engineering Technical Specifications	61

23.3 Standard Accessories	61

23.4 Installation Requirements	61

23.5 Documentation	62

23.6 Training	62

23.7 Warranty	62

24. HERTEL EXOPHTHALMOMETER	63

24.1 Clinical & Functional Specifications	63

24.2 Engineering Technical Specifications	63

24.3 Standard Accessories	63

24.4 Installation Requirements	63

24.5 Documentation	63

24.6 Training	63

24.7 Warranty	63

25. OPHTHALMIC EXAMINATION UNIT (REFRACTION CHAIR AND STAND)	65

25.1 Clinical & Functional Specifications	65

25.2 Engineering Technical Specifications	65

25.3 Standard Accessories	65

25.4 Installation Requirements	66

25.5 Documentation	66

25.6 Training	66

25.7 Warranty	66

26. VISUAL ACUITY PROJECTOR / DIGITAL ACUITY CHART	68

26.1 Clinical & Functional Specifications	68

26.2 Engineering Technical Specifications	68

26.3 Standard Accessories	68

26.4 Installation Requirements	69

26.5 Documentation	69

26.6 Training	69

26.7 Warranty	69

27. TRIAL LENS SET WITH TRIAL FRAME	70

27.1 Clinical & Functional Specifications	70

27.2 Engineering Technical Specifications	70

27.3 Standard Accessories	71

27.4 Installation Requirements	71

27.5 Documentation	71

27.6 Training	71

27.7 Warranty	71

28. SYNOPTOPHORE (MAJOR AMBLYOSCOPE)	72

28.1 Clinical & Functional Specifications	72

28.2 Engineering Technical Specifications	72

28.3 Standard Accessories	72

28.4 Installation Requirements	72

28.5 Documentation	73

28.6 Training	73

28.7 Warranty	73

29. HESS SCREEN	74

29.1 Clinical & Functional Specifications	74

29.2 Engineering Technical Specifications	74

29.3 Standard Accessories	74

29.4 Installation Requirements	74

29.5 Documentation	74

29.6 Training	75

29.7 Warranty	75

30. PRISM BAR SET (HORIZONTAL + VERTICAL)	76

30.1 Clinical & Functional Specifications	76

30.2 Engineering Technical Specifications	76

30.3 Standard Accessories	76

30.4 Installation Requirements	76

30.5 Documentation	76

30.6 Training	76

30.7 Warranty	76

31. TITMUS STEREOACUITY TEST	78

31.1 Clinical & Functional Specifications	78

31.2 Engineering Technical Specifications	78

31.3 Standard Accessories	78

31.4 Installation Requirements	78

31.5 Documentation	78

31.6 Training	78

31.7 Warranty	78

32. MADDOX WING	80

32.1 Clinical & Functional Specifications	80

32.2 Engineering Technical Specifications	80

32.3 Standard Accessories	80

32.4 Installation Requirements	80

32.5 Documentation	80

32.6 Training	80

32.7 Warranty	80

33. ISHIHARA TEST FOR COLOUR VISION	82

33.1 Clinical & Functional Specifications	82

33.2 Engineering Technical Specifications	82

33.3 Standard Accessories	82

33.4 Installation Requirements	82

33.5 Documentation	82

33.6 Training	82

33.7 Warranty	82

34. FRISBY STEREOTEST	84

34.1 Clinical & Functional Specifications	84

34.2 Engineering Technical Specifications	84

34.3 Standard Accessories	84

34.4 Installation Requirements	84

34.5 Documentation	84

34.6 Training	84

34.7 Warranty	84

35. RAF NEAR-POINT RULE	85

35.1 Clinical & Functional Specifications	85

35.2 Engineering Technical Specifications	85

35.3 Standard Accessories	85

35.4 Installation Requirements	85

35.5 Documentation	85

35.6 Training	85

35.7 Warranty	85

36. WORTH FOUR-DOT TEST WITH RED-GREEN GOGGLES	87

36.1 Clinical & Functional Specifications	87

36.2 Engineering Technical Specifications	87

36.3 Standard Accessories	87

36.4 Installation Requirements	87

36.5 Documentation	87

36.6 Training	87

36.7 Warranty	87

37. KAY PICTURE TEST	88

37.1 Clinical & Functional Specifications	88

37.2 Engineering Technical Specifications	88

37.3 Standard Accessories	88

37.4 Installation Requirements	88

37.5 Documentation	88

37.6 Training	88

37.7 Warranty	88

38. LogMAR NEAR VISUAL ACUITY CHART	90

38.1 Clinical & Functional Specifications	90

38.2 Engineering Technical Specifications	90

38.3 Standard Accessories	90

38.4 Installation Requirements	90

38.5 Documentation	90

38.6 Training	90

38.7 Warranty	90

39. LEA GRATING PADDLES	92

39.1 Clinical & Functional Specifications	92

39.2 Engineering Technical Specifications	92

39.3 Standard Accessories	92

39.4 Installation Requirements	92

39.5 Documentation	92

39.6 Training	92

39.7 Warranty	93

40. SPIELMANN TRANSLUCENT OCCLUDER	94

40.1 Clinical & Functional Specifications	94

40.2 Engineering Technical Specifications	94

40.3 Standard Accessories	94

40.4 Installation Requirements	94

40.5 Documentation	94

40.6 Training	94

40.7 Warranty	94

41. DIRECT GONIOSCOPY LENS (KOEPPE-TYPE)	96

41.1 Clinical & Functional Specifications	96

41.2 Engineering Technical Specifications	96

41.3 Standard Accessories	96

41.4 Installation Requirements	96

41.5 Documentation	96

41.6 Training	97

41.7 Warranty	97

42. GOLDMANN THREE-MIRROR LENS	98

42.1 Clinical & Functional Specifications	98

42.2 Engineering Technical Specifications	98

42.3 Standard Accessories	98

42.4 Installation Requirements	98

42.5 Documentation	99

42.6 Training	99

42.7 Warranty	99

43. ABRAHAM YAG CAPSULOTOMY CONTACT LENS	100

43.1 Clinical & Functional Specifications	100

43.2 Engineering Technical Specifications	100

43.3 Standard Accessories	100

43.4 Installation Requirements	100

43.5 Documentation	101

43.6 Training	101

43.7 Warranty	101

44. PERIPHERAL IRIDOTOMY (ABRAHAM) CONTACT LENS	102

44.1 Clinical & Functional Specifications	102

44.2 Engineering Technical Specifications	102

44.3 Standard Accessories	102

44.4 Installation Requirements	102

44.5 Documentation	103

44.6 Training	103

44.7 Warranty	103

45. PANFUNDOSCOPIC CONTACT LENS	104

45.1 Clinical & Functional Specifications	104

45.2 Engineering Technical Specifications	104

45.3 Standard Accessories	104

45.4 Installation Requirements	105

45.5 Documentation	105

45.6 Training	105

45.7 Warranty	105

46. MAINSTER RETINAL LASER LENS	106

46.1 Clinical & Functional Specifications	106

46.2 Engineering Technical Specifications	106

46.3 Standard Accessories	106

46.4 Installation Requirements	107

46.5 Documentation	107

46.6 Training	107

46.7 Warranty	107

MAJOR SURGICAL & LASER PLATFORMS

# 1. PHACOEMULSIFICATION UNIT

Quantity: 1

## 1.1 Clinical & Functional Specifications

The unit shall be a state-of-the-art microprocessor-controlled ophthalmic surgical system intended for cataract extraction by phacoemulsification, anterior vitrectomy, bipolar coagulation, and irrigation/aspiration in adult and paediatric patients. It shall support the complete surgical workflow including sculpting, quadrant/chop removal, epinuclear removal, cortical clean-up, capsular polishing, and anterior vitrectomy. The system shall provide selectable ultrasound modulation modes (continuous, pulsed, burst, and hyperpulse or equivalent modulated modes) to reduce cumulative dissipated energy and thermal load on ocular tissues. It shall provide fluidic control based on either peristaltic, Venturi, or hybrid pumps with real-time chamber stability management, occlusion detection, and post-occlusion surge suppression. Surgeon-programmable memory shall allow storage of multiple named surgeon profiles with independent step sequences.

## 1.2 Engineering Technical Specifications

- Display: Colour TFT/LED touchscreen, minimum 15-inch diagonal, resolution minimum 1024 × 768 pixels, tilt/swivel adjustable, viewable from surgeon and circulating-nurse positions.

- Control system: Microprocessor / embedded real-time OS with graphical user interface; multilingual capability where available.

- Pump technology: Peristaltic, Venturi, or hybrid; user-selectable or dual-pump configuration preferred.

- Vacuum range: Minimum 0 to 600 mmHg, adjustable in increments not exceeding 1 mmHg.

- Aspiration flow rate: Minimum 0 to 60 mL/min, adjustable.

- Bottle height / IOP control: Motorised, adjustable range minimum 0—110 cm above eye level, or equivalent active IOP compensation.

- Ultrasound frequency: Typical range 28—45 kHz.

- Ultrasound power: 0—100 %, adjustable in 1 % increments.

- US modulation modes: Continuous, pulsed (minimum up to 100 pps), burst, and modulated/hyperpulse modes.

- Anterior vitrectomy cutter rate: Minimum up to 800 cpm (guillotine type).

- Bipolar coagulation: Isolated output, adjustable 0—100 %, frequency in the range 500 kHz—2 MHz.

- Foot pedal: Dual-linear, wireless or wired, programmable, minimum 4 lateral switches, IPX8 rating.

- Memory: Minimum 20 surgeon profiles, each with multiple procedure steps.

- Data export: USB and/or network export of case logs.

- Power supply: 100—240 VAC, 50/60 Hz, single-phase.

- Battery/UPS backup: Internal or integrated UPS providing minimum 10 minutes safe shut-down / continuation of active case.

- Weight (console): Approximately 40—90 kg depending on configuration.

- Dimensions: Mobile console suitable for standard operating theatre; overall footprint compatible with routine OR layout.

- Environmental: Operating 10—40 °C, 30—75 % RH non-condensing.

- Compliance: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-2 (HF surgical), ISO 13485, ISO 14971, CE MDR and/or US FDA clearance.

## 1.3 Standard Accessories

a) Phacoemulsification handpieces (piezoelectric, 2 units) — Piezoelectric ultrasonic transducer; titanium tip compatible; frequency matched to console; autoclavable at 134 °C; supplied with irrigation sleeves; ergonomic balanced grip; cable with sealed silicone strain relief.

b) Anterior vitrectomy cutter handpiece — Guillotine-type, 20G or 23G compatible; cut rate compatible with console (minimum up to 800 cpm); autoclavable; pneumatic or electric drive as per console architecture.

c) Bipolar coagulation forceps with connecting cable — Reusable stainless-steel bipolar forceps; insulated shafts; autoclavable at 134 °C; supplied with silicone-jacketed cable minimum 3 m length; compatible with console coagulation output.

d) Irrigation/Aspiration handpieces (straight and angled coaxial) — Coaxial design; 0.3 mm aspiration port typical; autoclavable; supplied as matched pair (straight + curved/angled); polished tip.

e) Dual-linear multi-functional foot pedal control unit — Dual linear (pitch and yaw) programmable pedal; minimum 4 side switches; IPX8; wired or wireless with minimum 4 h battery life if wireless.

f) Height-adjustable motorised IV pole assembly — Integrated into console; motorised travel range minimum 0—110 cm above eye level; controllable from GUI and/or foot pedal.

g) Reusable test chambers for fluidic calibration — Set of transparent chambers compatible with tubing sets for pre-op priming and calibration verification.

h) Specialised torque wrench for tip tightening — Calibrated torque wrench matched to phaco tip specification; corrosion-resistant; autoclavable.

i) Integrated equipment trolley with anti-static locking casters — Anti-static, minimum 4 casters with at least 2 lockable; load capacity suitable for console + all accessories; smooth mobility in OR environment.

## 1.4 Installation Requirements

Single-phase 100—240 VAC, 50/60 Hz supply with dedicated isolated medical socket; earth leakage protected; UPS-backed circuit recommended. Ambient OR conditions 18—24 °C, 30—60 % RH. Level floor; minimum 1.0 m clearance around console for service access. Sterile-service compatibility for handpiece reprocessing (autoclave capability at 134 °C required in CSSD).

## 1.5 Documentation

User Manual (English), Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), Calibration Certificate, Preventive Maintenance Schedule, Installation & Commissioning Reports, Declaration of Conformity (CE/FDA), original equipment list with serial numbers.

## 1.6 Training

On-site clinical (surgeon + scrub nurse) training minimum 5 working days; biomedical engineering training including fluidics, handpiece service, calibration, and troubleshooting minimum 3 working days; refresher training within warranty period.

## 1.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, handpiece replacement/repair (excluding user damage), preventive maintenance (minimum twice yearly), breakdown maintenance, software updates, and technical support. Guaranteed after-sales spares availability minimum 10 years.

# 2. OPHTHALMIC OPERATING MICROSCOPE (WITH BIOM)

Quantity: 2

## 2.1 Clinical & Functional Specifications

Floor-mounted ophthalmic surgical microscope intended for anterior segment (cataract, cornea, glaucoma) and posterior segment (vitreoretinal) surgery. It shall provide coaxial red-reflex illumination, apochromatic optics, motorised zoom/focus/X-Y, and integration with a non-contact wide-angle viewing (BIOM-type) system for vitreoretinal work. The stand shall provide stable, vibration-free positioning with electromagnetic braking, and shall support surgeon and assistant simultaneous viewing.

## 2.2 Engineering Technical Specifications

- Optics: Apochromatic; parallel or convergent binocular; inclinable binocular tube, tilt range minimum 0°—180° or 0°—210°.

- Magnification: Motorised zoom, magnification factor range minimum 0.4× to 2.4× (giving typical total magnification 4×—25× with standard eyepiece/objective).

- Eyepieces: Wide-field, 10× or 12.5×, adjustable dioptre ±5 D minimum, adjustable interpupillary distance 55—75 mm.

- Objective lens: Motorised fine focus with working distance selectable/interchangeable in the range 175—225 mm.

- X-Y translation: Motorised, travel range minimum ±40 mm in both axes.

- Illumination: Coaxial red-reflex + oblique/spot illumination; LED or xenon; adjustable intensity 0—100 %; colour temperature typical 3200—5600 K; minimum 100,000 lux at working distance for LED. Blue-light and phototoxicity filters selectable.

- BIOM/non-contact wide-angle module: Integrated or attachable; field of view minimum 100°; motorised focus; image inverter (SDI or equivalent).

- Stand: Floor stand with counter-balanced arm; electromagnetic brakes on all joints; anti-static lockable casters; reach compatible with standard OR table.

- Foot control: Multi-functional, waterproof (IPX8), programmable, minimum 6 functions.

- Assistant scope: Co-observer stereo binocular attachment.

- Video/documentation port: Beam splitter port(s) with C-mount and/or integrated HD/4K camera capability.

- Power supply: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 15004, ISO 10936, CE MDR / FDA.

## 2.3 Standard Accessories

a) Non-contact wide-angle BIOM system — Motorised focus; field of view minimum 100°; integrated image inverter; sterilisable adjustment caps.

b) High-stability floor stand with electromagnetic brakes and caster wheels — Counter-balanced multi-joint arm; electromagnetic braking on all rotational axes; lockable anti-static casters (minimum 4).

c) Motorised zoom, focus, and X-Y translation mechanics — Fully motorised, foot-pedal controlled; X-Y ±40 mm minimum; focus travel minimum ±25 mm.

d) Multi-functional waterproof foot pedal — IPX8; minimum 6 programmable functions; ergonomic layout.

e) Red-reflex coaxial illumination LED/Xenon module — Coaxial red-reflex + oblique illumination; adjustable intensity; blue-light filter selectable; long-life LED (\>20,000 h) or xenon.

f) Sterilisable rubber caps for manual adjustment handles — Autoclavable at 134 °C; full set for all surgeon-manipulated handles.

g) Assistant co-observer view-tube attachment — Stereo binocular; matched magnification and focus; same optical quality as primary observer.

h) Heavy-duty industrial protective vinyl cover — Full stand and optics cover; flame-retardant; anti-static.

## 2.4 Installation Requirements

Single-phase 100—240 VAC dedicated circuit; earth-leakage protected; UPS recommended. Level OR floor with sufficient load capacity for base (approx. 150—250 kg). Minimum ceiling height 2.7 m. Vibration-isolated area preferred. Network drop for video documentation if integrated.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Optics Alignment Certificate, Electrical Safety Test Report (IEC 62353), Factory Test Certificate, Installation & Commissioning Reports, Preventive Maintenance Schedule.

## 2.6 Training

Clinical training for surgeons and OR staff (minimum 3 days). Biomedical engineering training: optical alignment verification, motor calibration, lamp/LED replacement, brake service (minimum 3 days).

## 2.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, LED/xenon lamps, preventive maintenance (twice yearly), software updates, and technical support. Spares availability minimum 10 years.

# 3. YAG LASER (POSTERIOR CAPSULOTOMY)

Quantity: 1

## 3.1 Clinical & Functional Specifications

Nd:YAG photodisruption laser system for posterior capsulotomy, peripheral iridotomy, and vitreolysis. Integrated slit-lamp delivery with coaxial dual-diode aiming beam, precise focus offset control, and configurable single/multi-burst firing modes.

## 3.2 Engineering Technical Specifications

- Laser medium: Nd:YAG, wavelength 1064 nm.

- Mode: Q-switched, single/multi-pulse, TEM₀₀ or equivalent.

- Pulse energy: Adjustable minimum 0.3 mJ to 10 mJ (per pulse).

- Pulse duration: Typical 3—5 ns.

- Repetition rate: Single, double, triple burst selectable; minimum burst up to 3 pulses.

- Aiming beam: Dual-diode red (630—670 nm), adjustable intensity, coaxial parallax-free.

- Focus offset: Adjustable posterior/anterior offset minimum ±500 µm in defined increments.

- Slit-lamp delivery: Integrated; magnification minimum 3 steps (typical 10×, 16×, 25×); binocular converging optics.

- Cone angle: Minimum 16°.

- Safety: Emergency stop, key-switch enable, safety interlock, laser-in-use indicator, external interlock connector.

- Display: Digital control panel with treatment counters, energy per shot, total energy.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-2-22 (laser safety), IEC 60825-1 (Class 4 laser), CE MDR / FDA.

## 3.3 Standard Accessories

a) Integrated high-precision slit lamp delivery chassis — Converging optics, adjustable magnification, chin/forehead rest.

b) Dual-diode aiming beam module — Coaxial red, adjustable intensity, parallax-free with treatment beam.

c) Control console with digital display — Programmable energy, offset, burst; treatment log.

d) Single/dual-pedal firing foot switch — Waterproof (IPX8); ergonomic; cable-connected.

e) Safety interlock connection cable / bypass plug — Compliant with IEC 60825-1 Class 4 requirements; door-interlock compatible.

f) High-visibility laser warning signage board — Bilingual where required; complies with IEC 60825 signage requirements.

## 3.4 Installation Requirements

Dedicated laser room with controlled access; door interlock recommended; laser-safe window covering or windowless room; warning signage at entry; laser-safety officer protocol. Single-phase 100—240 VAC dedicated circuit; UPS recommended.

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Laser Output Calibration Certificate, Electrical Safety Test Report, Laser Safety Compliance Certificate (IEC 60825), Factory Test Certificate, Installation & Commissioning Reports, PM Schedule.

## 3.6 Training

Clinical training for ophthalmologists (minimum 2 days) including laser safety. Biomedical engineering training on output verification, alignment, calibration (minimum 2 days).

## 3.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including flashlamp/diode module), calibration, preventive maintenance, software updates, technical support. Spares availability minimum 10 years.

# 4. ARGON / MULTI-WAVELENGTH PHOTOCOAGULATION LASER

Quantity: 1

## 4.1 Clinical & Functional Specifications

Solid-state or gas multi-wavelength photocoagulation laser for panretinal photocoagulation, focal/grid macular treatment, and treatment of retinal tears and vascular disease. It shall provide slit-lamp delivery with laser-safe operator filters and configurable pulse parameters, and optional continuous-wave and micropulse-equivalent modes where the platform supports it.

## 4.2 Engineering Technical Specifications

- Wavelength: Green 532 nm (mandatory); additional wavelengths (yellow 577 nm, red 670 nm) selectable/optional as per configuration.

- Mode: Continuous wave (CW) mandatory; micropulse or subthreshold modulated mode preferred.

- Power output: Adjustable minimum 50 mW to 2000 mW.

- Exposure duration: Adjustable range minimum 10 ms to 3000 ms.

- Interval time: Adjustable minimum 10 ms to 1000 ms.

- Spot size: Adjustable range minimum 50—1000 µm (at slit-lamp).

- Aiming beam: Red diode, adjustable intensity.

- Delivery: Slit-lamp adapter (mandatory); optional laser indirect ophthalmoscope or endolaser fibre connector.

- Safety: Operator-side automatic safety filter synchronised with treatment beam; emergency stop; key switch; door interlock connector.

- Display: Touchscreen or membrane control with treatment logs; remote/tablet control optional.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-2-22, IEC 60825-1 (Class 4), CE MDR / FDA.

## 4.3 Standard Accessories

a) Slit-lamp attachment laser delivery adapter — Compatible with the specified slit-lamp; integrated operator safety filter.

b) Laser-safe protective eye filters for operator microscope paths — Automatic in-line filters synchronised to firing.

c) Compact main laser generator console — Digital parameter display; programmable presets; treatment logging.

d) Waterproof heavy-duty firing foot switch pedal — IPX8; ergonomic.

e) Remote control touchscreen display pad bracket — Wireless or wired remote for parameter setting.

f) Safety interlock plug and laser room notification sign — Compliant with IEC 60825-1.

## 4.4 Installation Requirements

Dedicated controlled-access laser room; door interlock; warning signage; laser-safe eyewear for personnel. Single-phase 100—240 VAC dedicated circuit; UPS recommended.

## 4.5 Documentation

User Manual, Service Manual, Parts Catalogue, Laser Output Calibration Certificate, Electrical Safety Test Report, Laser Safety Compliance Certificate, Factory Test Certificate, Installation & Commissioning Reports, PM Schedule.

## 4.6 Training

Clinical training (minimum 3 days) including laser safety and treatment planning. Biomedical engineering training (minimum 2 days) on calibration, output verification, alignment.

## 4.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including laser diode/tube module), calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 5. POSTERIOR VITRECTOMY UNIT

Quantity: 1

## 5.1 Clinical & Functional Specifications

Integrated posterior segment surgical platform for pars plana vitrectomy, membrane peeling, endolaser, endodiathermy, and silicone oil / gas / fluid exchange. It shall provide high-speed vitrectomy cutting, precision fluidics with IOP control, high-intensity endo-illumination, integrated bipolar diathermy, and configurable foot-pedal control. Compatibility with 20G, 23G, 25G, and preferably 27G vitrectomy systems is required.

## 5.2 Engineering Technical Specifications

- Cutter type: Pneumatic guillotine, dual-drive (or equivalent high-speed technology).

- Cut rate: Minimum up to 10,000 cpm; higher preferred.

- Vacuum range: Minimum 0—650 mmHg, adjustable in 1 mmHg increments.

- Flow rate: Adjustable, compatible with 20G—27G cutters.

- IOP control: Active/pressurised infusion with adjustable IOP range minimum 0—120 mmHg.

- Endo-illumination: Xenon or LED, minimum 150 W xenon-equivalent output; minimum 2 illumination ports; colour filters (amber/green) selectable; light output adjustable 0—100 %.

- Bipolar diathermy: Isolated output, adjustable 0—100 %, frequency 500 kHz—2 MHz range.

- Silicone oil injection/extraction: Motorised; compatible with 1000 cSt and 5000 cSt silicone oil.

- Air/fluid exchange: Integrated pump; adjustable pressure 0—100 mmHg minimum.

- Foot pedal: Multi-linear, programmable, wireless or wired, IPX8.

- Display: Colour TFT/LED touchscreen, minimum 15-inch, resolution minimum 1024 × 768.

- Surgeon profiles: Minimum 20.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-2, ISO 13485, CE MDR / FDA.

## 5.3 Standard Accessories

a) High-speed pneumatic vitrectomy cutter handpieces (2 units) — Compatible with 23G/25G at minimum; matched to console cut rate; single-use or reprocessable per manufacturer specification.

b) Heavy-duty integrated gas-driven air compressor module — Oil-free medical-grade compressor; internal or dockable; delivers pneumatic pressure required for cutter operation.

c) Xenon/LED high-intensity endo-illumination console — Minimum 150 W xenon-equivalent; dual port; adjustable intensity; colour filters.

d) Multi-functional linear programmable foot pedal — Multi-linear (dual-axis minimum); IPX8; programmable.

e) Automated viscous silicone oil injection/extraction pump — Compatible with 1000 and 5000 cSt oils; motorised.

f) Integrated automated fluid infusion electronic IV pole — Motorised; travel range compatible with active IOP control; controllable from GUI/foot pedal.

g) Reusable bipolar diathermy forceps with cord — Autoclavable at 134 °C; insulated shafts; cable minimum 3 m.

h) Equipment rolling cart chassis with lockable wheels — Anti-static casters, minimum 2 locking; load-rated for full console assembly.

## 5.4 Installation Requirements

Single-phase 100—240 VAC dedicated circuit; UPS-backed circuit strongly recommended. OR ambient 18—24 °C, 30—60 % RH. Compressed medical air supply optional (system provides its own compressor). Minimum 1.0 m clearance around unit for service.

## 5.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Calibration Certificate, Installation & Commissioning Reports, PM Schedule, Declaration of Conformity.

## 5.6 Training

Clinical training for vitreoretinal surgeons and OR staff (minimum 5 days). Biomedical engineering training on fluidics, pneumatics, illumination module, and calibration (minimum 3 days).

## 5.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, xenon bulb/LED module replacement, PM (twice yearly), software updates, technical support. Spares availability minimum 10 years.

# 6. MICROPULSE TRANS-SCLERAL CYCLO-PHOTOCOAGULATION LASER

Quantity: 1

## 6.1 Clinical & Functional Specifications

Diode laser system delivering micropulse (subthreshold) trans-scleral cyclophotocoagulation for glaucoma management. It shall deliver energy in a train of short pulses with programmable duty cycle to reduce collateral thermal damage while achieving ciliary body modulation.

## 6.2 Engineering Technical Specifications

- Wavelength: 810 nm diode (nominal).

- Mode: Micropulse (mandatory); continuous-wave mode optional.

- Power output: Adjustable minimum 0—3000 mW.

- Micropulse duty cycle: Adjustable, typical range 10—50 %.

- Pulse duration (envelope): Adjustable minimum 100 ms to 10 s (total sweep time).

- Delivery: Fibre-optic to dedicated trans-scleral probe.

- Aiming beam: Diode, adjustable intensity.

- Safety: Emergency stop; key switch; door interlock; treatment log.

- Display: Touchscreen with parameter display.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-2-22, IEC 60825-1 (Class 4), CE MDR / FDA.

## 6.3 Standard Accessories

a) Main laser generator console with MicroPulse-mode software — Programmable duty cycle, power, envelope duration; treatment logging.

b) Sub-cyclophotocoagulation fibre delivery probe — Compatible with console; ergonomic handpiece; sterilisable or single-use per manufacturer specification.

c) Heavy-duty waterproof laser firing foot switch pedal — IPX8; ergonomic.

d) Patient safety interlock plug / bypass assembly — Complies with IEC 60825-1.

e) Laser room protective warning display card sign — Complies with IEC 60825 signage requirements.

## 6.4 Installation Requirements

Controlled-access laser room; door interlock; laser-safe eyewear; warning signage. Single-phase 100—240 VAC; UPS recommended.

## 6.5 Documentation

User Manual, Service Manual, Parts Catalogue, Laser Output Calibration Certificate, Electrical Safety Test Report, Laser Safety Compliance Certificate, Factory Test Certificate, Installation & Commissioning Reports, PM Schedule.

## 6.6 Training

Clinical (minimum 2 days) including laser safety and treatment protocols. Biomedical engineering (minimum 2 days) on calibration and probe interface.

## 6.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including diode module), calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 7. OPHTHALMIC CRYOTHERAPY UNIT

Quantity: 1

## 7.1 Clinical & Functional Specifications

Gas-driven (N₂O or CO₂) cryotherapy console for retinal cryopexy, cyclocryotherapy, and treatment of eyelid and conjunctival lesions. It shall provide rapid probe-tip freezing to therapeutic temperatures with automated defrost, non-electrical foot-pedal firing, and interchangeable probes for retinal and glaucoma applications.

## 7.2 Engineering Technical Specifications

- Refrigerant: N₂O or CO₂ (as per configuration).

- Probe tip temperature: Minimum —70 °C (N₂O) or —50 °C (CO₂) at tip.

- Freeze time to therapeutic temperature: Typical ≤ 5 seconds.

- Defrost: Automated; typical ≤ 5 seconds.

- Cylinder pressure input: Compatible with standard N₂O/CO₂ medical cylinders; regulator with pressure gauge.

- Console: Microprocessor-controlled; digital pressure and temperature display.

- Foot pedal: Non-electrical pneumatic (mandatory for cryo firing).

- Probes: Retinal probe and glaucoma probe (both interchangeable), autoclavable/sterilisable.

- Power: 100—240 VAC for console electronics.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR / FDA.

## 7.3 Standard Accessories

a) Microprocessor cryo-generator cabinet — Digital display; pressure regulator; probe port(s).

b) Reusable N₂O/CO₂ cryo-retina probe handle — Autoclavable; ergonomic; tip dimensions suitable for retinal cryopexy.

c) Reusable cryo-glaucoma probe handle — Autoclavable; tip dimensions suitable for cyclocryotherapy.

d) High-pressure gas cylinder connection hoses with pin adapters — Compliant with medical gas connection standards (pin-index or DIN); minimum 2 m length.

e) Non-electrical pneumatic firing foot pedal — Rugged; sealed for OR use.

f) Automated probe defrost cycle mechanism — Integrated in console.

## 7.4 Installation Requirements

Single-phase 100—240 VAC for console electronics; medical-grade N₂O or CO₂ cylinder with regulator; well-ventilated OR/procedure room; secure cylinder storage.

## 7.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Probe Temperature Verification Certificate, Installation & Commissioning Reports, PM Schedule.

## 7.6 Training

Clinical (minimum 2 days) on cryotherapy technique and safety. Biomedical engineering (minimum 1 day) on gas handling, probe service, leak testing.

## 7.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including probes for manufacturing defect), PM, technical support. Spares availability minimum 10 years.

DIAGNOSTIC IMAGING & REFRACTION INSTRUMENTS

# 8. OPTICAL COHERENCE TOMOGRAPHY (OCT)

Quantity: 1

## 8.1 Clinical & Functional Specifications

High-resolution spectral-domain or swept-source OCT system for non-invasive, non-contact cross-sectional imaging of the retina, optic nerve head, macula, and anterior segment. It shall provide quantitative retinal nerve fibre layer (RNFL), ganglion cell, and macular thickness analysis with normative database comparison, and shall support progression analysis for glaucoma and retinal disease. OCT-Angiography (OCT-A) capability is preferred where the platform supports it. Integrated eye-tracking shall compensate for micro-saccadic motion during acquisition.

## 8.2 Engineering Technical Specifications

- Technology: Spectral-Domain (SD-OCT) or Swept-Source (SS-OCT).

- Light source wavelength: SD-OCT typical 830—870 nm; SS-OCT typical 1050—1060 nm.

- Axial (depth) resolution: Minimum 5 µm in tissue.

- Transverse resolution: Typical 15—20 µm.

- Scan speed (A-scans/sec): Minimum 50,000 A-scans/second for SD; higher for SS.

- Scan depth: Minimum 2 mm in tissue.

- Scan width: Minimum 12 mm.

- Eye tracking: Integrated real-time active eye tracking with motion correction.

- Scan patterns: Line, cross, radial, raster/volume, circle (RNFL), macular cube, wide-field, anterior segment; all software-selectable.

- Analysis modules: RNFL, optic nerve head, ganglion cell / GCC, macular thickness, normative database, progression analysis, anterior segment; OCT-A preferred.

- Display: Dual medical-grade monitors, minimum 23-inch, resolution minimum 1920 × 1080.

- Workstation: Medical-grade PC, minimum 8-core processor, minimum 16 GB RAM, minimum 1 TB SSD storage, optical drive optional.

- Connectivity: DICOM 3.0, HL7, Ethernet (Gigabit), USB, PDF export.

- Chin/forehead rest: Motorised height adjustment.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 15004, DICOM 3.0, CE MDR / FDA.

## 8.3 Standard Accessories

a) Automated high-speed eye-tracking camera sensor module — Integrated active tracker; motion compensation during acquisition.

b) Motorised height-adjustable patient chin-rest mount — Motorised vertical travel; adjustable forehead strap.

c) High-performance processing workstation with dual monitors — Minimum specifications as above; medical-grade colour-calibrated monitors.

d) High-resolution colour fundus viewing alignment camera — Integrated fundus reference image for scan localisation.

e) Calibration test-eye standard validation kit — Model eye with known optical properties for QC.

f) Dedicated glaucoma and retina analysis software licence — Full-feature licence; normative database; progression module.

g) Uninterruptible power supply (UPS) module — Sized for full system; minimum 15 minutes runtime at full load.

## 8.4 Installation Requirements

Dedicated ophthalmic diagnostic room; ambient temperature 18—26 °C, 30—70 % RH non-condensing; controlled ambient lighting (low-light acquisition environment). Single-phase 100—240 VAC dedicated circuit; UPS-backed; earth-leakage protected. Network drop with DICOM connectivity to PACS. Minimum room size approximately 2.5 × 3 m. Vibration-free surface for the instrument table.

## 8.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report (IEC 62353), Software Licence Certificate, DICOM Conformance Statement, Installation & Commissioning Reports, PM Schedule.

## 8.6 Training

Clinical training for ophthalmologists and technicians (minimum 3 days) including acquisition, analysis, and reporting. Biomedical engineering training (minimum 2 days) on optical calibration verification, workstation service, and network integration.

## 8.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including light-source module), calibration, PM (twice yearly), software updates and upgrades within warranty, technical support. Spares availability minimum 10 years.

# 9. FUNDUS CAMERA

Quantity: 1

## 9.1 Clinical & Functional Specifications

Non-mydriatic or mydriatic digital fundus camera for high-resolution colour retinal photography, with capability for red-free and (where supported) auto-fluorescence imaging. It shall support routine screening (diabetic retinopathy, hypertensive retinopathy, glaucoma), documentation of retinal pathology, and integration with a PACS/DICOM environment. Non-mydriatic capability preferred to enable screening without pharmacological dilation.

## 9.2 Engineering Technical Specifications

- Imaging sensor: High-resolution CMOS or CCD, minimum 12 megapixels effective.

- Field of view: Minimum 40°, with wide-field/mosaic option preferred.

- Illumination: LED or xenon flash; flash energy adjustable; infrared illumination for alignment in non-mydriatic mode.

- Working distance: Typical 35—45 mm.

- Minimum pupil diameter (non-mydriatic mode): Approximately 3.3 mm or better.

- Focus: Motorised or manual with dioptre compensation range minimum —20 D to +20 D.

- Imaging modes: Colour, red-free (green), infrared; fundus auto-fluorescence preferred.

- Display: Colour LCD/LED, minimum 19-inch medical monitor.

- Workstation: Medical-grade PC, minimum 8 GB RAM, minimum 500 GB storage.

- Connectivity: DICOM 3.0, HL7, Ethernet, USB, PDF export.

- Chin/forehead rest: Motorised height adjustment; adjustable forehead strap.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 15004, DICOM 3.0, CE MDR / FDA.

## 9.3 Standard Accessories

a) High-resolution digital SLR/CMOS camera back attachment — Minimum 12 MP; matched to optical head; interchangeable per manufacturer configuration.

b) Motorised 3D chin-rest tracking assembly — Motorised X-Y-Z alignment; adjustable chin/forehead support.

c) High-performance imaging computer tower with medical display monitor — Colour-calibrated monitor; medical-grade PC.

d) Dedicated retinal imaging analysis software licence — Full-feature software with database, reporting, DICOM export.

e) Main system isolation transformer and UPS module — Medical isolation transformer; UPS minimum 15 minutes runtime.

f) Translucent protective dust cover sheet — Full-system cover; anti-static.

## 9.4 Installation Requirements

Ophthalmic diagnostic room with dim ambient lighting; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC dedicated circuit; UPS-backed; earth-leakage protected. Network drop with DICOM/PACS connectivity. Vibration-free surface for instrument table.

## 9.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report, DICOM Conformance Statement, Software Licence Certificate, Installation & Commissioning Reports, PM Schedule.

## 9.6 Training

Clinical training (minimum 3 days) on acquisition, analysis, and reporting. Biomedical engineering training (minimum 2 days) on optical alignment, flash/LED service, workstation and network integration.

## 9.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including flash/LED module), calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 10. FUNDUS FLUORESCEIN ANGIOGRAPHY (FFA) UNIT

Quantity: 1

## 10.1 Clinical & Functional Specifications

Digital fundus camera platform equipped with excitation and barrier filters for intravenous fluorescein angiography (and indocyanine green angiography where supported), providing high-speed time-sequenced imaging of retinal and choroidal vasculature. The platform shall provide synchronisation with injection timing, high frame-rate acquisition during arteriovenous phases, and a dedicated angiography analysis and reporting workstation.

## 10.2 Engineering Technical Specifications

- Imaging sensor: High-sensitivity CMOS/CCD, minimum 12 MP effective.

- Field of view: Minimum 40°; wide-field option preferred.

- Excitation filter (FA): Blue, central wavelength approximately 465—490 nm.

- Barrier filter (FA): Yellow-green, central wavelength approximately 520—540 nm.

- ICG filters (if supported): Excitation approximately 795 nm; barrier approximately 835 nm.

- Frame rate (angiography): Minimum 5 frames/second for arteriovenous phase; higher preferred.

- Illumination: Xenon flash or high-intensity LED with matched filters.

- Focus/dioptre: —20 D to +20 D minimum.

- Display: Dual medical-grade monitors, minimum 19-inch each.

- Workstation: Medical-grade PC, minimum 8-core processor, minimum 16 GB RAM, minimum 1 TB SSD.

- Timer/injector sync: Integrated angiography timer with external injector-sync port.

- Connectivity: DICOM 3.0, HL7, Ethernet, USB.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 15004, DICOM 3.0, CE MDR / FDA.

## 10.3 Standard Accessories

a) High-speed digital capture camera head with barrier filters — Matched excitation/barrier filter set; high sensitivity for angiography.

b) Multi-mode digital exciter filter wheel assembly — Motorised or manual filter wheel; FA and ICG (where supported) positions.

c) Dual-monitor diagnostic reporting workstation with processing CPU — Colour-calibrated medical monitors; medical-grade PC.

d) Injector sync timer connection cable port — External sync port for angiographic injector; integrated timer.

e) Dedicated vascular angiography mapping database software licence — Full-feature software including time-sequenced viewing, mosaic, and comparison.

f) Uninterruptible power supply (UPS) module — Sized for full system; minimum 15 minutes runtime.

## 10.4 Installation Requirements

Dedicated diagnostic room with dim ambient lighting; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC dedicated circuit; UPS-backed. Network drop with DICOM/PACS connectivity. Emergency medical support and resuscitation kit within immediate proximity (fluorescein/ICG adverse reactions).

## 10.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Filter Spectral Certificate, Electrical Safety Test Report, DICOM Conformance Statement, Software Licence Certificate, Installation & Commissioning Reports, PM Schedule.

## 10.6 Training

Clinical training (minimum 3 days) on FA/ICG acquisition, phase interpretation, safety, and adverse-event management. Biomedical engineering training (minimum 2 days) on filter alignment, flash service, workstation integration.

## 10.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including flash module and filters), calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 11. OPTICAL BIOMETER (IOL MASTER TYPE)

Quantity: 1

## 11.1 Clinical & Functional Specifications

Non-contact optical biometer for pre-cataract-surgery ocular measurements including axial length, corneal curvature (keratometry), anterior chamber depth, lens thickness, and white-to-white distance. It shall include a comprehensive suite of modern IOL calculation formulae (including at minimum SRK/T, Hoffer Q, Holladay, Haigis, and Barrett Universal II or equivalent) and shall support toric IOL calculation and post-refractive-surgery cases.

## 11.2 Engineering Technical Specifications

- Technology: Optical low-coherence interferometry / partial coherence interferometry / swept-source optical biometry.

- Axial length measurement range: Minimum 14 mm to 38 mm.

- Axial length accuracy: Within ±30 µm typical.

- Keratometry range: Minimum 5 mm to 10 mm radius (equivalent to approximately 33.75—67.5 D).

- Keratometry accuracy: Within ±0.25 D.

- Anterior chamber depth range: Minimum 1.5 mm to 6.5 mm.

- Lens thickness measurement: Supported (SS-OCT-based biometers).

- White-to-white: Supported; accuracy within ±0.25 mm typical.

- IOL formulae: Minimum SRK/T, Hoffer Q, Holladay I/II, Haigis, Barrett Universal II (or equivalent modern formulae); toric calculator; post-refractive module.

- Alignment: Motorised joystick-controlled 3D alignment.

- Fixation targets: Internal LED fixation array.

- Display: Medical-grade monitor minimum 19-inch.

- Workstation: Integrated or dedicated PC, minimum 8 GB RAM.

- Connectivity: DICOM 3.0, Ethernet, USB.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 15004, DICOM 3.0, CE MDR / FDA.

## 11.3 Standard Accessories

a) Motorised joystick-controlled 3D alignment cradle — X-Y-Z motorised; joystick control.

b) Internal fixation LED target illumination arrays — Multi-position LED fixation for eccentric measurements.

c) High-performance processing workstation tower and medical monitor — Medical-grade PC and colour-calibrated monitor.

d) Reusable calibration test-eye reference cylinder — Model eye with certified axial length and K values for QC.

e) Integrated advanced IOL calculation formula software package — Full-feature licence including modern formulae, toric, post-refractive, and reporting.

f) Laser report printer with paper tray — Networked or USB-connected office-grade laser printer; A4.

## 11.4 Installation Requirements

Ophthalmic diagnostic room; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC dedicated circuit; UPS recommended. Network drop for DICOM/EMR integration. Vibration-free instrument table.

## 11.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report, DICOM Conformance Statement, Software Licence Certificate, Installation & Commissioning Reports, PM Schedule.

## 11.6 Training

Clinical training (minimum 2 days) on measurement technique, IOL formula selection, toric and post-refractive cases. Biomedical engineering training (minimum 1 day) on calibration verification and workstation service.

## 11.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, calibration, PM, software updates (including new IOL formulae released within warranty period), technical support. Spares availability minimum 10 years.

# 12. CORNEAL TOPOGRAPHY / TOMOGRAPHY UNIT

Quantity: 1

## 12.1 Clinical & Functional Specifications

Non-contact corneal topographer/tomographer for anterior corneal surface mapping and (where supported) posterior corneal surface and full corneal pachymetry. It shall provide Placido-disc-based topography and, where the platform is a tomographer, Scheimpflug or OCT-based tomography with pachymetry and elevation maps. Clinical modules shall include keratoconus screening, contact lens fitting, refractive surgery planning, and toric IOL planning.

## 12.2 Engineering Technical Specifications

- Technology: Placido-disc topography (mandatory); Scheimpflug or OCT tomography preferred for full corneal mapping.

- Number of Placido rings: Minimum 20 rings (typical 22—32).

- Measurement points: Minimum 6,000 measurement points across the cornea.

- Coverage: Minimum central 8 mm of cornea; wide-field coverage preferred.

- Curvature measurement range: Minimum 5—10 mm radius (approximately 33.75—67.5 D).

- Accuracy (curvature): Within ±0.25 D typical.

- Pachymetry (tomographers): Range minimum 300—900 µm; accuracy within ±5 µm typical.

- Analysis modules: Axial, tangential, elevation (anterior and, where supported, posterior), pachymetry map, keratoconus screening indices, contact-lens fitting, refractive planning, toric IOL planning.

- Chin-rest: Motorised height adjustment.

- Display: Colour medical-grade monitor minimum 19-inch.

- Workstation: Dedicated PC, minimum 8 GB RAM, minimum 500 GB storage.

- Connectivity: DICOM 3.0, Ethernet, USB.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 15004, ISO 19980 (topographers), DICOM 3.0, CE MDR / FDA.

## 12.3 Standard Accessories

a) High-density multi-ring Placido disc projector cone — Minimum 20 rings; illuminated LED.

b) Motorised adjustment patient chin rest assembly — Motorised vertical adjustment; adjustable forehead strap.

c) Dedicated processing computer tower and medical display — Medical-grade PC and colour-calibrated monitor.

d) Reusable calibration verification sphere block tool — Certified reference sphere(s) for QC.

e) Advanced keratoconus screening analysis software licence — Full-feature software including keratoconus indices, contact-lens fitting, refractive planning.

f) Heavy-duty protective dust sheet cover — Full-instrument cover; anti-static.

## 12.4 Installation Requirements

Ophthalmic diagnostic room; ambient 18—26 °C, 30—70 % RH; controlled lighting. Single-phase 100—240 VAC dedicated circuit; UPS recommended. Network drop for DICOM/EMR integration. Vibration-free instrument table.

## 12.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report, DICOM Conformance Statement, Software Licence Certificate, Installation & Commissioning Reports, PM Schedule.

## 12.6 Training

Clinical training (minimum 2 days) on acquisition, map interpretation, keratoconus screening, and refractive/toric planning. Biomedical engineering training (minimum 1 day) on calibration and workstation service.

## 12.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 13. HUMPHREY VISUAL FIELD ANALYZER

Quantity: 1

## 13.1 Clinical & Functional Specifications

Automated static bowl perimeter for quantitative measurement of the visual field for glaucoma diagnosis and follow-up, and for neurological visual field defects. It shall provide standard automated perimetry (SAP), short-wavelength automated perimetry (SWAP) or equivalent where supported, and shall include normative databases, glaucoma progression analysis (event-based and trend-based), and reliability indices (fixation losses, false positives, false negatives).

## 13.2 Engineering Technical Specifications

- Bowl type: Hemispheric white bowl.

- Bowl radius: Approximately 300 mm (standard).

- Background illumination: 31.5 apostilbs (asb) or equivalent (10 cd/m²), standardised.

- Stimulus size: Goldmann sizes I to V, selectable.

- Stimulus intensity range: Minimum 0.08 to 10,000 asb (approximately 0—51 dB).

- Stimulus duration: Adjustable, typical 100—200 ms.

- Stimulus colours: White (mandatory); blue-on-yellow (SWAP) preferred.

- Test strategies: Full threshold, SITA-standard equivalent (fast threshold), SITA-fast equivalent, suprathreshold screening; kinetic (where supported).

- Test patterns: Central 10-2, 24-2, 30-2, macular, peripheral, full-field 60/120; neurological patterns.

- Fixation monitoring: Gaze tracking (mandatory), Heijl-Krakau blind-spot method, video monitoring.

- Analysis: Reliability indices, total deviation, pattern deviation, glaucoma hemifield test, mean deviation (MD), pattern standard deviation (PSD), visual field index (VFI), progression analysis (event- and trend-based).

- Display: Colour LCD/LED touchscreen minimum 15-inch.

- Workstation: Integrated PC; minimum 8 GB RAM; minimum 500 GB storage.

- Connectivity: DICOM 3.0 or vendor equivalent, Ethernet, USB.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 12866, CE MDR / FDA.

## 13.3 Standard Accessories

a) Integrated motorised chin rest and forehead support assembly — Motorised vertical adjustment; adjustable forehead strap.

b) Patient response button switch with coiled connecting cable — Ergonomic press-button; coiled cable minimum 1.5 m extended.

c) Integrated optical gaze tracking system module — Real-time gaze/pupil tracking; reliability recording.

d) External thermal graphics printer with paper roll — Compatible with system; supplied with initial paper stock.

e) Ergonomic alphanumeric keyboard and optical mouse — Standard USB peripherals.

f) Red/Blue/White stimulus filters for specialised testing — For SWAP or specialised protocols where platform supports it.

g) Goldmann standard calibration verification tool kit — For periodic calibration verification.

h) Main uninterruptible power supply (UPS) module — Minimum 15 minutes runtime.

## 13.4 Installation Requirements

Dedicated quiet, low-ambient-light diagnostic room; wall colour matte non-reflective preferred; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC dedicated circuit; UPS-backed. Network drop for EMR/DICOM integration. Comfortable patient chair.

## 13.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Photometric Calibration Certificate, Electrical Safety Test Report, Software Licence Certificate, Installation & Commissioning Reports, PM Schedule.

## 13.6 Training

Clinical training (minimum 2 days) on test selection, interpretation, reliability assessment, and progression analysis. Biomedical engineering training (minimum 1 day) on photometric verification, gaze tracker service, workstation integration.

## 13.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including stimulus bulb/LED), photometric calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 14. AUTO-REFRACTOMETER / AUTO-KERATOMETER

Quantity: 2

## 14.1 Clinical & Functional Specifications

Combined objective auto-refractometer and auto-keratometer for automated measurement of refractive error (sphere, cylinder, axis) and corneal curvature/keratometry. It shall provide fast, non-invasive, non-contact measurements suitable for adult and paediatric patients, with fixation targets appropriate for children.

## 14.2 Engineering Technical Specifications

- Sphere measurement range: Minimum —25.00 D to +22.00 D.

- Cylinder measurement range: Minimum —10.00 D to +10.00 D.

- Axis measurement range: 0° to 180°, in 1° increments.

- Sphere/cylinder step: Selectable, minimum 0.25 D.

- Keratometry radius range: Minimum 5 mm to 10 mm.

- Keratometry power range: Approximately 33.75 D to 67.5 D.

- Keratometry accuracy: Within ±0.25 D.

- Pupil diameter measurement: Minimum range 2.0 to 8.5 mm.

- Minimum pupil for measurement: Approximately 2.0 mm.

- Vertex distance: Selectable, typical 0, 12, 13.5, 15 mm.

- Alignment: Motorised 3D auto-alignment and auto-tracking with auto-shot.

- Fixation target: Internal graphic fixation with fogging; child-friendly target selectable.

- Display: Colour LCD touchscreen minimum 7-inch.

- Printer: Integrated thermal strip printer.

- Connectivity: RS-232, USB, Ethernet.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 10342 (refractometers), ISO 10343 (keratometers), CE MDR / FDA.

## 14.3 Standard Accessories

a) Motorised 3D alignment and tracking sensor system — X-Y-Z motorised; auto-alignment; auto-shot.

b) High-resolution colour LCD touchscreen user interface panel — Minimum 7-inch; capacitive or resistive touch.

c) Integrated thermal strip printer with 2 rolls of paper — Standard thermal roll; supplied with initial stock.

d) Standard calibration test eye sphere tool — Certified model eye for QC.

e) Reusable adjustable chin-rest pins and headrest band — Standard sanitary pins with disposable sleeves; adjustable strap.

f) Protective translucent dust cover jacket — Full-instrument cover; anti-static.

## 14.4 Installation Requirements

Refraction lane or dedicated diagnostic room; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC standard socket; UPS recommended. Sturdy motorised or fixed instrument table.

## 14.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report, Installation & Commissioning Reports, PM Schedule.

## 14.6 Training

Clinical training (minimum 1 day) on acquisition and paediatric technique. Biomedical engineering training (minimum 1 day) on calibration verification and printer service.

## 14.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 15. NON-CONTACT (AIR-PUFF) TONOMETER

Quantity: 1

## 15.1 Clinical & Functional Specifications

Non-contact air-puff tonometer for non-invasive measurement of intraocular pressure (IOP) without corneal contact or topical anaesthesia. It shall include auto-alignment, auto-shot, and averaging of multiple measurements with reliability indication. Corneal-compensated IOP output preferred where supported.

## 15.2 Engineering Technical Specifications

- Measurement range: Minimum 0 to 60 mmHg.

- Measurement step: Selectable, minimum 1 mmHg.

- Accuracy: Within ±2 mmHg (0—30 mmHg range) typical.

- Working distance: Approximately 11 mm.

- Air pulse: Automatic soft-puff with adjustment for measured IOP (soft-puff range control).

- Alignment: Motorised 3D auto-alignment; auto-shot.

- Averaging: Automatic multi-measurement averaging with reliability indication.

- Display: Colour LCD touchscreen minimum 5-inch.

- Printer: Integrated thermal strip printer.

- Connectivity: RS-232, USB, Ethernet.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 8612 (tonometers), CE MDR / FDA.

## 15.3 Standard Accessories

a) Auto-alignment 3D puff-positioning camera sensor system — Motorised 3D alignment with auto-shot.

b) Integrated thermal strip printer with 2 rolls of paper — Standard thermal roll; supplied with initial stock.

c) Standard test-eye verification tool cylinder — Certified reference for QC.

d) Reusable adjustable chin-rest chin pins set — Standard pins with disposable sleeves.

e) Translucent protective dust cover shroud — Full-instrument cover; anti-static.

## 15.4 Installation Requirements

Ophthalmic diagnostic/refraction lane; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC standard socket; UPS recommended. Instrument table (motorised preferred).

## 15.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, IOP Calibration Certificate, Electrical Safety Test Report, Installation & Commissioning Reports, PM Schedule.

## 15.6 Training

Clinical training (minimum ½ day) on technique. Biomedical engineering training (minimum ½ day) on calibration and service.

## 15.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 16. LENS METER — DIGITAL / AUTOMATIC

Quantity: 2

## 16.1 Clinical & Functional Specifications

Automatic digital lensmeter for measurement of single-vision, bifocal, trifocal, and progressive spectacle lenses and contact lenses, providing sphere, cylinder, axis, prism, and add power. Automatic progressive lens identification and UV/blue-light transmittance measurement are preferred features.

## 16.2 Engineering Technical Specifications

- Sphere range: Minimum —25.00 D to +25.00 D.

- Cylinder range: Minimum —10.00 D to +10.00 D.

- Axis: 0° to 180°, 1° increments.

- Add power range: Minimum 0 to +10.00 D.

- Prism range: Minimum 0 to 10 Δ.

- Step: Selectable, minimum 0.01 D or 0.25 D.

- Progressive lens identification: Automatic mapping (preferred).

- UV transmission measurement: Supported (preferred).

- Contact lens measurement: Supported with appropriate holder.

- Display: Colour LCD tilt-screen, minimum 5.7-inch.

- Printer: Integrated thermal strip printer.

- Connectivity: RS-232, USB, Ethernet.

- Power: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 8598 (lensmeters), CE MDR / FDA.

## 16.3 Standard Accessories

a) Integrated progressive lens identification software licence — Full-feature software for auto-detection and mapping of progressive lenses.

b) High-resolution colour LCD tilt-screen display interface — Minimum 5.7-inch tiltable display.

c) Built-in UV transmission and blue-light measurement module — Integrated measurement of UV/blue-light transmittance.

d) Integrated thermal strip printer with paper roll — Standard thermal roll; supplied with initial stock.

e) Removable lens marking pen ink cartridge kit — Compatible marker cartridges; spare cartridges supplied.

f) Standard calibration verification block plate — Certified reference for QC.

## 16.4 Installation Requirements

Optical dispensary or diagnostic room; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC standard socket. Stable work surface.

## 16.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Electrical Safety Test Report, Installation & Commissioning Reports, PM Schedule.

## 16.6 Training

Clinical/dispensing training (minimum ½ day). Biomedical engineering training (minimum ½ day) on calibration and service.

## 16.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts, calibration, PM, software updates, technical support. Spares availability minimum 10 years.

# 17. LENS METER — MANUAL

Quantity: 1

## 17.1 Clinical & Functional Specifications

Manual (analogue) lensmeter for measurement of sphere, cylinder, axis, prism, and add power for single-vision, bifocal, and trifocal spectacle lenses. Serves as backup instrument and for verification/teaching.

## 17.2 Engineering Technical Specifications

- Sphere range: Minimum —25.00 D to +25.00 D.

- Cylinder range: Minimum ±6.00 D or greater.

- Axis: 0° to 180°, 1° increments.

- Prism range: Minimum 0 to 5 Δ (integrated or auxiliary prism compensator).

- Reading resolution: 0.25 D or better.

- Target: Multi-dot corona projected target.

- Illumination: Internal incandescent or LED lamp.

- Eyepiece: Focusable, dioptre adjustment ±5 D minimum.

- Lens holder: Adjustable platform with non-marring clamp.

- Power: 100—240 VAC for illumination; where LED-only, low-voltage adapter supplied.

- Compliance: IEC 60601-1 (or IEC 61010 as applicable), ISO 8598, CE.

## 17.3 Standard Accessories

a) Internal scale display with focusable eyepiece ring — Adjustable dioptre eyepiece with clear scale visibility.

b) Multi-dot corona target optical projection alignment — Standard corona target for reading focus and axis.

c) Adjustable lens holder platform with non-marring clamp feet — Non-marring elastomer feet; adjustable clamp.

d) Integrated marking pin ink pad well assembly — Three-pin marker with refillable ink well; supplied with initial ink.

e) Replacement illumination lamp bulb — Minimum 1 spare bulb/LED matched to instrument.

## 17.4 Installation Requirements

Optical dispensary or diagnostic room; ambient 18—26 °C, 30—70 % RH. Standard mains socket where powered; stable work surface.

## 17.5 Documentation

User Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report (where powered), Installation Report.

## 17.6 Training

Dispensing/refraction training (minimum ½ day).

## 17.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including lamp), technical support. Spares availability minimum 10 years.

SLIT LAMP, TONOMETERS, OPHTHALMOSCOPES, RETINOSCOPE, EXOPHTHALMOMETER, EXAMINATION UNIT

# 18. SLIT LAMP BIOMICROSCOPE WITH GOLDMANN TONOMETER

Quantity: 3

## 18.1 Clinical & Functional Specifications

Binocular stereoscopic slit-lamp biomicroscope for detailed examination of the anterior segment (lids, conjunctiva, cornea, anterior chamber, iris, lens) and, with auxiliary lenses, of the vitreous and posterior segment. It shall include an integrated Goldmann applanation tonometer for gold-standard intraocular pressure measurement. Illumination and magnification shall be independently adjustable, with cobalt-blue and red-free (green) filters for fluorescein staining and vascular examination.

## 18.2 Engineering Technical Specifications

- Microscope type: Galilean or Greenough binocular; converging or parallel optics.

- Magnification: Minimum 5 steps (typical 6×, 10×, 16×, 25×, 40×) via drum or motorised zoom.

- Total magnification range: Minimum 6× to 40×.

- Eyepieces: Wide-field, 10× or 12.5×; dioptre adjustment ±5 D minimum; adjustable interpupillary distance minimum 55—75 mm.

- Slit width: Continuously adjustable, minimum 0 to 14 mm.

- Slit length: Continuously adjustable, minimum 1 to 14 mm.

- Slit rotation: 0°—180°, continuous.

- Slit inclination: Minimum 0°—20° or greater.

- Filters: Cobalt blue, red-free (green), heat-absorbing, neutral density, yellow (barrier for fluorescein observation).

- Illumination: LED (preferred) or halogen; adjustable intensity 0—100 %; long-life LED ≥ 20,000 h.

- Base movement (X-Y-Z): Joystick control with fine focus; travel minimum ±40 mm X-Y, ±30 mm Z.

- Chin/forehead rest: Adjustable; disposable chin-rest paper compatible.

- Tonometer mounting: R-type or T-type mount for Goldmann applanation tonometer.

- Power supply: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 10939 (slit lamps), ISO 8612 (tonometers), CE MDR / FDA.

## 18.3 Standard Accessories

a) Goldmann applanation tonometer (R-type or T-type mount) — Prism-based applanation tonometer; graduated in mmHg; compatible with slit lamp mount; measurement range minimum 0—60 mmHg; accuracy within ±1 mmHg after calibration.

b) Tonometer calibration verification check bar — Certified check-weight bar for periodic calibration verification.

c) Reusable applanation tonometer cones (set of 2) — Optical-grade prism cones; autoclavable or chemically disinfectable per manufacturer's protocol; supplied as set of 2.

d) Integrated high-intensity LED illumination module — Long-life LED (minimum 20,000 h); adjustable intensity 0—100 %; colour temperature typical 4000—5600 K.

e) Adjustable chin-rest assembly with paper sheets pack — Adjustable vertical travel; supplied with initial stock of disposable chin-rest paper (minimum 1 pack per unit).

f) Yellow contrast filter and cobalt blue illumination filter — Standard yellow barrier filter for enhanced fluorescein visualisation; cobalt blue for excitation.

g) Mechanical focusing test rod bar — Metal test rod for eyepiece dioptre focusing and reticle alignment.

h) Protective breathable dust cover sheet — Full-instrument cover; anti-static; breathable fabric.

## 18.4 Installation Requirements

Ophthalmic examination lane; controllable ambient lighting (dim environment for slit-lamp examination); ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC standard socket; earth-leakage protected. Motorised or fixed instrument table (compatible with slit lamp base). Comfortable patient chair.

## 18.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Optical Calibration Certificate, Tonometer Calibration Certificate, Electrical Safety Test Report (IEC 62353), Installation & Commissioning Reports, PM Schedule.

## 18.6 Training

Clinical training for ophthalmologists (minimum 1 day) on slit-lamp technique and tonometry. Biomedical engineering training (minimum 1 day) on optical alignment, illumination service, tonometer calibration.

## 18.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including LED module and tonometer prism), calibration, PM, technical support. Spares availability minimum 10 years.

# 19. DIRECT OPHTHALMOSCOPE WITH RECHARGEABLE HANDLE

Quantity: 5

## 19.1 Clinical & Functional Specifications

Handheld direct ophthalmoscope for monocular examination of the ocular fundus (optic disc, macula, retinal vasculature) and anterior segment, with rechargeable handle. Multiple aperture and filter selections shall be provided for varied clinical assessments (small pupil, fixation, red-free, cobalt blue).

## 19.2 Engineering Technical Specifications

- Illumination: LED (preferred) or xenon/halogen; long-life LED ≥ 20,000 h; colour temperature typical 3200—5600 K; adjustable intensity via rheostat.

- Apertures: Minimum 5 selectable — small spot, large spot, fixation target/graticule, slit, red-free.

- Filters: Cobalt blue (for fluorescein staining assessment), red-free (green).

- Dioptre lens wheel: Range minimum —25 D to +25 D (positive and negative), single-dioptre increments in central range.

- Handle: Rechargeable lithium-ion; capacity providing minimum 2 hours continuous use per full charge.

- Charging time: Maximum 4 hours full charge.

- Charging cradle: Desk-top; simultaneous charging of 2 handles (for twin cradle).

- Weight (optical head): Approximately 100—250 g.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 10943 (ophthalmoscopes), CE MDR / FDA.

## 19.3 Standard Accessories

a) Ophthalmoscope optical head with micro-aperture selection ring — Minimum 5 apertures/filters; smooth-rotating selection ring; interchangeable with matched handle.

b) Heavy-duty desk-top twin-charging cradle station — Two-slot desk cradle; charge indicator per slot; universal 100—240 VAC input.

c) Rechargeable lithium-ion battery handle — One per optical head; minimum 2 h continuous use; charge indicator.

d) Long-life replacement LED/Xenon bulbs (2 units per set) — Matched to optical head; long-life type where LED.

e) Zippered rigid protective storage pouch — Fitted foam interior; accommodates head, handle, cradle, and spares.

## 19.4 Installation Requirements

Standard mains socket 100—240 VAC for charging cradle. Suitable dedicated space in clinic/ward for cradle storage.

## 19.5 Documentation

User Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Battery Certificate (lithium-ion safety compliance), Warranty Certificate.

## 19.6 Training

Clinical technique training (minimum ½ day) — combined with indirect ophthalmoscope and retinoscope training where scheduled together.

## 19.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including battery and LED/bulb), technical support. Spares availability minimum 10 years.

# 20. BINOCULAR INDIRECT OPHTHALMOSCOPE

Quantity: 3

## 20.1 Clinical & Functional Specifications

Head-mounted binocular indirect ophthalmoscope for stereoscopic wide-field examination of the peripheral retina and vitreous, used with an auxiliary condensing lens (typically +20 D, +28 D, or +30 D). It shall provide stable, ergonomic head-mounting, comfortable interpupillary/convergence adjustment, and high-intensity coaxial illumination.

## 20.2 Engineering Technical Specifications

- Illumination: LED (preferred) or xenon/halogen; adjustable intensity 0—100 %; long-life LED ≥ 20,000 h.

- Illumination modes: Small, medium, large spot; diffuser, red-free (green), cobalt blue filters selectable.

- Optics: Binocular with adjustable interpupillary distance minimum 47—75 mm; adjustable convergence.

- Headband: Padded, adjustable circumference; balanced weight distribution.

- Power source: Rechargeable lithium-ion belt pack (mandatory); mains cable operation optional.

- Battery runtime: Minimum 2 hours continuous use per full charge.

- Charging time: Maximum 4 hours.

- Weight (head-mounted assembly): Approximately 500—800 g (excluding battery pack).

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 10943, CE MDR / FDA.

## 20.3 Standard Accessories

a) Lightweight ergonomic headband assembly with optical pod — Padded adjustable headband; balanced optical pod; interpupillary and convergence adjustment.

b) Rechargeable lithium-ion battery pack with belt clip — Belt-mounted pack; minimum 2 h runtime; charge indicator.

c) Desk-top battery charger station module — Compatible charger for belt battery; universal 100—240 VAC input.

d) Scleral depressor thimble (Adult and Pediatric sizes) — Stainless steel; smooth polished tip; supplied as a matched pair (adult + paediatric).

e) High-reflectivity teaching mirror attachment — Beam-splitter/teaching mirror for observer training; matched to headset.

f) Rigid padded protective storage case — Fitted foam interior; accommodates headset, battery, charger, scleral depressors, and accessories.

## 20.4 Installation Requirements

Standard mains socket 100—240 VAC for charger; suitable clinic storage.

## 20.5 Documentation

User Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Battery Certificate, Warranty Certificate.

## 20.6 Training

Clinical technique training (minimum 1 day) on indirect ophthalmoscopy with condensing lenses (may be combined with retinal-lens training).

## 20.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including battery, LED/bulb, and belt pack), technical support. Spares availability minimum 10 years.

# 21. STREAK RETINOSCOPE WITH RECHARGEABLE HANDLE

Quantity: 5

## 21.1 Clinical & Functional Specifications

Handheld streak retinoscope for objective determination of refractive error through observation of retinal reflex movement. It shall provide 360° streak rotation, sleeve control for plane/concave mirror effect and streak-width/divergence adjustment, and a rechargeable handle.

## 21.2 Engineering Technical Specifications

- Illumination: LED (preferred) or halogen; adjustable intensity; colour temperature typical 3200—5600 K; long-life LED ≥ 20,000 h.

- Streak rotation: Continuous 360° with smooth control ring.

- Sleeve control: Continuous adjustment of streak width/vergence and plane/concave mirror effect.

- Fixation cards: Provided for dynamic retinoscopy at multiple viewing distances.

- Handle: Rechargeable lithium-ion; compatible with matched charging base.

- Battery runtime: Minimum 2 hours continuous use per full charge.

- Charging time: Maximum 4 hours.

- Charging cradle: Desk-top; simultaneous charging of 2 handles (for twin base).

- Weight (optical head): Approximately 100—250 g.

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 10943, CE MDR / FDA.

## 21.3 Standard Accessories

a) Retinoscope optical head with 360° rotation streak control — Smooth continuous streak rotation; sleeve for vergence and mirror-effect control; interchangeable with matched handle.

b) Reusable fixation cards set for dynamic retinoscopy — Graduated cards for near-distance dynamic retinoscopy; laminated for cleaning.

c) Desk-top dual battery charging base station — Two-slot desk cradle; charge indicator per slot; universal 100—240 VAC input.

d) Rechargeable lithium-ion battery handle — One per optical head; minimum 2 h continuous use; charge indicator.

e) Rigid protective padded carrying container — Fitted foam interior; accommodates head, handle, charger, cards, and spares.

## 21.4 Installation Requirements

Standard mains socket 100—240 VAC for charging cradle; suitable clinic storage.

## 21.5 Documentation

User Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Battery Certificate, Warranty Certificate.

## 21.6 Training

Clinical technique training (minimum ½ day) — combined with direct/indirect ophthalmoscope training where scheduled together.

## 21.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including battery and LED/bulb), technical support. Spares availability minimum 10 years.

# 22. SCHIØTZ INDENTATION TONOMETER

Quantity: 1

## 22.1 Clinical & Functional Specifications

Mechanical indentation tonometer for supine intraocular pressure measurement, particularly when applanation tonometry is not feasible (bedside/OR use, corneal irregularity). It provides an analogue scale reading correlated to IOP via published conversion tables and the supplied weight set.

## 22.2 Engineering Technical Specifications

- Type: Mechanical indentation tonometer with jeweled bearing.

- Weight set: Interchangeable — 5.5 g, 7.5 g, 10.0 g (mandatory); optional 15.0 g.

- Scale: Graduated 0—20 units, readable to 0.5 units or better.

- Footplate: Polished corneal footplate; interchangeable/serviceable.

- Calibration: Test block/convex dome for zero verification.

- Materials: Stainless steel body; jeweled pointer bearing.

- Compliance: ISO 8612 (tonometers), CE where applicable.

## 22.3 Standard Accessories

a) Interchangeable measurement weights set (5.5 g, 7.5 g, 10 g) — Precision-machined weights; matched to instrument body; supplied as complete set.

b) Convex calibration test block metal dome base — Machined convex dome for zero verification; polished finish.

c) Jeweled pointer scale instrument pointer block — Precision jeweled bearing pointer; corresponds to weight-based indentation.

d) Velvet-lined rigid metal protective storage case — Fitted interior for instrument, weights, and test block; hinged lid with clasp.

## 22.4 Installation Requirements

None — portable bedside/OR/clinic instrument. Instrument disinfection compatibility per hospital infection-control protocol.

## 22.5 Documentation

User Manual, Factory Calibration Certificate, Conversion Table (Friedenwald or equivalent), Warranty Certificate.

## 22.6 Training

Clinical technique training (minimum ½ day) — combined with other tonometer training.

## 22.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects, calibration, and replacement of defective weights or pointer assembly. Spares availability minimum 10 years.

# 23. HAND-HELD TONOMETER

Quantity: 1

## 23.1 Clinical & Functional Specifications

Portable handheld digital tonometer for intraocular pressure measurement without topical anaesthesia, suitable for paediatric patients, bedside/ward use, operating room use, and screening. Rebound- or applanation-based technology as per platform.

## 23.2 Engineering Technical Specifications

- Technology: Rebound tonometry or handheld applanation.

- Measurement range: Minimum 5 to 50 mmHg.

- Accuracy: Within ±2 mmHg (5—30 mmHg) typical; ±3 mmHg (30—50 mmHg).

- Measurement mode: Automatic averaging of multiple readings; reliability indication.

- Anaesthesia: Not required for rebound type; topical for applanation type.

- Probe: Disposable single-use probe (rebound) or reusable sterilisable prism (applanation) per platform.

- Display: Digital LCD/OLED, minimum 1-inch, backlit.

- Battery: Rechargeable lithium-ion or standard replaceable cells.

- Battery runtime: Minimum 500 measurements per charge/battery set.

- Weight: Approximately 150—300 g.

- Data storage: Minimum 10 measurements internal memory.

- Connectivity: USB and/or Bluetooth (preferred).

- Compliance: IEC 60601-1, IEC 60601-1-2, ISO 8612, CE MDR / FDA.

## 23.3 Standard Accessories

a) Rechargeable lithium-ion battery module handle base — Integrated or replaceable; minimum 500-measurement runtime.

b) Desktop charging dock station with mains power adapter — Universal 100—240 VAC adapter; charge indicator.

c) Reusable or disposable tonometer tips verification kit — Initial supply of probes appropriate to platform (minimum starter pack); calibration verification kit.

d) Standard calibration check bar weight block — Certified verification device.

e) Rigid plastic hard protective transit shell — Fitted case for instrument, dock, probes, and accessories.

## 23.4 Installation Requirements

Standard mains socket 100—240 VAC for charging; clinic/ward storage.

## 23.5 Documentation

User Manual, Parts Catalogue, Factory Test Certificate, Calibration Certificate, Electrical Safety Test Report, Warranty Certificate.

## 23.6 Training

Clinical technique training (minimum ½ day).

## 23.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including battery), calibration, technical support. Spares availability minimum 10 years.

# 24. HERTEL EXOPHTHALMOMETER

Quantity: 1

## 24.1 Clinical & Functional Specifications

Optical instrument for measurement of anteroposterior globe position (proptosis/enophthalmos) relative to the lateral orbital rims, used in the assessment of thyroid eye disease, orbital tumours, trauma, and post-surgical follow-up.

## 24.2 Engineering Technical Specifications

- Type: Hertel-style parallel-mirror exophthalmometer.

- Base measurement range: Minimum 85—115 mm (interorbital base scale).

- Proptosis measurement range: Minimum 0—30 mm.

- Reading resolution: 1 mm or better.

- Optics: Parallel high-reflectivity mirrors with engraved millimetre scale.

- Orbital rests: Soft-padded, adjustable position.

- Materials: Anodised aluminium or stainless-steel body.

- Compliance: CE where applicable; not electrically powered.

## 24.3 Standard Accessories

a) Parallel high-reflectivity mirror alignment system with scale — Precision-aligned mirrors with clearly engraved millimetre scale.

b) Engraved millimetre measurement tracking rules (both sides) — Left and right eye scales; engraved and paint-filled for clarity.

c) Soft-padded orbital rim rests / stabilisers — Skin-friendly padding; cleanable with hospital disinfectants.

d) Rigid velvet-lined protective container storage box — Fitted interior; hinged lid with clasp.

## 24.4 Installation Requirements

None — portable clinic instrument. Disinfection compatibility with standard clinic wipes.

## 24.5 Documentation

User Manual, Factory Calibration Certificate, Warranty Certificate.

## 24.6 Training

Clinical technique training (minimum ½ day, combined with orbital assessment training).

## 24.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects and calibration. Spares availability minimum 10 years.

# 25. OPHTHALMIC EXAMINATION UNIT (REFRACTION CHAIR AND STAND)

Quantity: 1

## 25.1 Clinical & Functional Specifications

Combined refraction lane unit comprising a motorised patient examination chair and an equipment stand supporting a phoropter/refractor arm, slit-lamp table (or auto-refractometer table), projector shelf, and integrated overhead reading lamp, with a central power distribution/control console. It shall provide a complete integrated refraction workstation with electrical outlets for the connected diagnostic instruments.

## 25.2 Engineering Technical Specifications

Patient chair

- Height adjustment: Motorised, travel range minimum 500—700 mm to maximum 900—1000 mm from floor.

- Foot-pedal control: Sealed membrane foot switch, height up/down.

- Load capacity: Minimum 150 kg.

- Upholstery: Fire-retardant, wipe-clean vinyl or medical-grade synthetic leather; seam-sealed.

- Armrests: Foldable/retractable.

- Rotation: 360° rotation with lock.

Equipment stand / column

- Configuration: Column-based with counter-balanced phoropter arm, motorised or manual slit-lamp/instrument table, projector shelf with swivel arm.

- Phoropter arm: Counter-balanced; travel and swivel to position phoropter in front of patient.

- Instrument table: Motorised height adjustment preferred; travel minimum 200 mm.

- Load capacity (instrument table): Minimum 20 kg.

- Overhead reading lamp: Low-voltage LED; adjustable arm; adjustable intensity.

- Power console: Integrated switch panel with individually switched outlets for connected instruments (minimum 4 outlets), all fused and earth-leakage protected.

Electrical

- Power supply: 100—240 VAC, 50/60 Hz.

- Total connected load: Rated to accommodate typical refraction lane equipment (chair + slit lamp + auto-refractometer + projector + phoropter illumination).

- Compliance: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-52 (medical beds/chairs) where applicable, CE MDR / FDA.

## 25.3 Standard Accessories

a) Motorised height-adjustable patient examination chair with foot pedal — Sealed foot-pedal control; load capacity minimum 150 kg; wipe-clean upholstery; foldable armrests; 360° rotation with lock.

b) Counter-balanced phoropter arm assembly pivot bracket — Counter-balanced; smooth articulation; compatible with standard phoropter mounting.

c) Projector shelf with integrated swivel arm holder mounting — Adjustable swivel arm; wall or column mount option; compatible with standard acuity projectors.

d) Integrated low-voltage overhead reading lamp fixture — LED; adjustable arm; adjustable intensity; long-life LED ≥ 20,000 h.

e) Main power control console panel with equipment sockets — Minimum 4 individually switched outlets; fused and earth-leakage protected; master on/off switch; illuminated indicators.

## 25.4 Installation Requirements

Refraction lane with minimum floor area approximately 3.0 × 2.5 m; wall-projection distance for acuity chart approximately 3—6 m or per projector requirement (with mirror system if compact lane). Single-phase 100—240 VAC dedicated circuit; earth-leakage protected. Level, load-bearing floor. Wall reinforcement for phoropter/projector mounting where applicable.

## 25.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), Load Test Certificate (chair), Installation & Commissioning Reports, PM Schedule.

## 25.6 Training

Clinical/operational training (minimum ½ day) on chair, stand, and console operation. Biomedical engineering training (minimum 1 day) on motor service, upholstery cleaning protocol, and electrical safety.

## 25.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including chair motor, upholstery for manufacturing defect, and lamp), PM, technical support. Spares availability minimum 10 years.

ORTHOPTIC, STRABISMUS, COLOUR VISION, AND VISUAL ACUITY ASSESSMENT INSTRUMENTS

# 26. VISUAL ACUITY PROJECTOR / DIGITAL ACUITY CHART

Quantity: 5

## 26.1 Clinical & Functional Specifications

Digital LCD/LED visual acuity chart (or optical projector) for distance visual acuity testing during subjective refraction, with programmable test sequences and multiple optotype sets (Snellen, LogMAR, tumbling E, Landolt C, numbers, children's pictures). It shall provide randomised optotype presentation to prevent memorisation, red-green and polarised testing for binocular balance, and remote-controlled operation from the examiner's position.

## 26.2 Engineering Technical Specifications

- Display type: High-resolution LED/LCD panel (mandatory); alternatively high-brightness DLP projector.

- Display size (LCD/LED): Minimum 19-inch diagonal; 24-inch or greater preferred.

- Resolution: Minimum 1920 × 1080 (Full HD).

- Testing distance: Configurable, typical 2.5 m to 6.0 m.

- Optotype types: Snellen, LogMAR/ETDRS, tumbling E, Landolt C, numbers, children's pictures — minimum 4 sets selectable.

- Acuity range: Minimum 20/200 (0.1) to 20/10 (2.0), or equivalent LogMAR scale.

- Test modes: Single letter, single line, full chart, randomised; masking of adjacent lines/letters.

- Special tests: Red-green duochrome, polarised binocular balance, astigmatism dial, cross-cylinder targets, contrast sensitivity (where supported), fixation targets for children.

- Contrast: Adjustable; minimum 5 contrast steps.

- Control: Wireless infrared or RF remote handset; range minimum 5 m.

- Connectivity: Integration port (RS-232/USB/Ethernet) for linkage with digital phoropter systems.

- Power supply: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, ISO 8596 (visual acuity), CE MDR / FDA.

## 26.3 Standard Accessories

a) High-resolution LED/LCD digital chart panel display — Minimum 19-inch; Full HD; anti-glare screen; wall-mountable.

b) Multi-functional wireless infrared remote controller handset — Backlit keys; range minimum 5 m; supplied with batteries.

c) Wall-mounting hardware bracket and levelling alignment kit — VESA-compatible or manufacturer-specific bracket; supplied with fixings and levelling tool.

d) Red-green polarised glasses filter set (2 pairs) — For duochrome and polarised binocular balance testing; scratch-resistant lenses.

e) Interface cable for digital refraction system linkage — RS-232 / USB / Ethernet cable as per platform; minimum 3 m length.

## 26.4 Installation Requirements

Refraction lane with wall-projection or wall-mounting distance appropriate to configured test distance (2.5—6.0 m direct, or shorter with mirror system). Wall reinforcement sufficient for panel weight. Standard mains socket 100—240 VAC. Controllable ambient lighting.

## 26.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Installation Report, Warranty Certificate.

## 26.6 Training

Clinical/operational training (minimum ½ day) on optotype sequences and remote control operation.

## 26.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including remote handset and panel), software updates, technical support. Spares availability minimum 10 years.

# 27. TRIAL LENS SET WITH TRIAL FRAME

Quantity: 5

## 27.1 Clinical & Functional Specifications

Complete trial lens set with universal adjustable trial frame(s) for subjective refraction, low-vision assessment, contact-lens over-refraction, and paediatric refraction. Lenses shall include spherical (positive and negative), cylindrical (positive and negative), prismatic, and auxiliary lenses/occluders, in a full graduated range suitable for routine and specialised ophthalmic refraction.

## 27.2 Engineering Technical Specifications

Trial frame

- Type: Universal adjustable; multi-cell (minimum 3 cells per eye — sphere, cylinder, auxiliary).

- Interpupillary distance (IPD) adjustment: Minimum 48—80 mm, continuously adjustable.

- Vertex distance adjustment: Adjustable; range minimum 0—20 mm.

- Pantoscopic tilt: Adjustable minimum ±10°.

- Cylinder axis rotation: 0°—180°, 1° increments, engraved scale.

- Bridge/nose pad: Adjustable; anatomically shaped; hypoallergenic materials.

- Temples: Adjustable length; sculpted for stability.

- Weight: Minimised, typical 40—70 g.

Paediatric trial frame

- Type: Fixed or limited-adjustment interpupillary distance; child-appropriate ergonomics.

- IPD range: Suitable for paediatric use (typical 46—58 mm).

Trial lens set

- Sphere range: Minimum —20.00 D to +20.00 D in 0.25 D steps in the central range (typical 0.25 D steps from —5.00 to +5.00 D; larger steps beyond).

- Cylinder range: Minimum —6.00 D to +6.00 D in 0.25 D steps (positive and negative cylinders).

- Prism range: Minimum 0.5 Δ to 10 Δ in graduated steps.

- Auxiliary lenses: Plano occluder, pinhole occluder, Maddox rod (red/white), red/green filters, cross-cylinder, polarising filters, plano transparent, plano frosted.

- Lens material: High-quality optical crown glass or optical-grade CR-39; anti-reflective coating preferred.

- Lens rim: Metal ring with clear numeric marking (dioptre value and sign colour-coded — typically red for negative, black for positive).

- Lens diameter: Standard trial lens diameter (typical 38 mm).

Case

- Type: Velvet-lined wooden or aluminium case with labelled compartments for each lens.

- Capacity: Minimum 232 lenses (typical full-set standard configuration).

- Compliance: ISO 9801 (trial lenses), CE where applicable.

## 27.3 Standard Accessories

a) Universal adjustable trial frame with multi-axis vertex adjustment — Adjustable IPD (48—80 mm), vertex distance, pantoscopic tilt; anatomical bridge and adjustable temples.

b) Pediatric trial frame with fixed interpupillary settings — Fixed or limited-adjustment IPD for paediatric use; lightweight construction.

c) Reusable trial lenses (graduated spheres, cylinders, prisms, occluders) — Full graduated set as above; colour-coded rims; clearly engraved values.

d) Velvet-lined wooden or aluminium case storage tray — Labelled compartments for each lens; hinged lid with secure clasp.

e) Microfibre optical lens cleaning cloth — Non-abrasive; supplied with initial stock.

## 27.4 Installation Requirements

None — clinic-portable set. Suitable dedicated storage space; controlled humidity to prevent lens fungal growth (30—70 % RH preferred).

## 27.5 Documentation

User Manual, Lens Inventory List (with dioptre values), Factory Calibration Certificate, Warranty Certificate.

## 27.6 Training

Clinical/refraction training (minimum ½ day) on trial frame adjustment and subjective refraction technique.

## 27.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects and free replacement of lenses with dioptre-value deviation outside specification. Spares availability minimum 10 years.

# 28. SYNOPTOPHORE (MAJOR AMBLYOSCOPE)

Quantity: 1

## 28.1 Clinical & Functional Specifications

Major amblyoscope (haploscope) for assessment and treatment of binocular vision anomalies, strabismus, amblyopia, and abnormal retinal correspondence. It shall enable independent presentation of images to each eye with independently adjustable tube position for measurement of angle of deviation (horizontal, vertical, torsional), and shall provide flashing/alternating illumination for anti-suppression and fusion training.

## 28.2 Engineering Technical Specifications

- Type: Twin-tube haploscopic amblyoscope with independent illumination and angular adjustment for each tube.

- Horizontal deviation range: Minimum —40° to +40°.

- Vertical deviation range: Minimum ±15°.

- Torsional (cyclorotation) range: Minimum ±20°.

- Angular reading resolution: 1° or better; scales engraved on adjustment arcs.

- Illumination: Independent per tube; adjustable intensity; auto-flash mode with adjustable flash rate (typical 0.5—5 Hz range) for anti-suppression.

- Chin/forehead rest: Adjustable; disposable chin-rest paper compatible.

- Slide inserts: Compatible with standard diagnostic slide sets (grades I, II, III — simultaneous perception, fusion, stereopsis).

- Haidinger brush attachment: Optional/integrated for macular function assessment.

- Power supply: 100—240 VAC, 50/60 Hz.

- Compliance: IEC 60601-1, IEC 60601-1-2, CE MDR / FDA.

## 28.3 Standard Accessories

a) Twelve pairs of standard diagnostic slide inserts (starter kit) — Grades I, II, III (simultaneous perception, fusion, stereopsis) slides; matched pairs; supplied in labelled storage box.

b) Automatic flash system timer control module — Adjustable flash rate; independent per-tube illumination flashing for anti-suppression training.

c) Integrated Haidinger brush effect optical attachment — For macular function assessment.

d) Adjustable headrest and chin-rest stabiliser bars — Adjustable height; comfortable padded rests.

e) Replacement illumination lamp bulbs set — Minimum 2 spare bulbs/LED modules matched to instrument.

## 28.4 Installation Requirements

Orthoptic clinic room with controllable ambient lighting; ambient 18—26 °C, 30—70 % RH. Single-phase 100—240 VAC standard socket. Stable instrument table.

## 28.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, Installation Report, Warranty Certificate.

## 28.6 Training

Clinical/orthoptic training (minimum 1 day) on measurement of angle, fusion/stereopsis testing, and anti-suppression exercises.

## 28.7 Warranty

Comprehensive minimum two (2) years including labour, all spare parts (including lamps and flash module), technical support. Spares availability minimum 10 years.

# 29. HESS SCREEN

Quantity: 1

## 29.1 Clinical & Functional Specifications

Diagnostic apparatus for evaluation of ocular motility and identification of underacting/overacting extraocular muscles, particularly in paralytic and restrictive strabismus. The patient views a red/green tangent screen through red/green dissociating goggles and localises presented targets, allowing plotting of the deviation across cardinal gaze positions.

## 29.2 Engineering Technical Specifications

- Screen type: Black-backed calibrated tangent grid screen (wall-mounted mat or fabric).

- Grid pattern: Standard Hess grid — central point plus points at 15° and 30° in the eight cardinal and oblique positions.

- Grid colour: High-contrast red on black (or grid pattern per standard Hess design).

- Screen size: Approximately 1.0 × 1.0 m (standard).

- Test distance: 50 cm (standard Hess distance).

- Red/green goggles: Dissociating red-green goggles (adult size); reusable, cleanable.

- Light pointers: Handheld red and green torch-style pointers for examiner and patient use.

- Plotting charts: Pre-printed patient plotting chart sheets.

- Compliance: CE where applicable.

## 29.3 Standard Accessories

a) Black-backed calibrated tangent grid screen wall mat — Standard Hess grid; approximately 1.0 × 1.0 m; wall-mountable.

b) Red/Green diagnostic examination goggles frame — Adult size dissociating red-green goggles; cleanable frame.

c) Handheld red light pointer torch handle — Red LED pointer; battery-operated; ergonomic handle.

d) Handheld green light pointer torch handle — Green LED pointer; battery-operated; ergonomic handle.

e) Patient plotting chart sheets pack (minimum 50 sheets) — Pre-printed plotting sheets; matched to screen grid pattern.

## 29.4 Installation Requirements

Orthoptic clinic with wall for screen mounting at eye level of seated patient at 50 cm test distance. Controllable ambient lighting (dim environment).

## 29.5 Documentation

User Manual, Warranty Certificate, Plotting Chart Reorder Reference.

## 29.6 Training

Clinical/orthoptic training (minimum ½ day) on Hess screen technique and chart interpretation.

## 29.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects, screen mat, goggles, and pointers.

# 30. PRISM BAR SET (HORIZONTAL + VERTICAL)

Quantity: 1

## 30.1 Clinical & Functional Specifications

Calibrated horizontal and vertical prism bar set for measurement of ocular deviations in strabismus by prism cover test, and for measurement of fusional vergence amplitudes.

## 30.2 Engineering Technical Specifications

- Horizontal prism bar range: Minimum 1 Δ to 40 Δ; graduated in progressively increasing steps (typical 1, 2, 4, 6, 8, 10, 12, 14, 16, 18, 20, 25, 30, 35, 40 Δ).

- Vertical prism bar range: Minimum 1 Δ to 25 Δ; graduated in progressively increasing steps (typical 1, 2, 3, 4, 6, 8, 10, 12, 14, 16, 18, 20, 25 Δ).

- Prism material: Optical-grade acrylic (PMMA) or optical crown glass; anti-reflective coating preferred.

- Calibration accuracy: Within ±0.5 Δ of nominal value.

- Bar construction: Rigid frame; each prism securely fixed with clearly engraved value.

- Compliance: ISO 8624 or equivalent optical standard where applicable; CE.

## 30.3 Standard Accessories

a) Calibrated horizontal prism bar stick (1—40 Δ) — Full graduated horizontal bar; engraved values; scratch-resistant.

b) Calibrated vertical prism bar stick (1—25 Δ) — Full graduated vertical bar; engraved values; scratch-resistant.

c) Rigid protective velvet-lined storage case — Fitted interior for both bars; hinged lid with clasp.

d) Optical microfibre cleaning cloth — Non-abrasive; supplied with initial stock.

## 30.4 Installation Requirements

None — clinic-portable. Suitable dedicated storage; controlled humidity (30—70 % RH preferred).

## 30.5 Documentation

User Manual, Factory Calibration Certificate, Warranty Certificate.

## 30.6 Training

Clinical/orthoptic training (minimum ½ day) on prism cover test technique.

## 30.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects and calibration.

# 31. TITMUS STEREOACUITY TEST

Quantity: 1

## 31.1 Clinical & Functional Specifications

Polarised stereoacuity test booklet with accompanying polarised viewers for measurement of gross and fine stereopsis (typically stereo-fly for gross stereopsis and graded circles/animals for fine stereopsis down to approximately 40 seconds of arc).

## 31.2 Engineering Technical Specifications

- Test type: Polarised (vectographic) stereoacuity test.

- Gross stereopsis test: Stereo-fly (approximately 3000 seconds of arc).

- Fine stereopsis test: Graded circles (typical 800 down to 40 seconds of arc).

- Paediatric component: Animal series (typical 400, 200, 100 seconds of arc).

- Test distance: 40 cm (standard).

- Viewer: Polarised viewer glasses (adult and paediatric sizes).

- Test plate material: Polarised laminated plates; wipe-clean surface.

- Compliance: ISO test standard for stereopsis where applicable; CE.

## 31.3 Standard Accessories

a) Standard stereo-fly booklet test plates — Complete Titmus-style plate set with gross fly, animals, and graded circles.

b) Child-size polarised viewer glasses frame — Paediatric-sized polarised glasses.

c) Adult-size polarised viewer glasses frame — Adult-sized polarised glasses.

d) Instruction user manual and scoring reference chart — Includes normative data and scoring interpretation.

## 31.4 Installation Requirements

None — clinic-portable. Controlled ambient lighting for consistent testing.

## 31.5 Documentation

User Manual, Warranty Certificate.

## 31.6 Training

Clinical/orthoptic technique training (minimum ½ day, combined with other stereopsis tests).

## 31.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 32. MADDOX WING

Quantity: 1

## 32.1 Clinical & Functional Specifications

Handheld dissociating instrument for measurement of heterophoria at near (approximately 33 cm) — horizontal, vertical, and torsional (cyclophoria) components — using septum dissociation.

## 32.2 Engineering Technical Specifications

- Type: Handheld Maddox wing instrument.

- Test distance: Approximately 33 cm (near).

- Measurement scales: Horizontal (esophoria/exophoria), vertical (hyperphoria/hypophoria), torsional (cyclophoria) — clearly labelled and calibrated.

- Horizontal scale range: Minimum ±25 Δ.

- Vertical scale range: Minimum ±10 Δ.

- Torsional scale range: Minimum ±30°.

- Construction: Rigid metal or high-impact polymer body; septum divider.

- Forehead rests: Adjustable/comfortable padded rests.

- Compliance: CE where applicable.

## 32.3 Standard Accessories

a) Calibrated horizontal/vertical/torsion scale plate template — Clearly labelled scales with engraved calibration marks.

b) Adjustable septum separator fin divider sheet — Central septum for right/left eye dissociation.

c) Integrated twin forehead rest pad plates — Padded forehead rests; skin-friendly, cleanable.

d) Rigid protective clear storage sleeve container — Fitted case; hinged lid or slip-cover.

## 32.4 Installation Requirements

None — clinic-portable.

## 32.5 Documentation

User Manual, Warranty Certificate.

## 32.6 Training

Clinical/orthoptic technique training (minimum ½ day, combined with other heterophoria tests).

## 32.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 33. ISHIHARA TEST FOR COLOUR VISION

Quantity: 1

## 33.1 Clinical & Functional Specifications

Pseudo-isochromatic plate test for screening and diagnosis of congenital red-green colour vision deficiencies (protan/deutan).

## 33.2 Engineering Technical Specifications

- Type: Standard Ishihara pseudo-isochromatic plate booklet.

- Number of plates: Minimum 38-plate configuration (mandatory); optional 14- and 24-plate abridged editions supplied additionally where applicable.

- Plate categories: Introductory plate, transformation plates, vanishing plates, hidden digit plates, classification plates (protan/deutan differentiation), tracing plates for illiterate/paediatric patients.

- Test distance: 75 cm (standard).

- Illumination requirement: Daylight or daylight-equivalent illumination (approximately 6500 K) — specified in user manual.

- Binding: Hardcover book format; wipe-clean laminated plates.

- Compliance: CE where applicable.

## 33.3 Standard Accessories

a) Standard 38-plate configuration pseudo-isochromatic book — Full 38-plate Ishihara set; laminated plates; hardcover binding.

b) Non-marring soft bristle tracing brush pointer — For paediatric/illiterate tracing plates.

c) Opaque black individual plate mask card sheet — Mask for isolating individual plates during testing.

d) Scoring evaluation results sheet pad — Pre-printed scoring sheets.

## 33.4 Installation Requirements

None — clinic-portable. Daylight or daylight-equivalent illumination required in testing area.

## 33.5 Documentation

User Manual, Warranty Certificate.

## 33.6 Training

Clinical technique training (minimum ½ day, combined with other colour vision tests).

## 33.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 34. FRISBY STEREOTEST

Quantity: 1

## 34.1 Clinical & Functional Specifications

Real-depth (non-polarised, non-glasses) stereoacuity test using three transparent perspex plates of graded thickness with printed random-shape patterns, for measurement of stereopsis particularly in young children and where polarised viewer wear is not feasible.

## 34.2 Engineering Technical Specifications

- Type: Real-depth stereoacuity test (no glasses required).

- Test plates: Three transparent perspex plates of graded thickness (typical 6 mm, 3 mm, 1 mm or manufacturer-standard set).

- Stereopsis range: Approximately 600 down to 15 seconds of arc, dependent on test distance (30 cm to 80 cm range).

- Plate material: Optical-grade perspex/acrylic; wipe-clean surface.

- Compliance: CE where applicable.

## 34.3 Standard Accessories

a) Three clear perspex test plates (different thickness profiles) — Complete graded set; wipe-clean surface.

b) High-visibility measurement validation fold-rule bar — For accurate test-distance measurement.

c) Velvet-lined presentation wooden or plastic storage case — Fitted interior for plates and rule.

d) Scoring and evaluation reference matrix sheet — Distance-to-arc-seconds reference chart.

## 34.4 Installation Requirements

None — clinic-portable.

## 34.5 Documentation

User Manual, Reference Chart, Warranty Certificate.

## 34.6 Training

Clinical/orthoptic training (minimum ½ day, combined with other stereopsis tests).

## 34.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 35. RAF NEAR-POINT RULE

Quantity: 1

## 35.1 Clinical & Functional Specifications

Royal Air Force (RAF) near-point rule for measurement of near point of accommodation (NPA) and near point of convergence (NPC), with rotating target cube presenting graded accommodative targets.

## 35.2 Engineering Technical Specifications

- Type: Sliding-target rule with rotating four-sided target cube.

- Rule length / measurement range: Minimum 0—50 cm.

- Scales: Linear distance (cm) and equivalent dioptric power scale.

- Target cube: Four faces with graded accommodative targets (typical N5, N8, N10, and a fixation dot/line).

- Cheek rest: Anatomically shaped cheek/cheekbone rest for stability.

- Construction: Rigid metal or high-impact polymer.

- Compliance: CE where applicable.

## 35.3 Standard Accessories

a) Four-sided rotatable accommodation test cube target block — Rotating target cube with graded accommodative targets.

b) Calibrated sliding convergence rule bar track — Smooth-sliding calibrated track with linear and dioptric scales.

c) Comfortable shaped cheek-rest stabiliser pad mount — Anatomically shaped cheek rest; skin-friendly.

d) Protective fabric transport slip pouch — Soft slip-pouch for storage/transport.

## 35.4 Installation Requirements

None — clinic-portable.

## 35.5 Documentation

User Manual, Warranty Certificate.

## 35.6 Training

Clinical/orthoptic training (minimum ½ day, combined with other near-vision tests).

## 35.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 36. WORTH FOUR-DOT TEST WITH RED-GREEN GOGGLES

Quantity: 1

## 36.1 Clinical & Functional Specifications

Handheld illuminated four-dot test (two green, one red, one white) viewed through red-green dissociating goggles for assessment of binocular sensory status — fusion, suppression, and diplopia — at distance and near.

## 36.2 Engineering Technical Specifications

- Type: Handheld battery-operated illuminated Worth four-dot torch.

- Dot configuration: Standard Worth pattern — two green dots (top and bottom or per convention), one red dot, one white dot.

- Illumination: LED (preferred); adjustable intensity or on/off; battery-operated.

- Battery: Standard alkaline AA (mandatory) or rechargeable equivalent.

- Dissociating goggles: Red/green goggles (adult size, paediatric size preferred).

- Compliance: CE where applicable.

## 36.3 Standard Accessories

a) Handheld battery-operated four-dot light torch handle — LED illumination; ergonomic grip; on/off switch.

b) Red/Green diagnostic viewer goggles frame unit — Dissociating red/green goggles; cleanable frame.

c) Rechargeable or standard alkaline AA battery pack — Initial supply of batteries.

d) Rigid protective plastic carrying container — Fitted case for torch, goggles, and batteries.

## 36.4 Installation Requirements

None — clinic-portable. Controllable ambient lighting for testing.

## 36.5 Documentation

User Manual, Warranty Certificate.

## 36.6 Training

Clinical/orthoptic training (minimum ½ day, combined with other binocular sensory tests).

## 36.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 37. KAY PICTURE TEST

Quantity: 1

## 37.1 Clinical & Functional Specifications

Paediatric visual acuity test using familiar picture optotypes (Kay pictures) for assessment of visual acuity in pre-verbal and pre-literate children, with matching keycard for children unable to name pictures verbally.

## 37.2 Engineering Technical Specifications

- Type: Standardised picture-optotype acuity test booklet with matching keycard.

- Acuity range: Minimum approximately 6/60 (LogMAR 1.0) to 6/6 (LogMAR 0.0), or as per Kay standard.

- Test distance: Configurable (typical 3 m or 6 m); near-version optional.

- Optotype set: Standard Kay picture set (recognisable everyday objects/animals).

- Matching keycard: Handheld keycard with the pictures for non-verbal matching.

- Booklet material: Wipe-clean laminated pages; robust binding.

- Compliance: ISO 8596 where applicable; CE.

## 37.3 Standard Accessories

a) Specialised child-picture acuity booklet plates binder — Kay picture set; laminated pages; robust binding.

b) Handheld matching keycard identification board plate — Matching keycard for non-verbal identification.

c) Distance evaluation measurement guide cord track — Test-distance measurement cord/tape.

d) Padded storage and carrying wallet pouch — Fitted padded wallet for booklet and keycard.

## 37.4 Installation Requirements

None — clinic-portable. Controllable ambient lighting.

## 37.5 Documentation

User Manual, Warranty Certificate.

## 37.6 Training

Clinical/orthoptic training (minimum ½ day, combined with paediatric acuity training).

## 37.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 38. LogMAR NEAR VISUAL ACUITY CHART

Quantity: 1

## 38.1 Clinical & Functional Specifications

Standardised near visual acuity chart in LogMAR (logarithm of the minimum angle of resolution) format for assessment of near visual acuity at a fixed test distance, using equal-progression optotype sizing and equal letters per line, suitable for adult and (with picture equivalents) paediatric assessment.

## 38.2 Engineering Technical Specifications

- Type: Double-sided high-contrast near-vision LogMAR chart.

- Acuity range: Minimum LogMAR +1.0 to —0.3 (equivalent Snellen approximately 6/60 to 6/3).

- Optotype progression: Equal LogMAR progression of 0.1 per line.

- Letters per line: 5 (standard LogMAR/ETDRS-style).

- Test distance: Fixed near distance (typical 40 cm) with attached umbilical cord.

- Chart material: High-contrast laminated plate; wipe-clean surface; non-glare finish.

- Compliance: ISO 8596; CE.

## 38.3 Standard Accessories

a) Double-sided high-contrast near-vision card plate panel — LogMAR chart both sides (e.g., letters one side, numbers or pictures other side).

b) Fixed-distance patient umbilical line cord — Attached cord for consistent 40 cm test distance.

c) Lightweight plastic desk support display easel stand — Foldable easel for tabletop use.

d) Protective clear acetate wipe-clean overlay sheets — Removable protective overlay for extended service life.

## 38.4 Installation Requirements

None — clinic-portable. Consistent illumination at chart plane recommended (approximately 480 lux or per LogMAR standard).

## 38.5 Documentation

User Manual, Warranty Certificate.

## 38.6 Training

Clinical/refraction training (minimum ½ day, combined with other acuity assessment training).

## 38.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 39. LEA GRATING PADDLES

Quantity: 1

## 39.1 Clinical & Functional Specifications

Set of graded high-contrast grating paddles for preferential-looking visual acuity assessment in infants, young children, and non-verbal patients, using the preferential-looking response to detect the highest resolvable spatial frequency.

## 39.2 Engineering Technical Specifications

- Type: Preferential-looking grating paddle set.

- Paddles: Minimum 4 graded paddles of increasing spatial frequency, plus a blank grey control paddle.

- Grating pattern: High-contrast square-wave grating (black/white) on one face; homogeneous grey (equal mean luminance) on reverse face.

- Spatial frequency range: Graded set (typical 0.25, 0.5, 1.0, 2.0, 4.0, 8.0 cycles per centimetre or manufacturer-standard progression).

- Handle: Ergonomic handle for examiner presentation.

- Material: Rigid printed board; laminated/wipe-clean surface.

- Compliance: CE where applicable.

## 39.3 Standard Accessories

a) Graduated high-contrast grating pattern paddles (set of 4) — Full graded set as above; laminated wipe-clean surface; ergonomic handles.

b) Soft micro-fibre protective storage pouch sleeves — Individual pouches per paddle for scratch protection.

c) Patient testing instruction protocol matrix sheet — Standardised testing protocol and reference/normative data.

## 39.4 Installation Requirements

None — clinic-portable. Controllable ambient lighting for consistent testing.

## 39.5 Documentation

User Manual, Testing Protocol Reference, Warranty Certificate.

## 39.6 Training

Clinical/orthoptic paediatric acuity training (minimum ½ day, combined with other paediatric acuity assessment training).

## 39.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

# 40. SPIELMANN TRANSLUCENT OCCLUDER

Quantity: 1

## 40.1 Clinical & Functional Specifications

Handheld frosted/translucent occluder that blocks form vision while allowing light perception, preventing "cheating" or fixation with the occluded eye while avoiding the total light-deprivation effect of an opaque occluder. Used in cover testing, refraction, and orthoptic assessment.

## 40.2 Engineering Technical Specifications

- Type: Handheld translucent (frosted) occluder wand.

- Occluder diameter: Approximately 40—60 mm.

- Optical property: Uniform light transmission with complete form obstruction (frosted diffuser).

- Handle: Ergonomic non-slip moulded handle; length approximately 100—140 mm.

- Material: Optical-grade frosted acrylic or ground glass diffuser; polymer handle.

- Compliance: CE where applicable.

## 40.3 Standard Accessories

a) Frosted translucent examination wand handle stick — Optical-grade frosted diffuser; polished edge.

b) Ergonomic non-slip moulded finger grip shaft — Comfortable ergonomic handle.

c) Protective soft vinyl storage slip case — Fitted slip case for handling and storage.

## 40.4 Installation Requirements

None — clinic-portable.

## 40.5 Documentation

User Manual, Warranty Certificate.

## 40.6 Training

Clinical/orthoptic training (minimum ¼ day, combined with cover test training).

## 40.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects.

OPHTHALMIC DIAGNOSTIC & LASER CONTACT LENSES

General note applicable to all lenses in this batch: All contact lenses shall be optical-grade with anti-reflective coating on non-contact surfaces where applicable, be compatible with standard ophthalmic coupling fluids (methylcellulose or equivalent), and be reprocessable in accordance with published ophthalmic contact-lens disinfection protocols. Materials in contact with the ocular surface shall be biocompatible per ISO 10993 and non-toxic. All lenses shall be supplied with a factory certificate of optical properties.

# 41. DIRECT GONIOSCOPY LENS (KOEPPE-TYPE)

Quantity: 1

## 41.1 Clinical & Functional Specifications

Direct (Koeppe-type) diagnostic contact lens for direct visualisation of the anterior chamber angle in the supine patient, primarily used in the operating room, examination under anaesthesia (particularly paediatric glaucoma assessment), and educational settings. It provides an undistorted panoramic view of the angle without the mirror-reflection required by indirect gonioscopy.

## 41.2 Engineering Technical Specifications

- Type: Direct gonioscopy lens (Koeppe-type), dome-shaped diagnostic contact lens.

- Viewing method: Direct (non-mirror) visualisation, used with binocular microscope or handheld loupe and separate light source.

- Field of view: Full 360° panoramic view of the anterior chamber angle.

- Magnification factor: Typical 1.5× (as per lens design).

- Contact diameter: Approximately 16—20 mm (adult) or manufacturer-standard size range.

- Optical material: Optical-grade PMMA or equivalent biocompatible optical polymer.

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Anti-reflective coating: On non-contact optical surface (preferred).

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993 (biocompatibility), ISO 15004 (ophthalmic instruments — fundamental requirements), CE MDR / FDA.

## 41.3 Standard Accessories

a) Direct view Koeppe diagnostic lens element dome — Precision-moulded dome; polished contact surface; engraved or laser-marked lens identification.

b) Rigid clear plastic protective storage shell box — Fitted interior with soft cushioning; hinged/snap-fit lid; individual compartment for lens.

c) Optical microfibre cleaning cloth ring — Non-abrasive; chemically compatible with standard ophthalmic cleaning agents.

## 41.4 Installation Requirements

None — clinic/OR-portable. Availability of standard ophthalmic coupling fluid and reprocessing facilities per hospital infection-control protocol.

## 41.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate, Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 41.6 Training

Clinical technique training (minimum ½ day, combined with other gonioscopy lens training).

## 41.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, mechanical, and coating). Spares availability minimum 10 years.

# 42. GOLDMANN THREE-MIRROR LENS

Quantity: 1

## 42.1 Clinical & Functional Specifications

Universal Goldmann three-mirror diagnostic contact lens for slit-lamp examination of the anterior chamber angle (gonioscopy), the peripheral retina, the equatorial retina, and the central fundus. The lens combines a central direct-viewing element with three internal mirrors at different inclinations to permit systematic examination of successive posterior-segment zones.

## 42.2 Engineering Technical Specifications

- Type: Universal Goldmann three-mirror contact lens.

- Mirror configuration: Three internal mirrors at standard Goldmann inclinations for:

-   Central fundus (direct central lens element)

-   Equatorial retina (mirror at approximately 73°)

-   Peripheral retina (mirror at approximately 67°)

-   Anterior chamber angle / iridocorneal angle (mirror at approximately 59°)

- Field of view per mirror: As per standard Goldmann geometry.

- Contact diameter: Standard adult size (approximately 17—18 mm).

- Optical material: Optical-grade glass or biocompatible optical polymer; PMMA acceptable.

- Anti-reflective coating: On non-contact optical surfaces (mandatory).

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993, ISO 15004, CE MDR / FDA.

## 42.3 Standard Accessories

a) Universal diagnostic three-mirror gonio lens ring housing — Complete Goldmann three-mirror lens with all three mirrors and central element in single-body construction.

b) Anti-reflective optical glass coating layer mask — Applied on non-contact surfaces; supplied as factory-integrated coating.

c) Rigid plastic fitted padded protective storage box — Fitted foam/velvet interior; hinged lid with clasp; individual compartment.

## 42.4 Installation Requirements

None — clinic-portable. Used with slit-lamp biomicroscope (specified separately). Standard ophthalmic coupling fluid and reprocessing facilities required.

## 42.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate, Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 42.6 Training

Clinical technique training (minimum ½ day, combined with other slit-lamp diagnostic lens training).

## 42.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, mirror alignment, and coating). Spares availability minimum 10 years.

# 43. ABRAHAM YAG CAPSULOTOMY CONTACT LENS

Quantity: 1

## 43.1 Clinical & Functional Specifications

Abraham-type high-power (+66 D nominal) button contact lens for Nd:YAG laser posterior capsulotomy. It stabilises the eye, minimises corneal shift and beam wander, increases posterior capsule beam convergence for reduced energy at the capsule with lower risk to the intraocular lens optic, and provides a magnified surgeon view.

## 43.2 Engineering Technical Specifications

- Type: Abraham-style YAG capsulotomy contact lens (posterior-focus configuration).

- Central button power: Nominal +66 D (as per Abraham standard).

- Beam convergence effect: Increases posterior capsule irradiance / decreases energy requirement at capsule; reduces beam divergence at the intraocular lens optic.

- Contact diameter: Standard adult size (approximately 17—18 mm).

- Central button diameter: Approximately 8—10 mm.

- Magnification factor: Approximately 1.4× at capsule plane.

- Laser wavelength compatibility: Optimised for Nd:YAG (1064 nm) laser transmission.

- Optical material: Optical-grade PMMA or equivalent laser-transmissive optical polymer.

- Anti-reflective coating: Broadband AR coating suitable for Nd:YAG and slit-lamp visible-light transmission.

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Handling ring: Anodised aluminium or equivalent ergonomic finger-holding ring.

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993, ISO 15004, IEC 60601-2-22 (laser safety, as accessory), CE MDR / FDA.

## 43.3 Standard Accessories

a) Laser-focused +66 D button lens element with centre mount — Precision-moulded +66 D button; broadband AR coating; laser-transmissive.

b) Anodised aluminium ergonomic finger holding ring shaft — Corrosion-resistant handling ring; smooth polished edges; compatible with ophthalmic cleaning agents.

c) Padded custom-fitted protective containment case — Rigid case with fitted foam/velvet interior; hinged lid with clasp.

## 43.4 Installation Requirements

None — clinic/OR-portable. Used with Nd:YAG laser system (specified separately). Standard ophthalmic coupling fluid and reprocessing facilities required.

## 43.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate, Laser Transmission Certificate, Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 43.6 Training

Clinical technique training (minimum ½ day, combined with YAG laser training).

## 43.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, coating, and mechanical). Spares availability minimum 10 years.

# 44. PERIPHERAL IRIDOTOMY (ABRAHAM) CONTACT LENS

Quantity: 1

## 44.1 Clinical & Functional Specifications

Abraham-type high-power button contact lens with an offset (decentred) button optimised for laser peripheral iridotomy, providing focused laser energy at the peripheral iris plane, magnification of the target zone, stabilisation of the eye, and reduction of energy required at the iris surface. Applicable with Nd:YAG laser (photodisruption iridotomy) and, where the platform supports it, argon/532 nm laser (thermal iridotomy).

## 44.2 Engineering Technical Specifications

- Type: Abraham-style peripheral iridotomy contact lens with off-centre (decentred) button.

- Button power: Nominal +66 D (as per Abraham standard).

- Button position: Off-centre / decentred, aligned for peripheral-iris target zone.

- Contact diameter: Standard adult size (approximately 17—18 mm).

- Button diameter: Approximately 8—10 mm.

- Magnification factor: Approximately 1.4× at iris plane.

- Laser compatibility: Compatible with Nd:YAG (1064 nm) and argon/frequency-doubled Nd:YAG (532 nm) wavelengths (broadband transmission).

- Optical material: Optical-grade PMMA or equivalent laser-transmissive optical polymer.

- Anti-reflective coating: Broadband AR coating suitable for both Nd:YAG and visible/argon wavelengths.

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Handling ring: Knurled outer alignment ring for stable positioning and orientation.

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993, ISO 15004, IEC 60601-2-22 (as accessory), CE MDR / FDA.

## 44.3 Standard Accessories

a) Laser-focused +66 D magnification button element with offset mount — Decentred +66 D button; broadband AR coating; laser-transmissive.

b) Knurled stabilisation outer alignment ring casing — Textured knurled ring for stable orientation and rotational alignment.

c) Custom-fitted velvet-lined protective container box — Rigid case with fitted velvet interior; hinged lid with clasp.

## 44.4 Installation Requirements

None — clinic/OR-portable. Used with Nd:YAG or argon laser system (specified separately). Standard ophthalmic coupling fluid and reprocessing facilities required.

## 44.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate, Laser Transmission Certificate (Nd:YAG and 532 nm), Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 44.6 Training

Clinical technique training (minimum ½ day, combined with laser peripheral iridotomy training).

## 44.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, coating, and mechanical). Spares availability minimum 10 years.

# 45. PANFUNDOSCOPIC CONTACT LENS

Quantity: 1

## 45.1 Clinical & Functional Specifications

Wide-field panfundoscopic diagnostic and laser contact lens for slit-lamp examination and laser treatment of the posterior pole and mid-peripheral retina, providing a wide inverted image of the fundus. Suitable for panretinal photocoagulation, focal/grid macular laser (with appropriate technique), and general wide-field fundus diagnostic examination.

## 45.2 Engineering Technical Specifications

- Type: Wide-field panfundoscopic contact lens (inverted, real-image type).

- Static field of view: Minimum 120°; dynamic field of view with tilt minimum 140°.

- Image type: Inverted, real, wide-field.

- Laser spot magnification factor (retina): Typical approximately 1.4× (approximate; per lens design).

- Image magnification (surgeon view): Typical approximately 0.7× (as per wide-field inverted design).

- Contact diameter: Standard adult size (approximately 17—18 mm).

- Laser wavelength compatibility: Broadband — compatible with argon (488/514 nm), frequency-doubled Nd:YAG (532 nm), yellow (577 nm), red (670 nm), and Nd:YAG (1064 nm) where used diagnostically.

- Optical material: Optical-grade glass or laser-transmissive optical polymer.

- Anti-reflective coating: Broadband AR coating (multi-wavelength).

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Handling ring: Textured rubber stabilisation grip ring housing.

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993, ISO 15004, IEC 60601-2-22 (as laser accessory), CE MDR / FDA.

## 45.3 Standard Accessories

a) Wide-field retinal diagnostic imaging lens optics glass dome — Precision-moulded wide-field optical element; broadband AR coating; laser-transmissive.

b) Textured rubber stabilisation grip ring wrap housing — Non-slip textured grip; autoclave-incompatible components clearly identified; compatible with standard ophthalmic cleaning agents.

c) Padded rigid plastic protective transport shell case — Fitted padded interior; hinged lid with clasp; individual compartment.

## 45.4 Installation Requirements

None — clinic/OR-portable. Used with slit-lamp biomicroscope and/or photocoagulation laser system (specified separately). Standard ophthalmic coupling fluid and reprocessing facilities required.

## 45.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate (including field-of-view and spot-magnification specification), Laser Transmission Certificate (multi-wavelength), Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 45.6 Training

Clinical technique training (minimum ½ day, combined with photocoagulation laser training).

## 45.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, coating, and mechanical). Spares availability minimum 10 years.

# 46. MAINSTER RETINAL LASER LENS

Quantity: 1

## 46.1 Clinical & Functional Specifications

Mainster-type high-magnification focal retinal laser contact lens for precision focal and grid macular photocoagulation, panretinal photocoagulation, and posterior pole laser treatment. It provides a precise, minimally-magnified spot at the retina with a clear, distortion-free image of the macula and posterior pole.

## 46.2 Engineering Technical Specifications

- Type: Mainster-style focal/standard retinal laser contact lens (upright real-image type).

- Static field of view: Approximately 90° (standard/focal configuration); wide-field variants approximately 118°—125° where supplied as such.

- Image type: Upright, real, high-resolution (compared with wide-field panfundoscopic design).

- Laser spot magnification factor (retina): Typical approximately 1.05× (standard focal configuration; near-1:1 for precise focal work).

- Image magnification (surgeon view): Typical approximately 0.96× (standard focal configuration).

- Contact diameter: Standard adult size (approximately 17—18 mm).

- Laser wavelength compatibility: Broadband — compatible with argon (488/514 nm), frequency-doubled Nd:YAG (532 nm), yellow (577 nm), red (670 nm), and diode (810 nm) as applicable.

- Optical material: Optical-grade glass with laser-transmissive design; multi-element optical assembly.

- Anti-reflective coating: Broadband AR coating (multi-wavelength).

- Contact surface: Precision-polished, biocompatible per ISO 10993.

- Handling ring: Precision knurled holding ring frame assembly.

- Coupling fluid compatibility: Compatible with standard ophthalmic viscous coupling agents.

- Reprocessing: Manual cleaning and chemical disinfection per manufacturer's protocol.

- Compliance: ISO 10993, ISO 15004, IEC 60601-2-22 (as laser accessory), CE MDR / FDA.

## 46.3 Standard Accessories

a) High-magnification focal retinal photocoagulation lens glass — Precision multi-element optical assembly; broadband AR coating; laser-transmissive; upright real-image design.

b) Precision knurled holding ring frame assembly — Textured knurled ring for stable finger placement and rotational stability.

c) Custom-fitted padded protective utility storage case — Rigid case with fitted padded interior; hinged lid with clasp; individual compartment.

## 46.4 Installation Requirements

None — clinic/OR-portable. Used with photocoagulation laser system delivered through slit-lamp adapter (specified separately). Standard ophthalmic coupling fluid and reprocessing facilities required.

## 46.5 Documentation

User Manual (including cleaning and disinfection protocol), Factory Optical Certificate (including field-of-view and spot-magnification specification), Laser Transmission Certificate (multi-wavelength), Biocompatibility Certificate (ISO 10993), Warranty Certificate.

## 46.6 Training

Clinical technique training (minimum ½ day, combined with focal/grid macular and panretinal photocoagulation laser training).

## 46.7 Warranty

Comprehensive minimum two (2) years covering manufacturing defects (optical, coating, and mechanical). Spares availability minimum 10 years.