TECHNICAL SPECIFICATIONS — PAEDIATRIC DEPARTMENT EQUIPMENT

Technical Tender Document

Table of Contents

1. PAEDIATRIC MULTIPARA PATIENT MONITOR	7

1.1 Clinical & Functional Specifications	7

1.2 Engineering Technical Specifications	7

1.3 Standard Accessories	8

1.4 Installation Requirements	9

1.5 Documentation	9

1.6 Training	9

1.7 Warranty	9

2. PAEDIATRIC ECG MACHINE	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	11

2.4 Installation Requirements	11

2.5 Documentation	11

2.6 Training	11

2.7 Warranty	12

3. ULTRASONIC NEBULIZER (PAEDIATRIC)	13

3.1 Clinical & Functional Specifications	13

3.2 Engineering Technical Specifications	13

3.3 Standard Accessories	13

3.4 Installation Requirements	14

3.5 Documentation	14

3.6 Training	14

3.7 Warranty	14

4. PAEDIATRIC BP APPARATUS	15

4.1 Clinical & Functional Specifications	15

4.2 Engineering Technical Specifications	15

4.3 Standard Accessories	15

4.4 Installation Requirements	16

4.5 Documentation	16

4.6 Training	16

4.7 Warranty	16

5. PAEDIATRIC OTOSCOPE	17

5.1 Clinical & Functional Specifications	17

5.2 Engineering Technical Specifications	17

5.3 Standard Accessories	17

5.4 Installation Requirements	18

5.5 Documentation	18

5.6 Training	18

5.7 Warranty	18

6. PAEDIATRIC PULSE OXIMETER	19

6.1 Clinical & Functional Specifications	19

6.2 Engineering Technical Specifications	19

6.3 Standard Accessories	19

6.4 Installation Requirements	20

6.5 Documentation	20

6.6 Training	20

6.7 Warranty	20

7. INFANT WEIGHING SCALE	21

7.1 Clinical & Functional Specifications	21

7.2 Engineering Technical Specifications	21

7.3 Standard Accessories	21

7.4 Installation Requirements	22

7.5 Documentation	22

7.6 Training	22

7.7 Warranty	22

8. PAEDIATRIC STANDING WEIGHING SCALE	23

8.1 Clinical & Functional Specifications	23

8.2 Engineering Technical Specifications	23

8.3 Standard Accessories	23

8.4 Installation Requirements	23

8.5 Documentation	24

8.6 Training	24

8.7 Warranty	24

9. LENGTH BOARD / INFANT METER	25

9.1 Clinical & Functional Specifications	25

9.2 Engineering Technical Specifications	25

9.3 Standard Accessories	25

9.4 Installation Requirements	25

9.5 Documentation	25

9.6 Training	25

9.7 Warranty	26

10. PAEDIATRIC NASAL SPECULUM	27

10.1 Clinical & Functional Specifications	27

10.2 Engineering Technical Specifications	27

10.3 Standard Accessories	27

10.4 Installation Requirements	27

10.5 Documentation	27

10.6 Training	27

10.7 Warranty	27

11. VACCINE COLD-CHAIN REFRIGERATOR	29

11.1 Clinical & Functional Specifications	29

11.2 Engineering Technical Specifications	29

11.3 Standard Accessories	29

11.4 Installation Requirements	30

11.5 Documentation	30

11.6 Training	30

11.7 Warranty	30

12. PAEDIATRIC RESUSCITATION TROLLEY	31

12.1 Clinical & Functional Specifications	31

12.2 Engineering Technical Specifications	31

12.3 Standard Accessories	31

12.4 Installation Requirements	32

12.5 Documentation	32

12.6 Training	32

12.7 Warranty	32

13. PAEDIATRIC EXAMINATION TABLE	33

13.1 Clinical & Functional Specifications	33

13.2 Engineering Technical Specifications	33

13.3 Standard Accessories	33

13.4 Installation Requirements	34

13.5 Documentation	34

13.6 Training	34

13.7 Warranty	34

14. PAEDIATRIC STETHOSCOPE	35

14.1 Clinical & Functional Specifications	35

14.2 Engineering Technical Specifications	35

14.3 Standard Accessories	35

14.4 Installation Requirements	35

14.5 Documentation	35

14.6 Training	35

14.7 Warranty	36

15. PAEDIATRIC GLUCOMETER	37

15.1 Clinical & Functional Specifications	37

15.2 Engineering Technical Specifications	37

15.3 Standard Accessories	37

15.4 Installation Requirements	38

15.5 Documentation	38

15.6 Training	38

15.7 Warranty	38

16. NEONATAL BABY WARMER	39

16.1 Clinical & Functional Specifications	39

16.2 Engineering Technical Specifications	39

16.3 Standard Accessories	40

16.4 Installation Requirements	40

16.5 Documentation	40

16.6 Training	40

16.7 Warranty	40

17. PAEDIATRIC INFUSION PUMP	41

17.1 Clinical & Functional Specifications	41

17.2 Engineering Technical Specifications	41

17.3 Standard Accessories	42

17.4 Installation Requirements	42

17.5 Documentation	42

17.6 Training	42

17.7 Warranty	42

18. NEONATAL INCUBATOR	44

18.1 Clinical & Functional Specifications	44

18.2 Engineering Technical Specifications	44

18.3 Standard Accessories	45

18.4 Installation Requirements	45

18.5 Documentation	45

18.6 Training	46

18.7 Warranty	46

19. LED PHOTOTHERAPY UNIT	47

19.1 Clinical & Functional Specifications	47

19.2 Engineering Technical Specifications	47

19.3 Standard Accessories	47

19.4 Installation Requirements	48

19.5 Documentation	48

19.6 Training	48

19.7 Warranty	48

20. BUBBLE CPAP	49

20.1 Clinical & Functional Specifications	49

20.2 Engineering Technical Specifications	49

20.3 Standard Accessories	49

20.4 Installation Requirements	50

20.5 Documentation	50

20.6 Training	50

20.7 Warranty	50

21. VENTILATOR WITH PAEDIATRIC & NEONATAL MODES	51

21.1 Clinical & Functional Specifications	51

21.2 Engineering Technical Specifications	51

21.3 Standard Accessories	52

21.4 Installation Requirements	53

21.5 Documentation	53

21.6 Training	53

21.7 Warranty	53

# 1. PAEDIATRIC MULTIPARA PATIENT MONITOR

Quantity: 10

## 1.1 Clinical & Functional Specifications

The equipment shall be a microprocessor-controlled, high-resolution multi-parameter bedside patient monitor specifically configured for continuous physiological monitoring of paediatric and neonatal patients in general wards, high-dependency units, paediatric ICU, and step-down care areas. It shall provide simultaneous real-time acquisition, display, trending, storage, and alarming of vital physiological parameters including ECG (3/5-lead), Heart Rate, Respiration Rate (via impedance pneumography), Non-Invasive Blood Pressure (NIBP), Functional Oxygen Saturation (SpO₂), Pulse Rate, and dual-channel Temperature. The monitor shall automatically recognize patient category (Adult / Paediatric / Neonatal) with corresponding alarm limit presets, measurement algorithms, and cuff inflation profiles. It shall support silent bedside operation with low-noise fanless or ultra-quiet cooling design, be suitable for use with electrosurgical units, and provide arrhythmia analysis, ST-segment analysis, and pacemaker detection.

## 1.2 Engineering Technical Specifications

Display: Color TFT LCD or equivalent flat-panel display; minimum 10-inch diagonal active viewing area; capacitive or resistive touchscreen; minimum resolution 1024 × 600 pixels or better; capable of displaying at least 6 simultaneous waveforms with numerical parameters; adjustable brightness; viewing angle not less than 160° horizontal and vertical.

Processor & Memory: Medical-grade embedded processor with sufficient computing power for real-time multi-parameter processing; minimum 72 hours of full-disclosure trend storage for all parameters; minimum 1000 NIBP measurement records and 100 alarm event records with waveform snapshots.

ECG

- Leads: 3-lead and 5-lead selectable

- Bandwidth: minimum 0.05—150 Hz (Diagnostic mode); Monitor mode 0.5—40 Hz selectable

- Sampling rate: minimum 500 samples/second/channel (1000 samples/sec preferred)

- Heart Rate range: 15—300 bpm (Adult/Paed), 15—350 bpm (Neonatal)

- Accuracy: ±1 bpm or ±1%, whichever is greater

- Input impedance: ≥ 5 MΩ; CMRR ≥ 90 dB

- Defibrillator-proof input protection; ESU protected

- Arrhythmia and ST-segment analysis included

Respiration (Impedance): Range 0—150 rpm; accuracy ±2 rpm; adjustable apnoea alarm delay 10—40 seconds.

NIBP

- Method: Oscillometric

- Modes: Manual, Automatic (interval selectable 1—480 min), STAT, Continuous

- Pressure range: Neonatal 10—135 mmHg; Paediatric 10—200 mmHg; Adult 10—270 mmHg

- Accuracy: within ±3 mmHg mean error

- Overpressure protection with independent hardware safety circuit

- Automatic cuff-size recognition or selectable patient mode

SpO₂

- Range: 0—100%; accuracy ±2% (70—100%)

- Motion-tolerant and low-perfusion capable algorithm

- Pulse rate range 25—250 bpm; accuracy ±2 bpm

- Perfusion Index display

Temperature: Dual channel; range 0—50 °C; accuracy ±0.1 °C; YSI-400 or equivalent compatible probes.

Alarms: Three-level priority (High/Medium/Low); visual and audible; adjustable limits per parameter and patient category; alarm silence, pause, and audio-off with visual indication; alarm history log retrievable.

Connectivity: Ethernet (RJ-45) for central monitoring; Wi-Fi capability (optional/configurable); minimum 2 × USB ports; nurse-call relay output; VGA/HDMI output; HL7 compatibility for HIS/EMR integration.

Power

- Input: 100—240 VAC, 50/60 Hz, auto-ranging

- Integrated rechargeable Li-ion battery; runtime minimum 2 hours continuous operation on full charge; recharge time maximum 6 hours

- Battery status indicator on display

Environmental: Operating 5—40 °C, 15—95% RH non-condensing; IP rating minimum IPX1 (drip-proof).

Physical: Compact bedside configuration; weight approximately 4—7 kg; suitable for pole, wall-rail, shelf, or trolley mounting.

Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-1-8 (Alarms), IEC 60601-2-27 (ECG), IEC 60601-2-30 (NIBP), IEC 60601-2-49 (Multifunction), ISO 80601-2-61 (SpO₂), ISO 13485, CE MDR / US FDA cleared.

## 1.3 Standard Accessories

3/5-lead ECG patient cable and reusable paediatric/neonatal lead wire set: Medical-grade shielded, defibrillator-protected trunk cable with AHA or IEC colour-coded lead wires; snap or grabber-style connectors; latex-free; total length approximately 2.5—3.5 m; reusable and cleanable with hospital-grade disinfectants.

Reusable paediatric SpO₂ finger sensor with interface extension cable: Optical dual-wavelength (Red/IR) reusable sensor sized for paediatric fingers; biocompatible medical-grade housing; cable length minimum 90 cm; extension cable minimum 2 m; compatible with monitor SpO₂ module; cleanable with alcohol-based wipes.

Reusable infant SpO₂ wrap sensor with soft adhesive strips: Y-shaped flexible wrap sensor for neonatal/infant foot/hand application; soft biocompatible material; replaceable disposable adhesive wraps included (starter set); cable length minimum 90 cm.

Reusable paediatric NIBP cuffs (small, medium, large paediatric) with air hose: Nylon or reinforced fabric bladder cuffs; latex-free; hook-and-loop closure; size markings and range indicators printed; twin-tube or single-tube configuration matching monitor; air hose length minimum 3 m with locking connector; cleanable and reusable.

Reusable skin temperature probe (surface/rectal compatible): YSI-400 series or equivalent thermistor probe; biocompatible medical-grade PVC-coated; cable length minimum 3 m; accuracy ±0.1 °C; usable for skin surface and rectal measurement.

Integrated rechargeable battery pack: Medical-grade sealed Li-ion battery pack; capacity sufficient for minimum 2 hours continuous monitoring; user-accessible or service-replaceable; overcharge/deep-discharge protection.

Bedside rail mounting bracket assembly with cable management hooks: Powder-coated metal or high-impact ABS; universal bed-rail clamp; adjustable tilt; integrated cable management hooks; load rating suitable for monitor weight.

## 1.4 Installation Requirements

Standard hospital wall socket, 100—240 VAC, 50/60 Hz, 3-pin earthed medical outlet; UPS-backed circuit recommended; ambient temperature 18—28 °C; humidity ≤ 80% RH; network port (RJ-45) at bedside for central monitoring connectivity; no special HVAC or shielding required; no medical gas connections required.

## 1.5 Documentation

User Manual (English), Service Manual, Illustrated Parts Catalogue, Factory Test & Calibration Certificate, Electrical Safety Test Report (IEC 62353), Preventive Maintenance Schedule, Installation & Commissioning Report, CE/FDA regulatory certificates.

## 1.6 Training

Comprehensive on-site clinical training for nursing and paediatric medical staff; biomedical engineering training including troubleshooting, module replacement, calibration verification; application training for advanced features (arrhythmia, ST-analysis); refresher session at 6 months.

## 1.7 Warranty

Minimum 2 years comprehensive warranty covering all labour, spare parts, preventive maintenance visits (2 per year minimum), breakdown maintenance with response time not exceeding 48 hours, software updates and patches, and 24/7 technical support.

# 2. PAEDIATRIC ECG MACHINE

Quantity: 2

## 2.1 Clinical & Functional Specifications

Portable microprocessor-based 12-lead diagnostic electrocardiograph optimized for paediatric and neonatal patients, providing simultaneous acquisition, display, storage, interpretation, and printing of 12-lead ECG traces. The equipment shall support automatic paediatric-specific interpretation algorithms with age- and sex-adjusted normal ranges, arrhythmia detection, and computerized measurement of ECG parameters (PR, QRS, QT, QTc). Modes shall include manual, automatic, rhythm strip, and continuous monitoring. Suitable for outpatient, inpatient, emergency, and bedside use.

## 2.2 Engineering Technical Specifications

Display: Color TFT LCD, minimum 7-inch diagonal; touchscreen preferred; resolution minimum 800 × 480 pixels; simultaneous display of at least 3, 6, or 12 leads selectable.

ECG Acquisition

- Leads: Standard 12-lead simultaneous acquisition

- Sampling rate: minimum 1000 samples/second/channel (2000 Hz or better preferred for paediatric high-frequency content)

- A/D conversion: minimum 16-bit resolution

- Frequency response: 0.05—150 Hz minimum

- Digital filters: AC (50/60 Hz notch), baseline drift, muscle artefact (EMG) — selectable

- Input impedance: ≥ 50 MΩ

- CMRR: ≥ 100 dB

- Dynamic range: ±5 mV minimum

- DC offset tolerance: ±300 mV or better

- Defibrillator protection: withstand 5 kV / 360 J

- Pacemaker detection: automatic

- Heart rate range: 15—350 bpm; accuracy ±1 bpm

Interpretation: On-board automatic ECG measurement and interpretation software with paediatric-specific criteria (age-based normal values from neonate through adolescent); serial comparison capability preferred.

Printer: Integrated high-resolution thermal array printer; print width minimum 210 mm (A4) or equivalent; print resolution minimum 200 dpi vertical × 500 dpi horizontal (or 8 dots/mm × 40 dots/mm); paper speed selectable 5, 10, 25, 50 mm/s; formats 3/6/12 channel; auto and manual mode.

Storage: Internal memory for minimum 200 ECG records; USB export; DICOM/PDF/XML/SCP-ECG export formats.

Connectivity: USB, Ethernet (RJ-45), Wi-Fi optional; HL7 and DICOM-ECG compatible for HIS/EMR integration.

Power: 100—240 VAC, 50/60 Hz; integrated rechargeable Li-ion battery; minimum 2 hours continuous operation or 50 ECG reports per full charge; recharge time maximum 6 hours.

Environmental: Operating 10—40 °C, 25—95% RH non-condensing.

Physical: Portable tabletop/trolley; weight approximately 3—6 kg.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-25 (Diagnostic ECG), IEC 60601-2-51, ISO 13485, CE MDR / FDA cleared.

## 2.3 Standard Accessories

10-lead patient ECG trunk cable with paediatric-optimized banana plugs: Shielded medical-grade multi-conductor cable, AHA/IEC colour-coded, length approximately 3—4 m, defibrillator-protected, latex-free, with 4 mm banana-style paediatric lead connectors; strain relief at both ends.

Reusable paediatric chest suction electrodes (set of 6, small diameter bulb): Small-diameter suction bulb electrodes optimized for paediatric chest; chrome-plated brass or silver/silver-chloride contact; latex-free bulb; reusable and autoclavable/chemically disinfectable.

Reusable paediatric limb extremity clamp electrodes (set of 4, small profile): Small-profile spring clamp electrodes with silver/silver-chloride or nickel-plated contact surface; adjustable tension; colour-coded; reusable.

Integrated high-resolution thermal matrix printer module: Built-in high-density thermal print head; resolution minimum 8 dots/mm × 40 dots/mm; paper-out and paper-jam sensors.

Thermal Z-fold chart paper (2 packs): Medical-grade high-contrast thermal ECG paper; A4 or equivalent width; Z-fold pack of approximately 200 sheets each; standard ECG grid preprinted.

Built-in rechargeable lithium-ion battery pack with medical-grade power cord: Sealed medical-grade Li-ion pack; minimum 2 hours runtime; medical-grade IEC 60601-compliant 3-pin power cord.

## 2.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz earthed outlet; UPS-backed circuit recommended; ambient temperature 18—28 °C; no special HVAC; tabletop or trolley placement.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate, Electrical Safety Test Report per IEC 62353, PM Schedule, Installation Report, CE/FDA certificate.

## 2.6 Training

Clinical training for paediatricians, cardiologists and nursing staff on device operation and interpretation; biomedical engineering training on maintenance, calibration verification, and troubleshooting.

## 2.7 Warranty

Minimum 2 years comprehensive covering labour, spare parts, PM, breakdown maintenance, software updates, and technical support.

# 3. ULTRASONIC NEBULIZER (PAEDIATRIC)

Quantity: 6

## 3.1 Clinical & Functional Specifications

Ultrasonic nebulizer for paediatric inhalation therapy that converts liquid medication into fine respirable aerosol using high-frequency piezoelectric transducer, delivering consistent particle size suitable for lower-airway deposition. The unit shall provide adjustable nebulization output and airflow to accommodate infants through adolescents with respiratory conditions including asthma, bronchiolitis, and cystic fibrosis. It shall operate silently, feature integrated timer control, and permit continuous or intermittent therapy.

## 3.2 Engineering Technical Specifications

Transducer: Piezoelectric ultrasonic; operating frequency minimum 1.5 MHz (typical 1.7—2.4 MHz).

Aerosol Performance

- Particle size (MMAD): within 1—5 µm respirable range; typical 3—5 µm

- Nebulization output: adjustable 0.5—3.0 mL/min or equivalent range

- Airflow: adjustable, typical 5—20 L/min

- Medication chamber capacity: minimum 10 mL

Control: Digital display; adjustable output, airflow, and timer (0—60 min selectable); auto shut-off at end of medication or timer expiry; low-water safety cutoff.

Noise Level: ≤ 45 dB(A) at 1 m.

Power: 100—240 VAC, 50/60 Hz; power consumption approximately 50—100 W.

Environmental: Operating 5—40 °C, 15—90% RH; IPX1 minimum.

Physical: Portable tabletop or mobile stand; weight approximately 3—6 kg.

Standards: IEC 60601-1, IEC 60601-1-2, ISO 27427 (Aerosol drug delivery devices), ISO 13485, CE MDR.

## 3.3 Standard Accessories

High-frequency ultrasonic transducer crystal module (pre-installed): Factory-installed piezoelectric crystal assembly with cooling water bath separation; user-replaceable service part.

Reusable aerosol medication chambers (set of 2, autoclavable): Medical-grade polycarbonate or polysulfone chambers; capacity minimum 10 mL; compatible with steam autoclave up to 134 °C or approved chemical disinfection.

Reusable paediatric aerosol face masks with soft elastic straps (2 pieces): Soft PVC or silicone; anatomically contoured for paediatric facial profile; latex-free; adjustable elastic head strap; cleanable and reusable.

Corrugated flexible aerosol delivery tubes (various lengths, 2 pieces): Medical-grade flexible corrugated tubing; lengths approximately 1 m and 1.8 m; kink-resistant; low-resistance; latex-free; cleanable.

Integrated digital airflow and timer control panel screen: Backlit LCD or LED display showing output level, airflow rate, remaining time, and error/service indicators.

Mobile stand mount bracket or tabletop rubber stabilization base pads: Powder-coated metal mobile stand with lockable castors OR anti-slip rubber base pads for tabletop use; user selectable per site requirement.

## 3.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz earthed outlet; potable-quality water for coupling reservoir; ambient temperature 18—30 °C; no medical gas required.

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report, PM Schedule, CE certificate.

## 3.6 Training

Clinical training for nursing/respiratory therapy staff; biomedical training for cleaning, disinfection, transducer replacement, and troubleshooting.

## 3.7 Warranty

Minimum 2 years comprehensive covering labour, spare parts (excluding consumables), PM, and technical support.

# 4. PAEDIATRIC BP APPARATUS

Quantity: 3

## 4.1 Clinical & Functional Specifications

Aneroid sphygmomanometer with heavy-duty gauge, calibrated inflation system, and paediatric-sized cuffs for accurate manual measurement of blood pressure in infants, toddlers, and older children. The device shall be robust, portable or wall-mounted, and shall provide reliable long-term calibration stability with easy field verification.

## 4.2 Engineering Technical Specifications

Manometer

- Type: Aneroid (mechanical), high-visibility dial

- Dial diameter: minimum 140 mm

- Measurement range: 0—300 mmHg

- Resolution: 2 mmHg graduations

- Accuracy: within ±3 mmHg over full range at ambient temperature

- Zero-set indicator (pin or arc)

- Shock-resistant movement, protected against overpressure to at least 600 mmHg

Inflation System: Heavy-duty latex-free rubber bulb with air-release valve capable of controlled fine deflation (2—3 mmHg/sec) and rapid full exhaust.

Mounting: Desk stand, wall-mount, or mobile stand configuration.

Standards: IEC 60601-1, ISO 81060-1 (Non-automated sphygmomanometers), EN 1060 series, ISO 13485, CE MDR.

## 4.3 Standard Accessories

Heavy-duty desk-top or wall-mount aneroid manometer gauge with high-visibility dial: Dial ≥ 140 mm; large-print numerals; anti-glare face; shock-resistant housing; mounting hardware included.

Reusable infant nylon blood pressure cuff with bladder: Latex-free reinforced nylon cuff sized for infants (arm circumference approximately 8—13 cm); double-tube bladder; hook-and-loop closure; range markings printed; cleanable.

Reusable toddler nylon blood pressure cuff with bladder: Toddler size (arm circumference approximately 12—19 cm); construction as above.

Reusable child nylon blood pressure cuff with bladder: Child size (arm circumference approximately 18—26 cm); construction as above.

Heavy-duty latex-free rubber inflation bulb with air release thumb screw valve: Chrome-plated brass thumb-screw fine control valve; smooth-action deflation; check valve for one-way inflation.

Coiled flexible twin-core air connecting hoses with quick-click metal couplers: Latex-free coiled or straight hoses with chrome-plated metal luer or bayonet-style connectors; length minimum 60 cm extended.

## 4.4 Installation Requirements

No electrical or utility requirements; wall-mount version requires standard wall fixing hardware.

## 4.5 Documentation

User Manual, Factory Calibration Certificate, Cleaning/Disinfection Guide, CE certificate.

## 4.6 Training

Brief clinical orientation on cuff sizing, positioning, and Korotkoff auscultation technique.

## 4.7 Warranty

Minimum 2 years covering manometer movement, cuffs, bulb, and hoses under normal clinical use.

# 5. PAEDIATRIC OTOSCOPE

Quantity: 2

## 5.1 Clinical & Functional Specifications

High-intensity fibre-optic LED otoscope for examination of the external auditory canal and tympanic membrane in paediatric patients, with integrated pneumatic port for mobility testing of the tympanic membrane. Suitable for outpatient, emergency, and inpatient use.

## 5.2 Engineering Technical Specifications

Illumination

- Source: White LED, minimum 30,000 lux at distal tip

- Colour temperature: 4000—6000 K (near-daylight)

- LED life: minimum 20,000 hours

- Fibre-optic light delivery ring (shadow-free, obstruction-free view)

Optics

- Magnification: minimum 3× (typical 3—5×)

- Wide-angle sealed viewing lens with swing-away or removable design for instrumentation

- Anti-fog / anti-reflection coating preferred

Pneumatic Port: Sealed insufflation port with bulb (bulb may be supplied separately) for tympanic membrane mobility testing.

Handle: Rechargeable lithium-ion or NiMH; runtime minimum 2 hours continuous illumination per full charge; charge time maximum 6 hours; textured non-slip grip.

Charger: Desktop charging dock, 100—240 VAC, 50/60 Hz.

Environmental: Operating 10—40 °C.

Standards: IEC 60601-1, ISO 10993 (biocompatibility of specula), ISO 13485, CE MDR.

## 5.3 Standard Accessories

Otoscope optical head with high-intensity LED fibre-optic illumination path: Sealed viewing head with LED source, fibre-optic light guide, magnifying lens, and pneumatic port; twist-lock connection to handle.

Integrated pneumatoscope port: Sealed insufflation nipple accepting standard rubber tubing and bulb for tympanic mobility testing.

Reusable paediatric polypropylene specula (set of 4, diameters 2.5 mm and 3 mm): Autoclavable or chemically disinfectable polypropylene specula in paediatric diameters; smooth atraumatic tip; twist-lock or bayonet fit.

Rechargeable lithium-ion battery handle: Sealed medical-grade Li-ion; runtime minimum 2 hours; knurled non-slip grip; low-battery indicator.

Desktop intelligent charging dock station with main power adapter cord: Auto-detect charging with overcharge protection; charge status LED indicators; 100—240 VAC medical-grade adapter.

Rigid plastic form-fitted protective storage containment case: High-impact plastic case with foam cutouts for otoscope head, handle, specula, and accessories.

## 5.4 Installation Requirements

Standard 100—240 VAC outlet for charging dock; no other utilities.

## 5.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, CE certificate.

## 5.6 Training

Brief clinical familiarization for paediatric and ENT staff; cleaning and disinfection training.

## 5.7 Warranty

Minimum 2 years comprehensive covering optical head, handle, battery, and charger.

# 6. PAEDIATRIC PULSE OXIMETER

Quantity: 10

## 6.1 Clinical & Functional Specifications

Portable handheld / tabletop pulse oximeter for continuous or spot-check non-invasive measurement of functional oxygen saturation (SpO₂) and pulse rate in paediatric and neonatal patients. The device shall feature motion-tolerant, low-perfusion capable measurement algorithm and provide clear numeric and plethysmographic display, adjustable alarms, and internal memory for trending.

## 6.2 Engineering Technical Specifications

Display: Color OLED or TFT LCD; minimum 2.4-inch diagonal; simultaneous display of SpO₂, pulse rate, plethysmographic waveform, pulse-strength bar, and battery status.

SpO₂

- Range: 0—100%

- Accuracy: ±2% (70—100%), unspecified below 70%

- Motion-tolerant and low-perfusion algorithm; PI display 0.02—20%

Pulse Rate: 25—250 bpm; accuracy ±2 bpm.

Alarms: High/low SpO₂ and pulse rate limits adjustable; visual and audible; alarm silence.

Memory: Internal storage for minimum 72 hours of trend data; USB export.

Power: Rechargeable Li-ion battery; runtime minimum 8 hours continuous; charge time maximum 4 hours; 100—240 VAC charging cradle.

Environmental: Operating 5—40 °C, 15—95% RH; IPX2 minimum.

Physical: Handheld or compact tabletop; weight approximately 300—800 g.

Standards: IEC 60601-1, IEC 60601-1-2, ISO 80601-2-61, ISO 13485, CE MDR / FDA cleared.

## 6.3 Standard Accessories

Reusable paediatric clip-type SpO₂ finger probe: Spring-loaded finger clip sized for paediatric fingers; dual-wavelength LEDs; biocompatible housing; latex-free; cable length minimum 90 cm.

Reusable neonatal/infant flexible Y-style wrap sensor with soft wraps: Y-shaped flexible sensor for foot/hand application; soft foam or silicone wraps included as starter set; cable length minimum 90 cm.

Protective rubber boot cover with integrated table kickstand: Impact-absorbing silicone/rubber sleeve encasing device housing; integrated flip-out kickstand for tabletop viewing angle.

Rechargeable battery pack set with desktop charging cradle station: Medical-grade Li-ion battery pack; desktop cradle with 100—240 VAC medical adapter; charge status LEDs.

Carrying lanyard and soft fabric transport pouch: Adjustable safety-release lanyard; padded fabric pouch with belt loop.

## 6.4 Installation Requirements

Standard 100—240 VAC outlet for charging; no other utilities.

## 6.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate, Electrical Safety Test Report, CE/FDA certificate.

## 6.6 Training

Clinical training for nursing staff on probe application (particularly on neonates), reading interpretation, and alarm management; biomedical training on cleaning and battery replacement.

## 6.7 Warranty

Minimum 2 years comprehensive covering unit, probes, battery, and charger.

# 7. INFANT WEIGHING SCALE

Quantity: 4

## 7.1 Clinical & Functional Specifications

Electronic digital infant weighing scale with contoured tray for accurate and safe weighing of neonates and infants, incorporating motion-compensating (tare/hold) function to obtain stable readings on moving infants. Suitable for maternity, paediatric wards, NICU, and outpatient clinics.

## 7.2 Engineering Technical Specifications

Weighing Range: 0—20 kg minimum (typical 0—20 kg or 0—25 kg).

Resolution: ≤ 5 g up to 10 kg; ≤ 10 g above 10 kg.

Accuracy: Class III per OIML R76 or equivalent; typical ±10 g.

Load Cell: Strain-gauge, temperature-compensated.

Functions: Auto-zero, tare, hold/motion-averaging for moving infants, unit conversion (kg/lb), auto power-off.

Display: Backlit LCD, digit height minimum 20 mm.

Height Measurement: Integrated mechanical measuring tape or digital slider rod, range 0—80 cm minimum, resolution 1 mm.

Power: 100—240 VAC medical adapter + internal rechargeable battery, runtime minimum 30 hours or 500 weighings per charge.

Physical: Tabletop; tray dimensions suitable for infants up to approximately 60 cm length; anti-slip base.

Standards: OIML R76 Class III, IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR.

## 7.3 Standard Accessories

Ergonomic contoured acrylic baby tray bed with high perimeter safety walls: Impact-resistant transparent or white acrylic tray with anatomically contoured surface; raised perimeter walls to prevent roll-off; smooth cleanable surface.

Integrated mechanical height-measuring tape or digital slider rod rule: Retractable measuring tape or sliding rod; range 0—80 cm minimum; resolution 1 mm; lockable position.

High-visibility backlit digital LCD indicator console screen: LCD with backlight; digits ≥ 20 mm; simultaneous display of weight, unit, and function indicators.

Automated electronic tare/hold function software for moving infants: Digital signal-processing motion-averaging algorithm providing stable reading on active infants; automatic activation.

AC power supply adapter and rechargeable internal backup battery pack: Medical-grade 100—240 VAC adapter; sealed rechargeable battery for at least 30 hours or 500 weighings per charge.

## 7.4 Installation Requirements

Flat, stable, vibration-free surface; standard 100—240 VAC outlet; ambient temperature 10—40 °C.

## 7.5 Documentation

User Manual, Service Manual, Factory Calibration Certificate traceable to national standards, Verification Certificate per OIML, CE certificate.

## 7.6 Training

Brief clinical orientation on positioning, tare/hold usage, and calibration verification.

## 7.7 Warranty

Minimum 2 years comprehensive covering load cell, electronics, battery, and adapter.

# 8. PAEDIATRIC STANDING WEIGHING SCALE

Quantity: 2

## 8.1 Clinical & Functional Specifications

Electronic digital standing weighing scale with integrated stadiometer for accurate weight and height measurement of ambulatory paediatric patients, suitable for outpatient clinics, wards, and school health screening.

## 8.2 Engineering Technical Specifications

Weighing Range: 0—150 kg minimum.

Resolution: ≤ 50 g up to 100 kg; ≤ 100 g above.

Accuracy: Class III per OIML R76; typical ±100 g.

Height Measurement: Telescopic mechanical rod, range 60—200 cm minimum; resolution 1 mm.

Functions: Auto-zero, tare, unit conversion, BMI calculation preferred.

Display: Backlit LCD at eye-level; digits ≥ 20 mm.

Power: 100—240 VAC medical adapter with rechargeable/backup battery.

Physical: Low-profile platform with anti-slip mat; platform load rating minimum 200 kg.

Standards: OIML R76 Class III, IEC 60601-1, ISO 13485, CE MDR.

## 8.3 Standard Accessories

Low-profile heavy-duty platform base with kid-friendly non-slip rubber mat: Powder-coated steel or reinforced plastic platform; height ≤ 60 mm above floor; textured non-slip rubber mat with child-friendly design.

Telescopic mechanical height-measuring rod column: Anodized aluminium or powder-coated steel rod; range 60—200 cm; graduated metric scale with 1 mm resolution; locking headpiece.

High-visibility digital LCD indicator console screen mounted at eye level: Backlit LCD with digit height ≥ 20 mm; mounted at approximately 130—150 cm above floor.

Automated electronic tare calibration function software parameters: Digital tare, auto-zero on startup, and internal calibration verification routine.

AC power supply adapter charger with medical-grade electrical power cord: IEC 60601-compliant 100—240 VAC medical adapter with 3-pin earthed cord.

## 8.4 Installation Requirements

Level, hard, vibration-free floor; standard 100—240 VAC outlet; adequate ceiling height for full-extension stadiometer (approximately 210 cm).

## 8.5 Documentation

User Manual, Service Manual, Factory Calibration Certificate, Verification per OIML, CE certificate.

## 8.6 Training

Clinical orientation on positioning, height measurement technique, and calibration verification.

## 8.7 Warranty

Minimum 2 years comprehensive.

# 9. LENGTH BOARD / INFANT METER

Quantity: 1

## 9.1 Clinical & Functional Specifications

Mechanical infant length board (infantometer) for accurate supine measurement of recumbent length in neonates and infants, with fixed headboard and sliding footboard providing reproducible measurements.

## 9.2 Engineering Technical Specifications

Measurement Range: 0—100 cm minimum.

Resolution: 1 mm.

Accuracy: ±2 mm.

Construction: Rigid non-warping high-density polymer or finished hardwood base.

Headboard: Fixed perpendicular panel with padded contact surface.

Footboard: Smooth-sliding carriage on precision rails with dual locking mechanism.

Scale: Engraved or laser-etched high-contrast metric scale, resistant to fading.

Physical: Portable; weight approximately 3—6 kg.

Standards: ISO 13485, CE MDR.

## 9.3 Standard Accessories

Rigid non-warping plastic or finished wooden base platform structure: Dimensionally stable base; smooth cleanable surface; non-warping under normal humidity variation.

Fixed vertical head-stop panel plate blocker: Perpendicular headboard rigidly fixed to base; padded contact face; corner-reinforced.

Smooth-sliding foot-paddle positioner carriage with dual locking screws: Precision-fit sliding carriage; two independent locking screws or lever locks; smooth low-friction glide.

Engraved high-contrast metric measuring rule scale (millimeter increments): Laser-etched or engraved permanent scale in millimetres; contrasting colour for readability.

## 9.4 Installation Requirements

Flat clean surface; no utilities required.

## 9.5 Documentation

User Manual, Calibration/Verification Certificate, CE certificate.

## 9.6 Training

Brief clinical orientation on proper infant positioning technique.

## 9.7 Warranty

Minimum 2 years covering structural defects.

# 10. PAEDIATRIC NASAL SPECULUM

Quantity: 2

## 10.1 Clinical & Functional Specifications

Precision paediatric nasal speculum instruments for examination of the nasal cavity in children, permitting atraumatic dilation and visualization. Both Hartmann and Vienna patterns are required in paediatric sizing.

## 10.2 Engineering Technical Specifications

Material: Surgical-grade stainless steel, AISI 420 or 304 or equivalent; passivated finish; corrosion-resistant.

Finish: Matte or satin non-reflective finish preferred to reduce glare.

Sterilization: Fully autoclavable at 134 °C; compatible with chemical disinfection.

Construction: Precision-hinged sprung handles; smooth atraumatic blade tips; ratchet or spring return mechanism.

Blades: Short paediatric-sized blades appropriate to age group.

Standards: ISO 7153-1 (Surgical instrument materials), ISO 13485, CE MDR.

## 10.3 Standard Accessories

Hartmann-style paediatric nasal speculum instrument with extra-fine short blades: Stainless steel Hartmann pattern; short paediatric blades; smooth atraumatic tips; sprung handle with ratchet or return spring.

Vienna-style child nasal speculum instrument with delicate sprung handle: Stainless steel Vienna pattern; delicate sprung handle sized for paediatric use; short blades.

Perforated stainless steel instrument storage box case: Autoclavable perforated tray with lid; sized to accommodate both speculums; corrosion-resistant.

## 10.4 Installation Requirements

None.

## 10.5 Documentation

Material Certificate, Sterilization Compatibility Certificate, CE certificate.

## 10.6 Training

Brief clinical orientation.

## 10.7 Warranty

Minimum 2 years covering material defects and hinge mechanism.

# 11. VACCINE COLD-CHAIN REFRIGERATOR

Quantity: 2

## 11.1 Clinical & Functional Specifications

Medical-grade upright vaccine refrigerator providing precisely controlled internal temperature of +2 °C to +8 °C for safe storage of vaccines and thermolabile biologicals in accordance with WHO PQS and cold-chain requirements. The unit shall feature microprocessor temperature control, continuous monitoring, alarm system, data logging, and holdover capability during power interruption.

## 11.2 Engineering Technical Specifications

Temperature Range: +2 °C to +8 °C internal, uniformly maintained.

Temperature Uniformity: ≤ ±3 °C variation across storage volume.

Ambient Operating Range: +10 °C to +43 °C (tropical rating preferred).

Control: Microprocessor with digital PID control; setpoint locked or password-protected.

Display: External digital display of current temperature; min/max memory.

Alarms: High/low temperature, door-open, power failure, sensor failure; audible and visual; battery-backed alarm operation for minimum 72 hours.

Data Logging: Integrated data logger with minimum 30-day rolling storage; USB export; optional chart recorder.

Capacity: Net storage volume minimum 100 litres (configurable per site requirement).

Insulation: CFC-free, HFC-free foam insulation; minimum 60 mm thickness.

Compressor: Hermetically sealed, CFC-free refrigerant (R600a or equivalent low-GWP); low-noise operation ≤ 45 dB(A).

Door: Double- or triple-pane tempered glass or insulated solid door; magnetic seal; key lock.

Shelving: Adjustable coated-wire or coated-steel shelves; front safety lips to prevent vial fall.

Holdover Time: Minimum 4 hours at +43 °C ambient without power (WHO PQS type-tested preferred).

Power: 100—240 VAC, 50/60 Hz; voltage stabilizer compatible.

Standards: WHO PQS E003/RF category, IEC 60601-1, IEC 60335-2-89, ISO 13485, CE MDR.

## 11.3 Standard Accessories

Microprocessor control digital display temperature tracking module system (2—8 °C): External digital controller with LCD/LED display showing current, min, and max temperature; setpoint lockout.

Integrated acoustic and visual temperature deviation sensor alarm with battery backup: Audible buzzer and flashing visual indicator; internal rechargeable battery for alarm operation ≥ 72 hours during mains failure.

CFC-free heavy-duty insulation cooling compressor motor assembly: Hermetically sealed compressor using CFC/HFC-free low-GWP refrigerant; energy-efficient; low-noise.

Adjustable internal wire organizing storage shelves with safety front lips: Coated wire/steel shelves; height-adjustable; front safety lip to prevent vials from falling; corrosion-resistant.

Integrated digital tracking chart recorder or data logging USB slot port: Continuous data logger with minimum 30-day rolling memory; USB or SD-card export; timestamped records.

Dual-pane vacuum tempered safety glass door with cylinder key lock: Insulated tempered glass door with magnetic gasket seal; cylinder or padlock-compatible lock; auto-close hinges preferred.

## 11.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz outlet on stable / UPS-backed circuit; voltage stabilizer recommended in unstable-supply areas; ambient temperature not exceeding 43 °C; minimum 100 mm free clearance around unit for ventilation; level floor.

## 11.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, WHO PQS Certificate (where claimed), Temperature Mapping / Qualification Report, Electrical Safety Test Report, PM Schedule, CE certificate.

## 11.6 Training

Clinical training for EPI/vaccine staff on daily temperature monitoring, alarm response, and cold-chain protocols; biomedical training on defrost, gasket replacement, calibration verification, and troubleshooting.

## 11.7 Warranty

Minimum 2 years comprehensive covering compressor (minimum 5 years on compressor preferred), electronics, alarms, and data logger.

# 12. PAEDIATRIC RESUSCITATION TROLLEY

Quantity: 4

## 12.1 Clinical & Functional Specifications

Mobile emergency resuscitation trolley (crash cart) configured for paediatric emergency response, providing organized storage of resuscitation drugs, airway equipment, defibrillator, IV supplies, and ancillary devices, with colour-coded drawer system compatible with Broselow paediatric emergency standards.

## 12.2 Engineering Technical Specifications

Construction: High-impact ABS engineered polymer or 304 stainless steel; anti-microbial cleanable surfaces; powder-coated where painted.

Drawer System: Minimum 4—6 drawers with central breakout lock system (single-action lock for all drawers with numbered tamper-evident tags).

Colour Coding: Drawer facades / handles colour-coded per Broselow-Luten paediatric emergency scheme.

Load Rating: Minimum 100 kg overall; individual drawer minimum 15 kg.

Wheels: 4 twin-wheel castors, minimum 125 mm diameter, two with total-lock brakes; anti-static preferred.

Top Work Surface: Impact- and stain-resistant; raised perimeter with rounded corners.

Ancillary Features: CPR backboard, IV pole, defibrillator shelf, sharps bracket, waste bin(s), oxygen cylinder holder (optional).

Dimensions: Approximately 700—900 mm (L) × 500—700 mm (W) × 900—1100 mm (H) — generalized.

Standards: IEC 60601-1 (for any electrical accessories), ISO 13485, CE MDR (where applicable to combined medical device).

## 12.3 Standard Accessories

High-impact ABS or stainless steel mobile cart with central breakout lock system: Ergonomic push handle; anti-microbial finish; central single-action lock for all drawers with tamper-evident numbered seals.

Color-coded drawer handles/facades compatible with Broselow resuscitation standards: Broselow-Luten colour scheme (grey, pink, red, purple, yellow, white, blue, orange, green); durable colour-fast labelling.

Cardiac resuscitation backboard (CPR board) sized for paediatric patients: Rigid radiolucent polymer board; paediatric dimensions; smooth cleanable surface; stored on cart side or under top.

Height-adjustable telescopic I.V. pole attachment assembly: Stainless steel telescopic IV pole; 4-hook top; height range approximately 1.3—2.2 m; lockable clamp to cart.

Defibrillator installation shelf with security strap ties: Reinforced top or side shelf sized for standard defibrillator footprint; hook-and-loop or webbing security straps.

Side-mounted sharps container bracket and dual medical waste bins: Universal sharps container bracket; two colour-coded (general/clinical waste) removable bins with hinged lids.

Internal modular drawer divider set for emergency drug organizing: Adjustable modular polymer dividers to configure drawer compartments per site protocol.

## 12.4 Installation Requirements

None (mobile); storage location with minimum 800 mm clear access; charging point if defibrillator is mounted.

## 12.5 Documentation

User Manual, Parts Catalogue, Cleaning/Disinfection Guide, Broselow compatibility statement, CE certificate.

## 12.6 Training

Clinical training on cart configuration, restocking protocol, lock/seal system, and Broselow colour-coded emergency response.

## 12.7 Warranty

Minimum 2 years covering structure, wheels, locks, and drawer mechanisms.

# 13. PAEDIATRIC EXAMINATION TABLE

Quantity: 3

## 13.1 Clinical & Functional Specifications

Fixed-height or adjustable paediatric examination table for outpatient and inpatient clinical examination of infants and children, incorporating integrated measurement rule, padded mattress with wipeable antimicrobial cover, step-stool access for toddlers, and paper roll dispenser for hygienic single-patient use.

## 13.2 Engineering Technical Specifications

Frame: Reinforced tubular steel with anti-rust powder-coated finish OR stainless steel; corrosion-resistant; load rating minimum 150 kg.

Mattress: High-density foam; seamless fluid-proof antimicrobial vinyl cover; minimum 50 mm thickness; contoured or flat.

Integrated Measurement: Metric measuring scale printed or embedded along mattress edge, range 0—100 cm.

Step-Stool: Slide-out or fold-out reinforced step for toddler self-access; load rating minimum 50 kg; anti-slip surface.

Paper Roll Dispenser: Integrated rod assembly accepting standard hospital examination paper rolls; easy front-loading.

Levelling Feet: Non-skid rubber pads with individual height adjustment.

Dimensions: Approximately 1500—1800 mm (L) × 600—700 mm (W) × 750—850 mm (H).

Standards: IEC 60601-1 (if any electrical components), ISO 13485, CE MDR.

## 13.3 Standard Accessories

Multi-section reinforced frame table structure with anti-rust powder coating: Robust steel or stainless-steel frame; anti-rust epoxy powder-coated finish; welded joints; load rating minimum 150 kg.

Integrated padded infant measuring rule system layout on mattress frame: Permanent metric measuring scale (0—100 cm) printed or embedded along one edge; contrasting colour for readability.

High-density seamless vinyl mattress pad with fluid-proof antimicrobial cover: Minimum 50 mm foam; seamless welded vinyl cover; antimicrobial, latex-free, cleanable with hospital-grade disinfectants.

Slide-out step-stool drawer platform for toddler access: Reinforced pull-out or fold-out step; load rating minimum 50 kg; anti-slip textured surface; smooth glide mechanism.

Integrated heavy-duty paper sheet roll dispenser rod assembly: Rear- or side-mounted rod with retention clip; accepts standard examination paper rolls; easy replacement.

Reinforced non-skid rubber floor levelling stabilizer pads: Individually height-adjustable feet; anti-slip rubber base; corrosion-resistant.

## 13.4 Installation Requirements

Level floor; minimum 800 mm clear access on approach sides; no electrical requirements for basic version.

## 13.5 Documentation

User Manual, Assembly Instructions, Cleaning Guide, CE certificate.

## 13.6 Training

Brief orientation on step-stool operation, paper roll loading, and cleaning protocol.

## 13.7 Warranty

Minimum 2 years covering frame, mattress, and mechanical assemblies.

# 14. PAEDIATRIC STETHOSCOPE

Quantity: 2

## 14.1 Clinical & Functional Specifications

Dual-head acoustic stethoscope with paediatric-sized chestpiece for auscultation of heart, lung, and bowel sounds in infants and children, providing high acoustic sensitivity and clear sound reproduction across low- and high-frequency ranges.

## 14.2 Engineering Technical Specifications

Chestpiece: Stainless steel dual-head (diaphragm + bell) sized for paediatric patients; diaphragm diameter approximately 30—35 mm; bell diameter approximately 20—25 mm; non-chill rim.

Diaphragm: Tunable or fixed high-sensitivity diaphragm.

Tubing: Single- or dual-lumen medical-grade PVC or thermoplastic; latex-free; length approximately 55—70 cm; non-chill exterior.

Binaural: Anatomically angled, adjustable-tension.

Eartips: Soft biocompatible sealing tips (1 pair) plus firm replacement tips (1 pair); snap-tight secure fit.

Standards: ISO 13485, CE MDR.

## 14.3 Standard Accessories

Single-lumen lightweight acoustic tubing line with non-chill exterior coating: Medical-grade PVC or thermoplastic tubing; latex-free; length 55—70 cm; non-chill soft outer coating.

Stainless steel miniature dual-head chestpiece optimized for children: Polished stainless steel; diaphragm and bell sides with rotating stem; paediatric dimensions.

Tunable child-size diaphragm assembly and small open bell cone: Tunable diaphragm allowing frequency selection by pressure; open bell for low-frequency sounds; both with non-chill rim.

Soft-sealing ergonomic eartips with snap-tight plastic structural adapters (1 pair): Soft biocompatible polymer sealing tips; snap-tight secure connection to binaural.

Firm seating replacement eartips set (1 pair): Firmer alternative tips; snap-tight fit.

## 14.4 Installation Requirements

None.

## 14.5 Documentation

User Manual, Cleaning Guide, CE certificate.

## 14.6 Training

Brief clinical orientation.

## 14.7 Warranty

Minimum 2 years covering diaphragm, tubing, and chestpiece.

# 15. PAEDIATRIC GLUCOMETER

Quantity: 4

## 15.1 Clinical & Functional Specifications

Portable handheld electronic glucometer for rapid capillary blood glucose measurement in paediatric and neonatal patients, using test-strip based electrochemical (glucose oxidase or dehydrogenase) methodology, providing clinically reliable results from small blood volumes appropriate for paediatric fingerstick.

## 15.2 Engineering Technical Specifications

Measurement Range: 20—600 mg/dL (1.1—33.3 mmol/L) minimum.

Sample Volume: ≤ 1.0 µL (small-volume capable for paediatric use).

Measurement Time: ≤ 10 seconds.

Accuracy: Compliant with ISO 15197:2013 — ±15 mg/dL for values \< 100 mg/dL and ±15% for values ≥ 100 mg/dL, in ≥ 95% of measurements.

Hematocrit Range: 20—70% tolerance.

Memory: Minimum 500 results with date/time stamp; pre-/post-meal marking; average calculation (7/14/30 days).

Display: Backlit LCD; large digits for easy reading.

Coding: No-coding technology preferred (automatic strip recognition).

Connectivity: USB or Bluetooth for data transfer (preferred).

Power: Long-life coin-cell or rechargeable battery; minimum 1000 tests per battery/charge.

Environmental: Operating 10—40 °C, 10—90% RH.

Standards: ISO 15197:2013, IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR / FDA cleared.

## 15.3 Standard Accessories

Rigid plastic carrying and storage case: Impact-resistant case with foam cutouts for glucometer, lancing device, strips, and control solution; secure latch closure.

Paediatric ultra-fine depth adjustable lancing device instrument: Spring-loaded lancing pen; depth selectable across minimum 5 settings for paediatric skin thickness; compatible with standard lancets.

Check strip or optical verification strip for calibration verification: Manufacturer-supplied verification strip enabling daily function check without blood sample.

Rechargeable or long-life coin-cell battery pack sets: Minimum 1000 tests per battery/charge; user-replaceable.

## 15.4 Installation Requirements

None; portable device.

## 15.5 Documentation

User Manual, Quick Reference Card, Factory Calibration Certificate, ISO 15197 Performance Report, CE/FDA certificate.

## 15.6 Training

Clinical training for nursing staff on paediatric fingerstick technique, quality control, and result interpretation.

## 15.7 Warranty

Minimum 2 years covering unit, lancing device, and case.

# 16. NEONATAL BABY WARMER

Quantity: 6

## 16.1 Clinical & Functional Specifications

Overhead radiant infant warmer providing controlled radiant heat to maintain normothermia in newborn and preterm infants during resuscitation, examination, and routine care. The equipment shall provide servo-controlled skin-temperature regulation, manual power mode, integrated examination light, APGAR timer, and X-ray tray for bedside imaging without patient transfer.

## 16.2 Engineering Technical Specifications

Heater: Quartz or ceramic infrared radiant heating element(s); minimum 400 W output; long-life element with minimum 5000 hours service life.

Heating Modes

- Servo (skin) mode: closed-loop control via skin sensor, setpoint range 34.0—37.5 °C, resolution 0.1 °C, accuracy ±0.3 °C

- Manual (power) mode: adjustable output 0—100% in increments of ≤ 10%

- Pre-warm mode

Skin Temperature Range: 30—42 °C displayed; measurement accuracy ±0.2 °C.

Alarms: High/low skin temperature, sensor failure/disconnection, heater failure, power failure, prolonged heater override (safety timer typically 15 min in manual mode); audible and visual, adjustable priority.

Display: Backlit digital display; set and actual temperatures, heater output %, elapsed time, alarms.

APGAR Timer: Integrated countdown/count-up timer with audible marker at 1, 5, 10 minutes.

Examination Light: Integrated LED procedure lamp; minimum 20,000 lux at mattress; independently switched.

Mattress Platform: Padded, radiolucent tray; approximate dimensions 600 × 400 mm; tilting mechanism ±10° minimum (Trendelenburg / reverse).

Side Shields: Folding transparent acrylic side walls with drop-down or hinged access.

X-Ray Tray: Radiolucent cassette tray beneath mattress, accessible from bedside without patient movement; accepts standard cassette sizes.

Storage: Dual lockable drawers.

Height: Fixed or adjustable; overall height approximately 190—220 cm.

Power: 100—240 VAC, 50/60 Hz; consumption approximately 600—800 W; earthed medical-grade cord.

Environmental: Operating 18—30 °C, 20—90% RH.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-21 (Infant radiant warmers), ISO 13485, CE MDR / FDA cleared.

## 16.3 Standard Accessories

Microprocessor-controlled quartz/ceramic infrared overhead radiant heating module: Quartz or ceramic IR emitter; minimum 400 W; long-life element; safety cutoff on overtemperature; protective mesh guard.

Fixed examination mattress platform with folding transparent acrylic side shields: Radiolucent padded platform approximately 600 × 400 mm; folding transparent acrylic side walls providing 360° access with drop-down mechanism.

Reusable skin temperature probe sensors with skin attachment foil patches (set of 2): Medical-grade thermistor probes; biocompatible cable ≥ 3 m; reflective foil adhesive patches for skin attachment (starter pack included).

Integrated LED examination spotlight procedure lamp assembly: LED lamp; minimum 20,000 lux at mattress; colour temperature 4000—5500 K; independent switch; long service life ≥ 20,000 hours.

Adjustable IV pole assembly and dual lockable storage drawer compartments: Stainless-steel IV pole with 4 hooks; two drawers on trolley base with individual locks.

X-ray cassette tray holder slot underneath the patient mattress bed: Radiolucent tray accepting standard cassette sizes; slide-out access from bedside; smooth roller mechanism.

## 16.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz 15A outlet on UPS or emergency-backed circuit; ambient temperature 18—28 °C; minimum overhead clearance 220 cm; free floor area for trolley access; adequate lighting.

## 16.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test & Calibration Certificate, Electrical Safety Test Report per IEC 62353, PM Schedule, Installation & Commissioning Report, CE/FDA certificate.

## 16.6 Training

Clinical training for NICU nursing and paediatric staff on servo/manual mode use, probe placement, alarm management, APGAR timer, and safety protocols; biomedical training on element replacement, calibration verification, and troubleshooting.

## 16.7 Warranty

Minimum 2 years comprehensive covering heater element, electronics, sensors, and structural components.

# 17. PAEDIATRIC INFUSION PUMP

Quantity: 6

## 17.1 Clinical & Functional Specifications

Volumetric infusion pump for accurate, controlled intravenous administration of fluids and medications to paediatric and neonatal patients, with micro-infusion capability supporting flow rates as low as 0.1 mL/h, comprehensive safety features (air-in-line, occlusion, free-flow protection), and drug-library-based dose-error-reduction system (DERS).

## 17.2 Engineering Technical Specifications

Infusion Mechanism: Peristaltic (linear or rotary) volumetric.

Flow Rate Range

- Minimum: 0.1 mL/h or lower (paediatric/neonatal micro-infusion)

- Maximum: minimum 999 mL/h

- Resolution: 0.1 mL/h up to 100 mL/h; 1 mL/h above

Flow Accuracy: ±5% or better over full range under specified conditions.

Volume to Be Infused (VTBI): 0.1—9999 mL selectable.

Delivery Modes: Continuous rate, VTBI/time, ml/kg/h, µg/kg/min, dose calculation, secondary/piggyback, intermittent, bolus, KVO (Keep Vein Open), taper up/down (preferred).

Safety

- Air-in-line detector: ultrasonic; adjustable sensitivity; single-bubble and cumulative

- Occlusion detection: upstream and downstream; adjustable pressure limits (e.g. 100—1000 mmHg)

- Anti-free-flow: automatic when cassette/door opened

- Anti-bolus after occlusion release

- Battery-backed continuous operation

Drug Library: On-board configurable DERS with soft and hard dose limits; paediatric/neonatal profiles.

Display: Color LCD or TFT, minimum 3.5-inch; clear numerical and graphical presentation.

Alarms: Air-in-line, occlusion (up/down), door open, low battery, near-end-of-infusion, end-of-infusion, KVO, malfunction; three-level priority; audible and visual.

Memory: Event log minimum 5000 entries with time stamp; retained during power loss.

Connectivity: USB and/or wireless for drug library update; nurse-call output; HL7 / IHE PCD-01 compatible preferred.

Power

- Input: 100—240 VAC, 50/60 Hz

- Internal rechargeable Li-ion battery; runtime minimum 6 hours at 25 mL/h

- Recharge time maximum 6 hours

Environmental: Operating 5—40 °C, 15—95% RH; IP rating minimum IPX3.

Physical: Compact stackable design; pole/rail mount; weight approximately 1.5—3.0 kg.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-8, IEC 60601-2-24 (Infusion pumps), ISO 13485, CE MDR / FDA cleared.

## 17.3 Standard Accessories

Heavy-duty universal pole mounting clamp with lock screw knob: Rotatable clamp compatible with IV poles 15—40 mm diameter; large ergonomic knob; quick-release preferred.

Micro-infusion delivery tracking software supporting increments down to 0.1 mL/h: On-board software supporting flow-rate resolution of 0.1 mL/h across full paediatric range; dose-rate modes.

Integrated air-in-line and ultra-sensitive occlusion sensor system: Ultrasonic air-in-line sensor with adjustable single-bubble and cumulative thresholds; upstream and downstream occlusion sensors with adjustable pressure limits.

High-visibility graphic user interface screen with safety dosage calculation modes: Colour LCD/TFT minimum 3.5-inch; guided programming; drug-library and dose-error-reduction system.

Internal high-capacity rechargeable battery pack: Sealed medical-grade Li-ion battery; minimum 6 hours runtime at 25 mL/h; user-serviceable.

Main electrical power supply cord with medical plug: IEC 60601-compliant 3-pin earthed medical power cord.

## 17.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz outlet at bedside; UPS-backed circuit preferred; IV pole with clamp-compatible tube; ambient temperature 18—28 °C.

## 17.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate, Electrical Safety Test Report per IEC 62353, Drug Library Configuration Report, PM Schedule, CE/FDA certificate.

## 17.6 Training

Clinical training on programming, drug library, safety features, and alarm management for paediatric/neonatal nursing staff; biomedical training on PM, battery replacement, and calibration verification.

## 17.7 Warranty

Minimum 2 years comprehensive covering pump mechanism, electronics, battery, and accessories; software updates included.

# 18. NEONATAL INCUBATOR

Quantity: 6

## 18.1 Clinical & Functional Specifications

Microprocessor-controlled neonatal incubator providing a controlled thermoneutral environment (air temperature, humidity, and oxygen) for preterm and sick newborn infants, with servo-controlled skin-temperature regulation, active humidification, and comprehensive alarm and safety systems. The incubator shall permit non-disruptive nursing access, integrated tilting mattress, X-ray access without patient handling, and backup battery operation during transport within the hospital.

## 18.2 Engineering Technical Specifications

Air Temperature

- Range: 20—39 °C (extendable to 39—40 °C with override for hyperthermia therapy)

- Resolution: 0.1 °C

- Accuracy: ±0.5 °C

- Uniformity within hood: ±0.8 °C

Skin Temperature (Servo)

- Range: 34.0—37.5 °C

- Resolution 0.1 °C; accuracy ±0.3 °C

- Dual skin-sensor capability (central + peripheral)

Humidity

- Servo-controlled active humidifier

- Range: 30—95% RH

- Accuracy: ±10%

- Automatic filling or large water reservoir

Oxygen (optional/configurable): Servo O₂ control with range 21—65%; accuracy ±3%; low-oxygen alarm.

Air Circulation: Low-velocity laminar airflow across mattress; noise level inside hood ≤ 50 dB(A).

Alarms: High/low air temperature, high/low skin temperature, sensor failure, humidity failure, fan failure, power failure, oxygen failure (if fitted); three-level priority per IEC 60601-1-8; audible and visual.

Hood: Double-walled clear acrylic canopy for thermal insulation and reduced condensation; minimum 4—6 hand-access ports with silicone iris/gasket seals; large front access door; tube grommets.

Mattress: Radiolucent padded mattress; motorized continuous tilting mechanism ±10° or better (Trendelenburg / reverse); integrated weighing scale optional/configurable.

X-Ray Tray: Radiolucent slide-out cassette tray beneath mattress accessible from outside hood; accepts standard cassette sizes.

Cabinet/Trolley: Height-adjustable base (electric or manual) preferred; lockable castors; storage drawers; anti-microbial cleanable surfaces.

Backup Power: Integrated rechargeable battery for continuous operation during intra-hospital transport; runtime minimum 30 minutes at set conditions.

Power: 100—240 VAC, 50/60 Hz; consumption approximately 500—800 W.

Environmental: Operating 18—30 °C, 10—90% RH.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-8, IEC 60601-2-19 (Infant incubators), IEC 60601-2-20 (Transport incubators — where applicable), ISO 13485, CE MDR / FDA cleared.

## 18.3 Standard Accessories

Double-walled acrylic isolation hood canopy with multi-port access hand hatches: Double-wall clear acrylic construction for thermal insulation; minimum 4—6 iris-sealed access ports; front-opening door; tube pass-through grommets; smooth cleanable surfaces.

Motorized continuous tilting examination mattress cradle mechanism: Electric or gas-assisted tilting mechanism providing continuous adjustment ±10° or better (Trendelenburg / reverse); smooth, quiet operation.

Servo-controlled integrated active air humidification chamber assembly: Heated humidifier with servo control; range 30—95% RH; large-capacity or auto-refill reservoir; easily removable and cleanable.

Reusable skin temperature probe sensors (set of 2, skin/air modes control): Medical-grade thermistor probes with biocompatible cable; length ≥ 3 m; reflective adhesive patches (starter pack); calibration traceable.

X-ray cassette slide tray platform accessible from outside the hood: Radiolucent slide-out cassette tray beneath mattress; accepts standard cassette sizes; enables X-ray without patient movement.

Heavy-duty mobile cabinet stand trolley with lockable wheels and storage drawers: Powder-coated steel or stainless-steel base with height-adjustment mechanism; minimum 2 lockable storage drawers; 4 castors with minimum 2 with brakes; anti-microbial surfaces.

Integrated backup emergency battery supply system: Sealed medical-grade rechargeable battery; runtime minimum 30 minutes; automatic changeover on mains failure with visual/audible indicator.

## 18.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz 15A outlet on UPS / emergency-backed circuit; ambient temperature 20—28 °C; humidity ≤ 70% RH; distilled or sterile water supply for humidifier (manual refill acceptable); minimum floor area approximately 2 m² per unit with adequate access; NICU-grade environmental hygiene.

## 18.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test & Calibration Certificate, Electrical Safety Test Report per IEC 62353, Installation & Commissioning Report, Environmental Performance Test Report (temperature uniformity, humidity, sound), PM Schedule, CE/FDA certificate.

## 18.6 Training

Comprehensive clinical training for NICU nursing and neonatology staff on modes of operation, servo/manual control, humidity management, alarm response, and safety protocols; biomedical training on PM, calibration verification, humidifier maintenance, filter replacement, and troubleshooting.

## 18.7 Warranty

Minimum 2 years comprehensive covering incubator, humidifier, sensors, motorized components, battery, and electronics; software updates included.

# 19. LED PHOTOTHERAPY UNIT

Quantity: 4

## 19.1 Clinical & Functional Specifications

High-intensity narrow-band blue LED phototherapy unit for treatment of neonatal hyperbilirubinaemia (jaundice), delivering therapeutic light in the 450—470 nm wavelength range corresponding to peak bilirubin absorption. The unit shall provide uniform irradiance over the treatment area, adjustable output, integrated hour counter for lamp-life monitoring, and mobile pedestal for use with incubators, warmers, or bassinets.

## 19.2 Engineering Technical Specifications

Light Source: High-intensity blue LED array; peak emission wavelength 450—470 nm; narrow-band spectrum optimized for bilirubin photodegradation.

Irradiance

- Intensive phototherapy: minimum 30 µW/cm²/nm at standard treatment distance

- Standard phototherapy: minimum 15 µW/cm²/nm

- Adjustable intensity settings (minimum 2 levels; preferred continuous or 3-level)

Treatment Area: Uniform irradiance over minimum 30 × 50 cm at recommended distance.

LED Life: Minimum 20,000 hours rated service life; no UV or IR emission; low heat generation.

Hour Counter: Integrated non-resettable (or protected) electronic hour counter displaying total lamp usage.

Height Adjustment: Gas-spring or mechanical height adjustment; range approximately 100—180 cm above floor.

Angle Adjustment: Tiltable lamp head for optimal positioning.

Mobile Stand: 4- or 5-castor pedestal with at least 2 lockable castors.

Power: 100—240 VAC, 50/60 Hz; low power consumption ≤ 100 W typical.

Noise Level: ≤ 45 dB(A) (fan-cooled units).

Environmental: Operating 10—40 °C, 20—90% RH.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-50 (Infant phototherapy equipment), ISO 13485, CE MDR / FDA cleared.

## 19.3 Standard Accessories

High-intensity blue LED light tube array module optimized for bilirubin degradation: Narrow-band blue LED array; peak wavelength 450—470 nm; uniform irradiance distribution; no UV/IR emission; rated life ≥ 20,000 hours.

Mobile pedestal stand with gas-spring assisted height adjustment and caster wheels: Powder-coated steel column with gas-spring or mechanical height adjustment (100—180 cm); 4 or 5 castors with minimum 2 lockable; stable weighted base.

Integrated electronic total lamp usage hour-counter display meter: Digital hour counter; tamper-resistant; visible on control panel; retained during power loss.

Opaque neonate protective eye-shield masks (starter pack of 5, adjustable sizes): Soft opaque foam or fabric eye masks with hook-and-loop adjustable strap; latex-free; sizes for preterm through term neonates; single-patient or reusable per manufacturer.

Detachable transparent acrylic shield plate cover guard: Clear acrylic shield below LED array preventing accidental contact and protecting LEDs; easily removable for cleaning.

## 19.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz outlet; ambient temperature 18—30 °C; no special HVAC or shielding.

## 19.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate including Irradiance Measurement Report, Electrical Safety Test Report per IEC 62353, PM Schedule, CE/FDA certificate.

## 19.6 Training

Clinical training for NICU nursing staff on positioning, irradiance measurement, treatment protocols, and eye protection; biomedical training on radiometer verification, LED replacement, and troubleshooting.

## 19.7 Warranty

Minimum 2 years comprehensive covering LED array, electronics, gas-spring mechanism, and castors; LED array warranted for minimum rated life.

# 20. BUBBLE CPAP

Quantity: 4

## 20.1 Clinical & Functional Specifications

Bubble Continuous Positive Airway Pressure (bCPAP) system for non-invasive respiratory support of neonates with respiratory distress syndrome (RDS), transient tachypnoea, and other neonatal respiratory conditions. The system shall provide precisely blended, heated, humidified gas at continuous positive airway pressure generated by underwater column depth, with independent control of flow, FiO₂, and pressure.

## 20.2 Engineering Technical Specifications

Gas Blending

- Precision air-oxygen blender

- FiO₂ range: 21—100%; accuracy ±3%

- Independent flow control with dual flowmeters (or single blended flowmeter): range 0—15 L/min

Pressure Generation

- Bubble/underwater column pressure generator

- CPAP pressure range: 3—10 cmH₂O (minimum), adjustable in 1 cmH₂O increments

- Clear graduated column for visual verification of pressure setting

- Overpressure safety relief

Humidification

- Heated humidifier with servo temperature control

- Chamber temperature 31—37 °C adjustable

- Auto-refill or manual-fill water chamber

- Heated wire circuit compatibility

Circuit: Heated wire dual-limb neonatal breathing circuit; reusable/autoclavable or single-use compatible.

Interfaces: Neonatal nasal prongs (multiple sizes) and nasal masks; secured with silicone bonnets.

Mobile Stand: Pole stand with lockable castors, circuit support arm, and humidifier/blender mounting.

Power: 100—240 VAC, 50/60 Hz (for humidifier); gas supply from central pipeline or cylinders.

Alarms: Humidifier temperature high/low, water low, power failure (audible/visual).

Environmental: Operating 18—30 °C, 20—90% RH.

Standards: IEC 60601-1, IEC 60601-1-2, ISO 80601-2-74 (Respiratory humidifiers), ISO 80601-2-87 or applicable CPAP standards, ISO 13485, CE MDR / FDA cleared.

## 20.3 Standard Accessories

Precision air-oxygen gas blender module with independent dual flowmeters: Mechanical or electronic air/O₂ blender; FiO₂ 21—100%, accuracy ±3%; dual Thorpe-tube or digital flowmeters, range 0—15 L/min; standard NIST or DISS gas inlet connectors.

Heated humidifier base unit with automatic-filling water chamber column: Servo-controlled heated humidifier; temperature setpoint 31—37 °C; auto-refill or large-capacity chamber; heated wire circuit compatible; low-water alarm.

Bubble generator column chamber with adjustable pressure regulatory water scale (3 to 10 cmH₂O): Clear graduated cylindrical column; adjustable submersion depth for CPAP pressure 3—10 cmH₂O; overpressure relief; secure mounting.

Heated wire patient breathing tube circuits (set of 2, reusable/washable): Dual-limb neonatal circuits with integrated heated wire; latex-free medical-grade material; reusable/autoclavable per manufacturer specification; low resistance.

Paediatric/neonatal nasal prongs and silicone infant bonnets (starter kit of various sizes): Soft silicone binasal prongs in multiple sizes (preterm to term) and nasal masks; silicone/fabric bonnets with adjustable straps for secure fixation; latex-free.

Mobile pole stand with lockable caster wheels and circuit support arm bracket: Powder-coated steel or stainless-steel pole; 4 castors with minimum 2 lockable; articulated circuit support arm to prevent circuit drag on infant.

## 20.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz outlet for humidifier; medical air and oxygen gas supply at 3.5—4.5 bar (50—65 psi) from central pipeline or cylinders with appropriate NIST/DISS terminals; distilled/sterile water for humidification; NICU environment with ambient temperature 20—28 °C.

## 20.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Test Report per IEC 62353, Gas System Compatibility Report, PM Schedule, CE/FDA certificate.

## 20.6 Training

Clinical training for NICU nursing and neonatology staff on CPAP setup, interface fitting, pressure/FiO₂ adjustment, humidification management, and troubleshooting; biomedical training on blender calibration, humidifier maintenance, and circuit reprocessing.

## 20.7 Warranty

Minimum 2 years comprehensive covering blender, humidifier, pressure generator, and mobile stand; consumables (prongs, bonnets) excluded after commissioning stock.

# 21. VENTILATOR WITH PAEDIATRIC & NEONATAL MODES

Quantity: 4

## 21.1 Clinical & Functional Specifications

Microprocessor-controlled intensive-care mechanical ventilator providing full range of invasive and non-invasive ventilation modes for neonatal, paediatric, and where required adult patients, with high-frequency ventilation capability, precise gas delivery, comprehensive monitoring, and integrated humidification. Suitable for NICU, PICU, and paediatric high-dependency use.

## 21.2 Engineering Technical Specifications

Patient Range: Neonatal (from ≥ 300 g), Paediatric, Adult selectable; automatic circuit compliance compensation.

Ventilation Modes (minimum)

- Volume Control (VC / VCV)

- Pressure Control (PC / PCV)

- SIMV (Volume and Pressure)

- Pressure Support (PSV)

- CPAP / PEEP

- PRVC or equivalent dual-control mode

- Non-Invasive Ventilation (NIV / NIV-PSV / NIV-CPAP)

- Apnoea backup ventilation

- High-Frequency Ventilation (HFV / HFOV) — as configured/optional

- Neonatal-specific modes (e.g., flow-cycled, pressure-limited time-cycled)

Parameters

- Tidal Volume: 2—2000 mL (neonatal range from 2 mL)

- Respiratory Rate: 1—150 bpm (neonatal to 150; adult 1—80)

- I:E Ratio: 1:9 to 4:1 selectable

- PEEP/CPAP: 0—35 cmH₂O

- Pressure Support: 0—60 cmH₂O

- FiO₂: 21—100%, accuracy ±3%

- Inspiratory Flow: up to 180 L/min

- Trigger: flow (0.1—15 L/min) and pressure (—0.5 to —20 cmH₂O)

Monitoring

- Airway pressure (peak, mean, plateau, PEEP)

- Delivered tidal volume, minute volume, spontaneous VT/MV

- Respiratory rate (total, spontaneous, mandatory)

- FiO₂, I:E, compliance, resistance, auto-PEEP

- Waveforms: pressure, flow, volume vs time; loops (P-V, F-V, P-F)

- Trends: minimum 72 hours

Display: Color TFT touchscreen; minimum 12-inch diagonal; simultaneous waveforms, numerics, and trends; adjustable brightness.

Alarms: Three-level priority per IEC 60601-1-8; comprehensive (high/low pressure, volume, MV, rate, FiO₂, apnoea, disconnection, occlusion, power, gas supply); audible and visual; history log ≥ 1000 events.

Gas Supply: Medical air and oxygen at 2.8—6.0 bar; standard NIST or DISS connectors; internal turbine option preferred for redundancy.

Humidification: Integrated or dedicated heated humidifier with servo control; compatible with heated-wire circuits.

Connectivity: RS-232, USB, Ethernet; HL7 / IHE PCD compatible; nurse-call output; VGA/HDMI mirror output preferred.

Power

- 100—240 VAC, 50/60 Hz

- Integrated rechargeable battery; runtime minimum 60 minutes at typical settings

- Automatic changeover; low-battery alarm

Environmental: Operating 10—40 °C, 15—95% RH.

Physical: Mobile trolley with lockable castors; articulated circuit support arm; weight approximately 60—100 kg with trolley.

Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-8, ISO 80601-2-12 (Critical care ventilators), ISO 80601-2-79 / -87 (Home care / non-invasive as applicable), IEC 62304 (software), ISO 14971 (risk management), ISO 13485, CE MDR / FDA cleared.

## 21.3 Standard Accessories

Heavy-duty mobile trolley with lockable casters and articulated circuit support arm: Powder-coated steel or aluminium mobile trolley; 4 castors (minimum 125 mm) with 2 lockable; articulated multi-joint circuit support arm; integrated cylinder holder (optional); storage shelves.

High-frequency ventilation (HFV) or proximal flow sensor tracking modules: HFV module (as configured) providing high-frequency oscillatory or jet ventilation OR proximal flow sensor module for accurate neonatal flow/volume measurement at the patient wye; interchangeable/upgradeable.

Reusable neonatal silicone breathing circuit with inline water traps: Neonatal-sized dual-limb silicone circuit; inline water traps; heated-wire compatible; autoclavable per manufacturer specification.

Reusable paediatric silicone breathing circuit with inline water traps: Paediatric-sized dual-limb silicone circuit; inline water traps; heated-wire compatible; autoclavable per manufacturer specification.

Integrated electronic or ultrasonic humidifier base unit with reusable chamber: Servo-controlled heated humidifier; temperature and humidity setpoints adjustable; reusable/autoclavable chamber; heated-wire circuit compatible; low-water and over-temperature alarms.

High-pressure oxygen and medical air inlet hoses with color-coded connectors: Medical-grade high-pressure gas hoses; O₂ (white/green per regional standard) and air (black/yellow) colour-coded; NIST or DISS standard terminals matched to hospital pipeline; length minimum 3 m.

Reusable neonatal proximal flow sensors (pack of 2) and expiratory valve membranes: Neonatal proximal flow sensors (differential-pressure or hot-wire type); pack of 2 reusable/autoclavable per manufacturer; replacement expiratory valve diaphragm membranes.

Built-in rechargeable backup battery pack for continuous emergency use: Sealed medical-grade Li-ion battery; runtime minimum 60 minutes at typical ventilation settings; automatic changeover; user-replaceable service part.

## 21.4 Installation Requirements

Standard 100—240 VAC 50/60 Hz 15A outlet on UPS / emergency-backed critical circuit; medical oxygen and medical air pipeline at 2.8—6.0 bar with appropriate NIST/DISS terminals within 3 m of installation; ambient temperature 20—28 °C; humidity ≤ 80% RH; NICU/PICU-grade environmental hygiene; network connection (RJ-45) recommended for HIS integration; minimum floor area 2 m² per unit with adequate access.

## 21.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test & Calibration Certificate, Electrical Safety Test Report per IEC 62353, Gas System Compatibility Report, Installation & Commissioning Report, Performance Verification Report, PM Schedule, Software Version and Update Log, CE/FDA certificate.

## 21.6 Training

Comprehensive clinical training for NICU/PICU nursing, respiratory therapy, neonatology, and paediatric intensive care staff on all ventilation modes, alarm management, weaning protocols, and troubleshooting; biomedical engineering training on PM, calibration verification, sensor replacement, humidifier maintenance, and software updates; refresher training at 6 and 12 months; application specialist support during initial commissioning period.

## 21.7 Warranty

Minimum 2 years comprehensive covering ventilator, humidifier, sensors, battery, trolley, and all electronic and pneumatic components; all software updates and patches included for warranty duration; response time not exceeding 24 hours for critical failures; minimum 2 preventive maintenance visits per year.