HOSPITAL TECHNICAL TENDER SPECIFICATIONS — PATIENT MONITORING EQUIPMENT

Technical Tender Document

Table of Contents

1. ICU-Grade Multipara Patient Monitor (ECG, SpO₂, NIBP, Dual IBP, EtCO₂, Dual-Temp, Cardiac Output)	6

1.1 Clinical & Functional Specifications	6

1.2 Engineering Technical Specifications	6

1.3 Standard Accessories	8

1.4 Installation Requirements	9

1.5 Documentation	9

1.6 Training	9

1.7 Warranty	9

2. HDU Multipara Patient Monitor (ECG, SpO₂, NIBP, EtCO₂, Temp)	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	10

2.4 Installation Requirements	11

2.5 Documentation	11

2.6 Training	11

2.7 Warranty	11

3. PACU Multipara Patient Monitor (ECG, SpO₂, NIBP, EtCO₂, Temp)	12

3.1 Clinical & Functional Specifications	12

3.2 Engineering Technical Specifications	12

3.3 Standard Accessories	12

3.4 Installation Requirements	13

3.5 Documentation	13

3.6 Training	13

3.7 Warranty	13

4. Multipara Patient Monitor — Resuscitation Bay	14

4.1 Clinical & Functional Specifications	14

4.2 Engineering Technical Specifications	14

4.3 Standard Accessories	14

4.4 Installation Requirements	15

4.5 Documentation	15

4.6 Training	15

4.7 Warranty	15

5. OT-Grade Multipara Patient Monitor	16

5.1 Clinical & Functional Specifications	16

5.2 Engineering Technical Specifications	16

5.3 Standard Accessories	17

5.4 Installation Requirements	17

5.5 Documentation	17

5.6 Training	17

5.7 Warranty	17

6. Ward-Grade Basic Multipara Patient Monitor	18

6.1 Clinical & Functional Specifications	18

6.2 Engineering Technical Specifications	18

6.3 Standard Accessories	18

6.4 Installation Requirements	19

6.5 Documentation	19

6.6 Training	19

6.7 Warranty	19

7. Patient Monitor — Basic Multi-Parameter (OBS)	20

7.1 Clinical & Functional Specifications	20

7.2 Engineering Technical Specifications	20

7.3 Standard Accessories	20

7.4 Installation Requirements	21

7.5 Documentation	21

7.6 Training	21

7.7 Warranty	21

8. Cardiac Monitor (Seven Parameters)	22

8.1 Clinical & Functional Specifications	22

8.2 Engineering Technical Specifications	22

8.3 Standard Accessories	22

8.4 Installation Requirements	23

8.5 Documentation	23

8.6 Training	23

8.7 Warranty	23

9. Cardiac Monitor (Five Parameters)	24

9.1 Clinical & Functional Specifications	24

9.2 Engineering Technical Specifications	24

9.3 Standard Accessories	24

9.4 Installation Requirements	25

9.5 Documentation	25

9.6 Training	25

9.7 Warranty	25

10. Hemodynamic Monitor with Cardiac Output	26

10.1 Clinical & Functional Specifications	26

10.2 Engineering Technical Specifications	26

10.3 Standard Accessories	26

10.4 Installation Requirements	27

10.5 Documentation	27

10.6 Training	27

10.7 Warranty	27

11. Central Monitoring Station	28

11.1 Clinical & Functional Specifications	28

11.2 Engineering Technical Specifications	28

11.3 Standard Accessories	28

11.4 Installation Requirements	29

11.5 Documentation	29

11.6 Training	29

11.7 Warranty	29

12. BIS Monitor	31

12.1 Clinical & Functional Specifications	31

12.2 Engineering Technical Specifications	31

12.3 Standard Accessories	31

12.4 Installation Requirements	32

12.5 Documentation	32

12.6 Training	32

12.7 Warranty	32

13. Neuro Monitoring (IONM)	33

13.1 Clinical & Functional Specifications	33

13.2 Engineering Technical Specifications	33

13.3 Standard Accessories	34

13.4 Installation Requirements	34

13.5 Documentation	34

13.6 Training	35

13.7 Warranty	35

14. Pulse Oximeter (Handheld, Portable)	36

14.1 Clinical & Functional Specifications	36

14.2 Engineering Technical Specifications	36

14.3 Standard Accessories	36

14.4 Installation Requirements	37

14.5 Documentation	37

14.6 Training	37

14.7 Warranty	37

15. Glucometer	38

15.1 Clinical & Functional Specifications	38

15.2 Engineering Technical Specifications	38

15.3 Standard Accessories	38

15.4 Installation Requirements	39

15.5 Documentation	39

15.6 Training	39

15.7 Warranty	39

# 1. ICU-Grade Multipara Patient Monitor (ECG, SpO₂, NIBP, Dual IBP, EtCO₂, Dual-Temp, Cardiac Output)

Quantity: 27

## 1.1 Clinical & Functional Specifications

The equipment shall be a high-acuity, modular or pre-configured multi-parameter bedside patient monitor intended for continuous physiological monitoring of critically ill adult, pediatric, and neonatal patients in Intensive Care Units, Coronary Care Units, and high-dependency environments. It shall provide continuous, simultaneous, real-time monitoring and display of the following clinical parameters: 3/5/12-lead ECG with ST-segment and arrhythmia analysis, heart rate (HR), respiration rate (RR) by thoracic impedance, functional oxygen saturation (SpO₂), pulse rate (PR), non-invasive blood pressure (NIBP), a minimum of two invasive blood pressure (IBP) channels, end-tidal CO₂ (EtCO₂) with inspired CO₂ and respiration rate, dual-channel temperature (T1/T2 and ΔT), and cardiac output (CO) by thermodilution technique with derived hemodynamic parameters.

The monitor shall support adult, pediatric, and neonatal patient categories with automatically adjusted default alarm limits, measurement ranges, and algorithms per selected category. It shall provide arrhythmia detection (minimum lethal and non-lethal categories), ST-segment analysis on all available ECG leads, pacemaker pulse detection and rejection, and QT/QTc interval analysis. The user interface shall support at least three user-selectable screen layouts (standard numeric-waveform view, large font view for distant viewing, trend/OxyCRG/hemodynamic calculation view, and split-screen).

The monitor shall provide continuous storage and retrieval of tabular and graphical trends for a minimum of 96 hours, alarm event review with associated waveforms, minimum 24-hour full-disclosure waveform storage for all displayed parameters, drug dose calculator, hemodynamic calculation package, ventilation calculation package, and oxygenation calculation package. It shall support networked operation with a central monitoring station and interoperability with the hospital's electronic medical record (EMR) using standard HL7 messaging.

## 1.2 Engineering Technical Specifications

Display: Color TFT LCD or equivalent flat-panel technology; minimum 15-inch diagonal active viewing area; minimum resolution 1280 × 800 pixels or better; capacitive multi-touch touchscreen with hard-key backup for critical functions; minimum viewing angle 160° horizontal/vertical; anti-glare treated surface; capable of simultaneous display of at least 8 waveforms and all corresponding numeric parameters.

Processor / Memory / Storage: Medical-grade embedded processor architecture with sufficient computing capacity for simultaneous real-time acquisition, analysis, and display of all specified parameters without waveform lag; internal solid-state non-volatile memory sufficient to retain configuration, trends, and full-disclosure data through complete power loss.

ECG

- Leads: Selectable 3-lead / 5-lead / 12-lead (upgradable)

- Bandwidth: Minimum 0.05 — 150 Hz (diagnostic mode); monitoring and surgical modes selectable

- Sampling rate: Minimum 500 samples/second/channel; 1000 samples/second/channel or better for diagnostic 12-lead

- HR range: 15 — 300 bpm (adult/pediatric); 15 — 350 bpm (neonate)

- HR accuracy: ±1 bpm or ±1%, whichever is greater

- ST analysis: −2.0 to +2.0 mV, accuracy ±0.02 mV or better

- Arrhythmia detection: Minimum lethal (asystole, VF, sustained VT) plus multiple non-lethal categories

- ESU protection and defibrillator-proof input; recovery within 5 seconds post-defibrillation

Respiration (Impedance): Range 0 — 150 rpm; accuracy ±2 rpm; apnea detection with selectable delay (10 — 40 s).

SpO₂: Measurement range 0 — 100%; accuracy ±2% (70 — 100%, adult/pediatric); ±3% (neonate); motion-tolerant and low-perfusion capable algorithm; pulse rate range 20 — 300 bpm.

NIBP: Oscillometric technique; adult/pediatric/neonatal modes; measurement range: Systolic 30 — 270 mmHg (adult), Diastolic 10 — 220 mmHg, MAP 20 — 235 mmHg; accuracy ±3 mmHg or per IEC 80601-2-30; modes: manual, automatic (selectable interval 1 — 480 minutes), STAT, continuous; over-pressure protection per patient category.

Dual IBP: Minimum 2 independent channels (expandable to 4); range −50 to +300 mmHg; accuracy ±2% or ±1 mmHg; zeroing range ±200 mmHg; compatible with standard disposable transducers (5 μV/V/mmHg sensitivity).

EtCO₂: Side-stream and/or main-stream capable; range 0 — 99 mmHg (0 — 13.2 kPa); accuracy ±2 mmHg (0—40 mmHg), ±5% of reading above; sampling flow (side-stream) approximately 50 — 200 mL/min selectable; RR range 0 — 150 rpm; automatic barometric compensation.

Dual Temperature: 2 independent channels with ΔT; range 0 — 50°C; accuracy ±0.1°C (25 — 45°C); YSI-400 series compatible.

Cardiac Output: Thermodilution technique; CO range 0.1 — 20 L/min; accuracy ±5% or ±0.2 L/min; injectate temperature range 0 — 27°C; derived parameters including CI, SV, SVI, SVR, SVRI, PVR, LCW, RCW, etc.

Alarms: Three-level priority (high/medium/low); visual and audible; audio-paused, alarm silence, and alarm-off with regulatory-compliant safeguards; nurse call relay output.

Power: Universal input 100 — 240 VAC, 50/60 Hz; internal rechargeable lithium-ion battery, minimum 2 hours runtime under typical monitoring load; battery charging time maximum 6 hours; battery status indicator.

Connectivity: Minimum 2 × USB, Ethernet (RJ45) with wired network capability, Wi-Fi (IEEE 802.11 a/b/g/n or better), nurse call output, defibrillator sync output, analog output (optional), HL7-capable interface; DICOM optional.

Environmental: Operating temperature 5 — 40°C; storage −20 to +60°C; relative humidity 15 — 95% non-condensing; altitude up to 3000 m.

Physical: Compact bedside/rail-mountable configuration; weight approximately 6 — 10 kg with modules; ingress protection minimum IPX1 (IPX2 or better preferred).

Applicable Standards: IEC 60601-1 (general safety), IEC 60601-1-2 (EMC), IEC 60601-1-6 (usability), IEC 60601-1-8 (alarms), IEC 60601-2-27 (ECG), IEC 60601-2-49 (multifunction monitors), IEC 80601-2-30 (NIBP), ISO 80601-2-55 (respiratory gas — capnography), ISO 80601-2-61 (pulse oximetry), IEC 62304 (medical software), IEC 62366 (usability engineering), ISO 13485, ISO 14971, CE MDR and/or US FDA 510(k) clearance.

## 1.3 Standard Accessories

- 3/5-lead ECG Patient Cable and Reusable Lead Wire Set — Medical-grade shielded trunk cable with defibrillation-protected, AAMI/IEC color-coded lead wires; minimum length 2.5 m; reusable; latex-free; compatible with standard snap or clip electrodes; ESU-interference-suppressed construction; autoclave-free chemical cleaning compatible.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Silicone or medical-grade soft-shell reusable finger clip; wavelengths compliant with standard red/infrared oximetry (\~660/940 nm); minimum cable length 1.0 m sensor plus 2.0 m extension cable; latex-free; low-level disinfectant compatible.

- Reusable Adult NIBP Cuff (Standard Size) with Air Hose — Nylon/polyurethane bladder-in-sleeve reusable cuff with hook-and-loop closure, adult standard size (approximately 25 — 35 cm arm circumference); dual-lumen or single-lumen air hose (as per system) minimum length 3.0 m; latex-free; wipe-clean compatible.

- Dual IBP Interface Cables — Two shielded reusable cables compatible with disposable transducers of internationally standardized 5 μV/V/mmHg sensitivity; minimum length 3.0 m; defibrillation protected; sterilization-free wipe-clean.

- Side-Stream or Main-Stream EtCO₂ Sensor Module with Sampling Line Kit — Complete CO₂ measurement assembly with water-trap or Nafion drying tube (side-stream); airway adapter (main-stream); starter pack of disposable sampling lines and adapters for adult and pediatric circuits.

- Reusable Skin Temperature Probe and Reusable Rectal/Esophageal Temperature Probe — YSI-400-series-compatible thermistor probes; skin probe with adhesive-compatible sensing surface; rectal/esophageal probe with smooth biocompatible tip; minimum cable length 3.0 m; chemical disinfection compatible.

- Cardiac Output (CO) Interface Cable, Thermistor Cable, and Injectate Temperature Probe — Complete CO measurement kit compatible with standard thermodilution catheters; in-line injectate temperature sensor for room-temperature and cold injectate; reusable cables minimum 3.0 m.

- Integrated Rechargeable Battery Pack — Sealed lithium-ion battery; minimum 2 hours continuous runtime under typical clinical load; minimum 500 charge/discharge cycles at ≥80% capacity retention; battery management system with over-charge and thermal protection.

- Bedside Rail Mounting Bracket or Wall Mount Assembly — Adjustable stainless-steel/aluminum mounting bracket compatible with standard hospital bedside rails (25 mm and 38 mm rail systems) or VESA-compatible wall mount; load rating suitable for monitor weight with modules and cables.

## 1.4 Installation Requirements

Dedicated 100 — 240 VAC, 50/60 Hz, 15A grounded medical-grade wall outlet per monitor location; connection to hospital UPS or isolated power system preferred; proper equipotential grounding per IEC 60601-1; network drop (Cat 6 minimum) or wireless coverage at each bed for central station connectivity; ambient temperature 18 — 27°C with relative humidity 30 — 70%; adequate ventilation clearance around device.

## 1.5 Documentation

Supplier shall provide: User Manual (English), Service Manual with schematics/board-level troubleshooting, Spare Parts Catalogue, Factory Test Certificate, Calibration Certificate, Electrical Safety Test Report per IEC 62353, EMC Compliance Certificate, Declaration of Conformity, Preventive Maintenance Schedule, Installation Report, and Commissioning Report — all in printed and digital formats.

## 1.6 Training

Supplier shall provide on-site: Clinical/Applications Training for nursing and medical staff; Biomedical Engineering Training covering configuration, network setup, and calibration; Preventive Maintenance and basic troubleshooting training with practical hands-on sessions; documented training completion certificates.

## 1.7 Warranty

Minimum two (2) years comprehensive warranty covering all parts, labor, on-site preventive and breakdown maintenance, software updates and patches, and 24/7 technical support. Extended AMC options shall be quoted separately.

# 2. HDU Multipara Patient Monitor (ECG, SpO₂, NIBP, EtCO₂, Temp)

Quantity: 10

## 2.1 Clinical & Functional Specifications

Mid-acuity multi-parameter bedside patient monitor for High-Dependency Units, step-down units, and sub-intensive care. Shall provide continuous simultaneous monitoring of ECG (3/5-lead), heart rate, respiration by impedance, SpO₂, pulse rate, NIBP, side-stream EtCO₂ with inspired CO₂ and airway respiration rate, and skin temperature. Supports adult, pediatric, and neonatal categories with category-appropriate alarm defaults, arrhythmia detection (minimum basic lethal categories), ST-segment monitoring, drug-dose calculator, and networking to central station. Shall provide tabular and graphical trends for minimum 72 hours, alarm/event review, and HL7-compatible EMR interfacing.

## 2.2 Engineering Technical Specifications

Display: Color TFT LCD; minimum 12-inch diagonal; minimum resolution 1024 × 768 pixels; touchscreen with hard-key backup; anti-glare surface; capable of displaying at least 6 waveforms simultaneously.

ECG: 3/5-lead selectable; bandwidth 0.05 — 150 Hz; sampling ≥ 500 samples/s/channel; HR range 15 — 300 bpm, accuracy ±1 bpm or ±1%; ST analysis −2.0 to +2.0 mV; basic arrhythmia detection; defibrillator-proof input; ESU protected.

Respiration: 0 — 150 rpm, ±2 rpm; apnea alarm 10 — 40 s selectable.

SpO₂: Range 0 — 100%, accuracy ±2% (70 — 100%); motion-tolerant algorithm; pulse rate 25 — 300 bpm.

NIBP: Oscillometric; adult/pediatric/neonatal modes; ranges as per IEC 80601-2-30; accuracy ±3 mmHg; manual/auto/STAT modes.

EtCO₂: Side-stream capable; range 0 — 99 mmHg; accuracy ±2 mmHg (0—40 mmHg), ±5% of reading above; sampling flow selectable 50 — 200 mL/min; automatic zero and baro-compensation.

Temperature: Minimum 1 channel (expandable); range 0 — 50°C; accuracy ±0.1°C; YSI-400 compatible.

Alarms: Three-level priority per IEC 60601-1-8; visual and audible; nurse call output.

Power: 100 — 240 VAC, 50/60 Hz; rechargeable lithium-ion battery ≥ 2 hours runtime; charging ≤ 6 hours.

Connectivity: USB ×2, Ethernet RJ45, Wi-Fi (802.11 a/b/g/n), HL7-capable, nurse call output.

Environmental: 5 — 40°C operating; humidity 15 — 95% non-condensing.

Physical: Weight approximately 5 — 8 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-8, IEC 60601-2-27, IEC 60601-2-49, IEC 80601-2-30, ISO 80601-2-55, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 2.3 Standard Accessories

- 3/5-lead ECG Patient Cable and Lead Wire Set — Defib-protected, AAMI/IEC color-coded, latex-free, min 2.5 m, reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; sensor + extension total min 3.0 m; disinfectant compatible.

- Reusable Adult NIBP Cuff with Interlocking Air Hose — Standard adult size (25—35 cm); latex-free bladder-in-sleeve; min 3.0 m hose.

- Side-Stream EtCO₂ Sensor Adapter and Starter Sampling Line Kit — Includes water trap, adult/pediatric adapters, and starter set of disposable sampling lines.

- Reusable Skin Temperature Probe — YSI-400 compatible thermistor; skin-surface sensor; min 3.0 m cable; chemical disinfection compatible.

- Integrated Rechargeable Lithium-Ion Battery — Min 2 h runtime, ≥ 500 cycles at 80% capacity, protected battery management circuit.

- Bedside Rail Mounting Bracket — Adjustable stainless-steel/aluminum bracket compatible with 25 mm/38 mm hospital rails.

## 2.4 Installation Requirements

100—240 VAC 50/60 Hz medical-grade outlet; UPS backup preferred; equipotential grounding per IEC 60601-1; Cat 6 network drop or Wi-Fi coverage; ambient 18—27°C, RH 30—70%.

## 2.5 Documentation

User manual, service manual, spare parts catalogue, factory test certificate, calibration certificate, electrical safety test report (IEC 62353), EMC compliance, Declaration of Conformity, PM schedule, installation and commissioning reports.

## 2.6 Training

Clinical, biomedical engineering, PM, and basic troubleshooting training on site, with completion certificates.

## 2.7 Warranty

Minimum 2 years comprehensive: parts, labor, PM, breakdown maintenance, software updates, technical support.

# 3. PACU Multipara Patient Monitor (ECG, SpO₂, NIBP, EtCO₂, Temp)

Quantity: 8

## 3.1 Clinical & Functional Specifications

Post-anesthesia care unit multi-parameter bedside patient monitor for continuous monitoring of recovering post-surgical adult/pediatric/neonatal patients. Continuous simultaneous monitoring of ECG (3/5-lead), HR, respiration, SpO₂, PR, NIBP, side-stream EtCO₂, and skin temperature. Category-based defaults, basic arrhythmia and ST-segment analysis, drug-dose calculator, trend storage ≥ 72 hours, alarm/event review, HL7-compatible EMR interfacing, and networking to central station. Suitable for use on rolling mobile stand or wall-mounted configuration in PACU bays.

## 3.2 Engineering Technical Specifications

Display: Color TFT LCD ≥ 12 inch; resolution ≥ 1024 × 768; touchscreen with hard-key backup; ≥ 6 waveforms simultaneous.

ECG: 3/5-lead; bandwidth 0.05—150 Hz; sampling ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; ST −2.0 to +2.0 mV; defib-proof; ESU protected.

Respiration: 0—150 rpm ±2 rpm; apnea 10—40 s.

SpO₂: 0—100%, ±2% (70—100%); motion-tolerant.

NIBP: Oscillometric adult/ped/neo; per IEC 80601-2-30; accuracy ±3 mmHg.

EtCO₂: Side-stream; 0—99 mmHg; ±2 mmHg (0—40), ±5% above; flow 50—200 mL/min selectable.

Temperature: 1 channel (expandable); 0—50°C; ±0.1°C; YSI-400.

Alarms: Three-level per IEC 60601-1-8; nurse call.

Power: 100—240 VAC 50/60 Hz; battery ≥ 2 h; charge ≤ 6 h.

Connectivity: USB ×2, Ethernet, Wi-Fi 802.11 a/b/g/n, HL7.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 5—8 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1 series, IEC 60601-2-27, IEC 60601-2-49, IEC 80601-2-30, ISO 80601-2-55, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 3.3 Standard Accessories

- 3/5-lead ECG Patient Cable and Lead Wire Set — Defib-protected, AAMI/IEC color-coded, latex-free, ≥ 2.5 m, reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Total min 3.0 m; low-level disinfectant compatible.

- Reusable Adult NIBP Cuff with Air Hose — Standard adult 25—35 cm; latex-free; min 3.0 m hose.

- Side-Stream EtCO₂ Module with Sampling Lines — Water trap, airway adapters, starter sampling line kit.

- Reusable Skin Temperature Probe — YSI-400 compatible; ≥ 3.0 m cable.

- Integrated Rechargeable Battery Pack — ≥ 2 h; ≥ 500 cycles.

- Rolling Mobile Stand with Basket or Wall Mount Kit — Antimicrobial-coated or stainless steel; 5-caster base with min 2 locking casters; adjustable height; integrated accessory basket; wall mount option VESA compatible.

## 3.4 Installation Requirements

Medical-grade 100—240 VAC outlets in each PACU bay; UPS preferred; equipotential grounding; network drop or Wi-Fi; ambient 18—27°C; RH 30—70%.

## 3.5 Documentation

Complete set: user/service manuals, parts catalogue, factory test and calibration certificates, IEC 62353 electrical safety test report, EMC compliance, PM schedule, installation and commissioning reports.

## 3.6 Training

Clinical, BME, PM, and troubleshooting training with certificates.

## 3.7 Warranty

Minimum 2 years comprehensive as above.

# 4. Multipara Patient Monitor — Resuscitation Bay

Quantity: 10

## 4.1 Clinical & Functional Specifications

Rugged emergency department resuscitation-bay multi-parameter monitor for rapid deployment on critically ill/trauma patients across all age categories. Continuous simultaneous monitoring of ECG (3/5-lead), HR, respiration, SpO₂, PR, NIBP, side-stream EtCO₂ (with support for intubated and non-intubated capnography), and temperature. Rapid startup (typically ≤ 30 seconds to monitoring), one-touch/quick-access ED workflow, defibrillator sync capability, category defaults, basic arrhythmia and ST-segment analysis, trend storage ≥ 72 hours, event/alarm review, HL7 EMR interfacing, and networking with central station. Ruggedized construction for high-throughput ED environment; swivel wall mount minimizes floor footprint.

## 4.2 Engineering Technical Specifications

Display: ≥ 12-inch color TFT LCD; ≥ 1024 × 768; touchscreen + hard-key; ≥ 6 waveforms; high-brightness anti-glare.

ECG: 3/5-lead; 0.05—150 Hz; ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; ST analysis; defib-proof; ESU protected; defibrillator sync output.

Respiration: 0—150 rpm ±2 rpm; apnea 10—40 s.

SpO₂: 0—100%, ±2% (70—100%); motion-tolerant, low-perfusion capable.

NIBP: Oscillometric; adult/ped/neo; per IEC 80601-2-30; ±3 mmHg; STAT mode.

EtCO₂: Side-stream; 0—99 mmHg; ±2 mmHg (0—40); flow 50—200 mL/min; supports intubated and non-intubated (nasal cannula) capnography.

Temperature: 1 channel (expandable); 0—50°C; ±0.1°C.

Alarms: Three-level per IEC 60601-1-8; visual + audible; nurse call output.

Power: 100—240 VAC 50/60 Hz; heavy-duty rechargeable Li-ion; ≥ 3 hours runtime under typical ED load; charge ≤ 6 h.

Connectivity: USB ×2, Ethernet, Wi-Fi 802.11 a/b/g/n, HL7, defib sync, nurse call.

Environmental: 0—40°C operating; RH 15—95% NC; shock/vibration tolerant for cart-based movement.

Physical: Ruggedized housing; ≈ 6—9 kg; ingress protection minimum IPX2 (IPX4 or better preferred for splash resistance).

Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-8, IEC 60601-2-27, IEC 60601-2-49, IEC 80601-2-30, ISO 80601-2-55, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 4.3 Standard Accessories

- 3/5-lead ECG Patient Cable and Lead Wire Set — Defib-protected, AAMI/IEC color-coded, latex-free, ≥ 2.5 m, reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; total ≥ 3.0 m; disinfectant compatible.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard (25—35 cm); latex-free; ≥ 3.0 m hose.

- Side-Stream EtCO₂ Sensor Module with Sampling Lines — Water trap, adult/pediatric adapters, intubated and non-intubated sampling line starter kit.

- Reusable Skin Temperature Probe — YSI-400 compatible; ≥ 3.0 m cable; chemical disinfection compatible.

- Heavy-Duty Integrated Rechargeable Battery — ≥ 3 h runtime; ≥ 500 cycles at 80% capacity; over-current/thermal protection.

- Swivel Wall Mount Bracket with Integrated Cable Management Basket — Adjustable swivel/tilt arm; stainless steel or powder-coated aluminum; load-rated for monitor plus modules; cable management basket integrated.

## 4.4 Installation Requirements

Medical-grade 100—240 VAC outlet at each resus bay; UPS backup strongly recommended; equipotential grounding per IEC 60601-1; network drop and Wi-Fi coverage; ambient 18—27°C; RH 30—70%.

## 4.5 Documentation

Full documentation set: user/service manuals, parts catalogue, factory test/calibration certificates, IEC 62353 report, EMC compliance, PM schedule, installation/commissioning reports.

## 4.6 Training

Clinical (emergency staff), BME, PM, troubleshooting training with certificates.

## 4.7 Warranty

Minimum 2 years comprehensive: parts, labor, PM, breakdown, software updates, technical support.

# 5. OT-Grade Multipara Patient Monitor

Quantity: 4

## 5.1 Clinical & Functional Specifications

Operating-theater-grade multi-parameter bedside patient monitor for continuous intraoperative monitoring of adult, pediatric, and neonatal patients. Continuous simultaneous monitoring of ECG (3/5-lead with ESU-shielded acquisition), HR with pacemaker detection, respiration, SpO₂, PR, NIBP, EtCO₂ (side-stream and/or main-stream capable), and temperature; expandable for IBP, anesthetic gas analysis, and neuromuscular transmission monitoring. Advanced arrhythmia and ST-segment analysis, drug-dose calculator, hemodynamic/ventilation/oxygenation calculation packages, trend storage ≥ 96 hours, event/alarm review, HL7 EMR and anesthesia information management system (AIMS) interoperability, and networking to central monitoring station. Designed to withstand and reject electrosurgical unit interference.

## 5.2 Engineering Technical Specifications

Display: ≥ 15-inch color TFT LCD; ≥ 1280 × 800; capacitive touchscreen + hard-key; ≥ 8 waveforms; anti-glare.

ECG: 3/5/12-lead selectable/upgradable; bandwidth 0.05—150 Hz; sampling ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; ST −2.0 to +2.0 mV, ±0.02 mV; arrhythmia detection; ESU protection compliant with IEC 60601-2-27; defibrillator-proof; sync output; pacemaker detection.

Respiration: 0—150 rpm ±2 rpm; apnea 10—40 s.

SpO₂: 0—100%, ±2% (70—100%); motion-tolerant, low-perfusion capable.

NIBP: Oscillometric adult/ped/neo; per IEC 80601-2-30; ±3 mmHg; STAT and continuous modes.

EtCO₂: Side-stream and/or main-stream; 0—99 mmHg; ±2 mmHg (0—40), ±5% above; flow 50—200 mL/min; automatic baro-compensation; supports low-flow anesthesia.

Temperature: ≥ 2 channels; 0—50°C; ±0.1°C; YSI-400.

IBP (Expandability): Support for minimum 2 IBP channels via modular upgrade.

Alarms: Three-level per IEC 60601-1-8; visual/audible; nurse-call/OT signal output.

Power: 100—240 VAC 50/60 Hz; internal Li-ion ≥ 2 h; charge ≤ 6 h.

Connectivity: USB ×2, Ethernet, Wi-Fi 802.11 a/b/g/n, HL7, defib sync, analog/serial anesthesia machine interface.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 6—10 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27 (with ESU compliance), 60601-2-49, IEC 80601-2-30, ISO 80601-2-55, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 5.3 Standard Accessories

- 5-lead ECG Patient Cable and Electrosurgical Interference-Shielded Lead Wire Set — High-density shielded trunk cable and lead wires specifically constructed for ESU interference suppression; defib-protected; AAMI/IEC color-coded; latex-free; ≥ 2.5 m; reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Dual-wavelength; total ≥ 3.0 m; disinfectant compatible.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard (25—35 cm); latex-free; ≥ 3.0 m hose.

- Side-Stream or Main-Stream EtCO₂ Sensor Kit — Complete assembly with water trap (side-stream) and/or airway adapter (main-stream); starter sampling line pack for adult/pediatric anesthesia circuits.

- Reusable Temperature Probe — YSI-400 compatible; skin and/or esophageal type; ≥ 3.0 m cable.

- Integrated Rechargeable Battery — ≥ 2 h; ≥ 500 cycles; protected BMS.

- Anesthesia Machine Mounting Arm or Rolling Mobile Stand — Articulating arm compatible with standard anesthesia machine mounting rails, or 5-caster antimicrobial mobile stand with locking casters and accessory basket.

## 5.4 Installation Requirements

Medical-grade 100—240 VAC 50/60 Hz outlets in each OT; isolated power system or UPS strongly recommended; equipotential grounding per IEC 60601-1; network drop (Cat 6) for AIMS/central station; ambient OT conditions typically 18—24°C, RH 40—60%; EMI/ESU-tolerant installation.

## 5.5 Documentation

Full set: user/service manuals, parts catalogue, factory test/calibration certificates, IEC 62353 report, EMC compliance including ESU testing, PM schedule, installation/commissioning reports.

## 5.6 Training

Clinical (anesthesia and OT staff), BME, PM, troubleshooting training with certificates.

## 5.7 Warranty

Minimum 2 years comprehensive: parts, labor, PM, breakdown, software updates, technical support.

# 6. Ward-Grade Basic Multipara Patient Monitor

Quantity: 66

## 6.1 Clinical & Functional Specifications

Basic multi-parameter bedside patient monitor for general ward/inpatient monitoring of stable adult and pediatric patients. Continuous simultaneous monitoring of ECG (3-lead), HR, respiration by impedance, SpO₂, PR, NIBP, and skin temperature. Basic arrhythmia detection, three-level alarms, tabular/graphical trend storage ≥ 48 hours, event/alarm review, and networking to central station with HL7 EMR interfacing. Simple, intuitive user interface for ward nursing staff.

## 6.2 Engineering Technical Specifications

Display: ≥ 10-inch color TFT LCD; ≥ 800 × 600 (1024 × 768 preferred); touchscreen and/or hard-key; ≥ 4 waveforms.

ECG: 3-lead (5-lead upgradable optional); 0.05—150 Hz; ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; basic arrhythmia; defib-proof.

Respiration: 0—150 rpm ±2 rpm; apnea 10—40 s.

SpO₂: 0—100%, ±2% (70—100%).

NIBP: Oscillometric adult/ped; per IEC 80601-2-30; ±3 mmHg; manual/auto modes.

Temperature: 1 channel; 0—50°C; ±0.1°C; YSI-400.

Alarms: Three-level per IEC 60601-1-8; visual/audible; nurse-call output.

Power: 100—240 VAC 50/60 Hz; Li-ion battery ≥ 2 h; charge ≤ 6 h.

Connectivity: USB ×1 minimum, Ethernet, Wi-Fi (optional 802.11 a/b/g/n), HL7-capable.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 3—6 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27, 60601-2-49, IEC 80601-2-30, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 6.3 Standard Accessories

- 3-lead ECG Patient Cable and Lead Wire Set — Defib-protected; AAMI/IEC color-coded; latex-free; ≥ 2.5 m; reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; total ≥ 3.0 m.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard 25—35 cm; latex-free; ≥ 3.0 m hose.

- Reusable Temperature Probe — YSI-400 compatible; skin type; ≥ 3.0 m cable.

- Integrated Rechargeable Battery Pack — ≥ 2 h; ≥ 500 cycles.

- Bedside Rail or Wall Mounting Kit — Adjustable bracket compatible with 25/38 mm hospital rails or VESA wall mount.

## 6.4 Installation Requirements

Medical-grade 100—240 VAC outlets per bed; equipotential grounding; network drop or Wi-Fi coverage; ambient 18—27°C; RH 30—70%.

## 6.5 Documentation

Complete standard set: user/service manuals, parts catalogue, factory test/calibration certificates, IEC 62353 report, EMC compliance, PM schedule, installation/commissioning reports.

## 6.6 Training

Clinical, BME, PM, and troubleshooting training with certificates.

## 6.7 Warranty

Minimum 2 years comprehensive.

# 7. Patient Monitor — Basic Multi-Parameter (OBS)

Quantity: 2

## 7.1 Clinical & Functional Specifications

Portable basic multi-parameter patient monitor for obstetrics/observation with continuous simultaneous monitoring of ECG (3-lead), HR, respiration, SpO₂, PR, NIBP, and skin temperature on adult and neonatal categories. Category defaults, three-level alarms, trend storage ≥ 48 hours, event/alarm review, HL7 EMR interfacing, and networking to central station. Mounted on rolling mobile stand for flexible use across OBS areas.

## 7.2 Engineering Technical Specifications

Display: ≥ 10-inch color TFT LCD; ≥ 800 × 600 (1024 × 768 preferred); touchscreen; ≥ 4 waveforms.

ECG: 3-lead; 0.05—150 Hz; ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; defib-proof.

Respiration: 0—150 rpm ±2 rpm; apnea alarm.

SpO₂: 0—100%, ±2% (70—100%); ±3% neonatal.

NIBP: Oscillometric adult/ped/neo; per IEC 80601-2-30; ±3 mmHg.

Temperature: 1 channel; 0—50°C; ±0.1°C; YSI-400.

Alarms: Three-level per IEC 60601-1-8.

Power: 100—240 VAC 50/60 Hz; Li-ion ≥ 2 h; charge ≤ 6 h.

Connectivity: USB, Ethernet, Wi-Fi (optional), HL7.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 3—6 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27, 60601-2-49, IEC 80601-2-30, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 7.3 Standard Accessories

- 3-lead ECG Patient Cable and Lead Wire Set — Defib-protected; AAMI/IEC color-coded; latex-free; ≥ 2.5 m; reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; total ≥ 3.0 m.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard 25—35 cm; latex-free; ≥ 3.0 m hose.

- Reusable Temperature Probe — YSI-400 compatible; ≥ 3.0 m cable.

- Integrated Rechargeable Battery — ≥ 2 h; ≥ 500 cycles.

- Rolling Mobile Stand with Storage Basket — 5-caster antimicrobial-finished base; ≥ 2 locking casters; adjustable height; accessory basket.

## 7.4 Installation Requirements

Medical-grade 100—240 VAC outlets; equipotential grounding; network drop or Wi-Fi; ambient 18—27°C; RH 30—70%.

## 7.5 Documentation

Complete standard set as above.

## 7.6 Training

Clinical, BME, PM, troubleshooting training with certificates.

## 7.7 Warranty

Minimum 2 years comprehensive.

# 8. Cardiac Monitor (Seven Parameters)

Quantity: 28

## 8.1 Clinical & Functional Specifications

Cardiac-focused seven-parameter bedside monitor for cardiology, CCU, and cardiac step-down units. Continuous simultaneous monitoring of ECG (3/5-lead), HR, respiration by impedance, SpO₂, PR, NIBP, temperature, dual IBP, and side-stream EtCO₂. Advanced arrhythmia detection (minimum lethal plus multiple non-lethal categories), ST-segment analysis on all available leads, QT/QTc interval analysis, pacemaker detection and rejection, hemodynamic calculation package, trend storage ≥ 96 hours, event/alarm review, HL7 EMR interfacing, and networking to central station.

## 8.2 Engineering Technical Specifications

Display: ≥ 12-inch color TFT LCD; ≥ 1024 × 768; touchscreen + hard-key; ≥ 6 waveforms.

ECG: 3/5-lead (12-lead upgradable); bandwidth 0.05—150 Hz; ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; ST −2.0 to +2.0 mV ±0.02 mV; advanced arrhythmia; QT/QTc; defib-proof; ESU protected.

Respiration: 0—150 rpm ±2 rpm; apnea 10—40 s.

SpO₂: 0—100%, ±2% (70—100%); motion-tolerant.

NIBP: Oscillometric adult/ped/neo; per IEC 80601-2-30; ±3 mmHg; manual/auto/STAT.

Dual IBP: 2 channels; range −50 to +300 mmHg; ±2% or ±1 mmHg; 5 μV/V/mmHg transducer compatible.

EtCO₂: Side-stream; 0—99 mmHg; ±2 mmHg (0—40), ±5% above; flow 50—200 mL/min.

Temperature: 1 channel (2 preferred); 0—50°C; ±0.1°C; YSI-400.

Alarms: Three-level per IEC 60601-1-8; nurse call.

Power: 100—240 VAC 50/60 Hz; Li-ion ≥ 2 h; charge ≤ 6 h.

Connectivity: USB ×2, Ethernet, Wi-Fi 802.11 a/b/g/n, HL7, defib sync.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 5—9 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27, 60601-2-34 (IBP), 60601-2-49, IEC 80601-2-30, ISO 80601-2-55, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 8.3 Standard Accessories

- 5-lead ECG Patient Cable and Lead Wire Set — Defib-protected; AAMI/IEC color-coded; latex-free; ≥ 2.5 m; reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; total ≥ 3.0 m.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard 25—35 cm; latex-free; ≥ 3.0 m hose.

- Reusable Temperature Probe — YSI-400 compatible; ≥ 3.0 m cable.

- Dual IBP Interface Cables — Two shielded reusable cables; 5 μV/V/mmHg compatible; defib-protected; ≥ 3.0 m.

- Side-Stream EtCO₂ Sampling Kit — Water trap, airway adapters, starter disposable sampling lines.

- Integrated Rechargeable Battery — ≥ 2 h; ≥ 500 cycles.

- Bedside Rail Mounting Bracket — Adjustable stainless steel/aluminum, 25/38 mm rail compatible.

## 8.4 Installation Requirements

Medical-grade 100—240 VAC outlets; UPS backup preferred; equipotential grounding; Cat 6 network drop or Wi-Fi; ambient 18—27°C; RH 30—70%.

## 8.5 Documentation

Complete standard set including IBP-related calibration and EMC test reports.

## 8.6 Training

Clinical (cardiology), BME, PM, troubleshooting training with certificates.

## 8.7 Warranty

Minimum 2 years comprehensive.

# 9. Cardiac Monitor (Five Parameters)

Quantity: 2

## 9.1 Clinical & Functional Specifications

Five-parameter cardiac monitor for continuous simultaneous monitoring of ECG (3/5-lead), HR, respiration by impedance, SpO₂, PR, NIBP, and temperature. Basic arrhythmia and ST-segment analysis, pacemaker detection, three-level alarms, trend storage ≥ 72 hours, event/alarm review, HL7 EMR interfacing, and networking to central station.

## 9.2 Engineering Technical Specifications

Display: ≥ 10-inch color TFT LCD; ≥ 1024 × 768; touchscreen and/or hard-key; ≥ 4 waveforms.

ECG: 3/5-lead; 0.05—150 Hz; ≥ 500 samples/s/channel; HR 15—300 bpm ±1 bpm/±1%; ST analysis; basic arrhythmia; defib-proof.

Respiration: 0—150 rpm ±2 rpm.

SpO₂: 0—100%, ±2% (70—100%).

NIBP: Oscillometric adult/ped/neo; per IEC 80601-2-30; ±3 mmHg.

Temperature: 1 channel; 0—50°C; ±0.1°C; YSI-400.

Alarms: Three-level per IEC 60601-1-8.

Power: 100—240 VAC 50/60 Hz; Li-ion ≥ 2 h; charge ≤ 6 h.

Connectivity: USB, Ethernet, Wi-Fi (optional), HL7.

Environmental: 5—40°C; RH 15—95% NC.

Physical: ≈ 4—7 kg; IPX1 minimum.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27, 60601-2-49, IEC 80601-2-30, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 9.3 Standard Accessories

- 3/5-lead ECG Patient Cable and Lead Wire Set — Defib-protected; AAMI/IEC color-coded; latex-free; ≥ 2.5 m; reusable.

- Reusable Adult SpO₂ Finger Sensor with Extension Cable — Standard dual-wavelength; total ≥ 3.0 m.

- Reusable Adult NIBP Cuff with Air Hose — Adult standard 25—35 cm; latex-free; ≥ 3.0 m hose.

- Reusable Temperature Probe — YSI-400 compatible; ≥ 3.0 m cable.

- Integrated Rechargeable Battery Pack — ≥ 2 h; ≥ 500 cycles.

- Bedside Rail Mounting Bracket — Adjustable; 25/38 mm rail compatible.

## 9.4 Installation Requirements

Medical-grade 100—240 VAC outlets; equipotential grounding; network drop or Wi-Fi; ambient 18—27°C; RH 30—70%.

## 9.5 Documentation

Complete standard set.

## 9.6 Training

Clinical, BME, PM, troubleshooting training with certificates.

## 9.7 Warranty

Minimum 2 years comprehensive.

# 10. Hemodynamic Monitor with Cardiac Output

Quantity: 2

## 10.1 Clinical & Functional Specifications

Advanced hemodynamic monitor for continuous and intermittent measurement of cardiac output and derived hemodynamic parameters in critically ill and complex cardiac patients. Shall support thermodilution technique (bolus and/or continuous), and, where offered, additional less-invasive techniques (pulse contour analysis or equivalent). Continuous simultaneous monitoring of ECG, HR, SpO₂, minimum 2 IBP channels, temperature, and cardiac output with automatic calculation of derived parameters: cardiac index (CI), stroke volume (SV), stroke volume index (SVI), systemic vascular resistance (SVR), SVRI, pulmonary vascular resistance (PVR), left/right cardiac work indices, oxygen delivery (DO₂), oxygen consumption (VO₂), and extraction ratio. Trend storage ≥ 96 hours; hemodynamic calculation package; HL7 EMR interfacing; touchscreen operation.

## 10.2 Engineering Technical Specifications

Display: ≥ 15-inch high-resolution color TFT LCD; ≥ 1280 × 800; capacitive multi-touch touchscreen; anti-glare.

Cardiac Output (Thermodilution): Range 0.1—20 L/min; accuracy ±5% or ±0.2 L/min; injectate temperature range 0—27°C; automatic averaging of multiple bolus measurements.

IBP: ≥ 2 channels (4 preferred); range −50 to +300 mmHg; ±2% or ±1 mmHg; 5 μV/V/mmHg transducer compatible.

ECG: 3/5-lead; 0.05—150 Hz; HR 15—300 bpm ±1 bpm/±1%; defib-proof.

SpO₂: 0—100%, ±2% (70—100%).

Temperature: ≥ 2 channels; 0—50°C; ±0.1°C.

Derived Hemodynamic Parameters: CI, SV, SVI, SVR, SVRI, PVR, LCW, RCW, LVSW, RVSW, DO₂, VO₂, O₂ER, and configurable additional parameters.

Alarms: Three-level per IEC 60601-1-8.

Power: 100—240 VAC 50/60 Hz; Li-ion ≥ 2 h; charge ≤ 6 h.

Connectivity: USB ×2, Ethernet, Wi-Fi 802.11 a/b/g/n, HL7-capable.

Environmental: 5—40°C; RH 15—95% NC.

Physical: Mobile stand-mounted; monitor unit ≈ 5—10 kg.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, 60601-2-27, 60601-2-34, 60601-2-49, IEC 80601-2-30, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 10.3 Standard Accessories

- Cardiac Output (CO) Interface Cable — Shielded reusable cable compatible with standard thermodilution catheters; defib-protected; ≥ 3.0 m.

- Thermistor Sensor Cable and Injectate Temperature Probe Kit — Complete kit with in-line injectate temperature sensor for room-temperature and cold injectate operation; reusable cables ≥ 3.0 m.

- Dual IBP Interface Cables — 2 shielded reusable cables; 5 μV/V/mmHg compatible; ≥ 3.0 m.

- High-Resolution Display Monitor with Touchscreen Control — Integrated as per system architecture; ≥ 15-inch; ≥ 1280 × 800; capacitive multi-touch.

- Integrated Rechargeable Battery Pack — ≥ 2 h runtime; ≥ 500 cycles.

- Mobile Roll Stand with Accessory Basket — 5-caster antimicrobial finish base; ≥ 2 locking casters; adjustable height; integrated basket; load-rated for monitor with modules.

## 10.4 Installation Requirements

Medical-grade 100—240 VAC outlets; UPS backup recommended; equipotential grounding per IEC 60601-1; network drop for HL7 EMR interfacing; ambient 18—27°C; RH 30—70%.

## 10.5 Documentation

Complete standard set with CO calibration documentation.

## 10.6 Training

Clinical (ICU/cardiology), BME, PM, troubleshooting training with certificates.

## 10.7 Warranty

Minimum 2 years comprehensive.

# 11. Central Monitoring Station

Quantity: 7

## 11.1 Clinical & Functional Specifications

Networked central monitoring station for continuous simultaneous surveillance of all connected bedside patient monitors within an assigned unit (ICU, CCU, HDU, wards, etc.). Shall display real-time waveforms and numeric parameters of all connected beds, provide centralized alarm management with priority classification, allow full-disclosure waveform review, trend and event review, arrhythmia and ST review, printed and electronic reporting, and bidirectional communication with bedside monitors and hospital EMR via HL7. Shall support at least the full bed capacity of the connected unit (bed licenses to match tender-specified capacity).

## 11.2 Engineering Technical Specifications

Workstation: Medical/industrial-grade tower workstation; multi-core processor (equivalent to modern quad-core or better); minimum 16 GB RAM; minimum 500 GB SSD storage; redundant storage or backup mechanism preferred; medical-grade or hospital-approved operating system with security patches supported.

Displays: Dual 24-inch (or larger) medical-grade LED/LCD monitors; minimum resolution 1920 × 1080 (Full HD) each; ≥ 250 cd/m² brightness; wide viewing angle ≥ 170°; anti-glare surface.

Software Capacity: License for full bed capacity of the connected unit; capable of simultaneous display of all connected beds with real-time waveforms and numeric parameters; alarm history; full-disclosure waveform storage minimum 72 hours per bed; trend storage minimum 168 hours per bed; ST and arrhythmia review; printable reports.

Networking: Wired Gigabit Ethernet; wireless via router or access point per site requirement; HL7 EMR gateway; DICOM-optional.

Alarms: Centralized alarm management with visual and audible indication; three-level priority per IEC 60601-1-8; alarm forwarding to nurse station and mobile devices (optional).

Peripherals: Alphanumeric keyboard, optical mouse, laser report printer, UPS with surge protection.

Power: 100—240 VAC 50/60 Hz; connection to hospital UPS.

Environmental: Nurse station indoor conditions 5—40°C; RH 15—95% NC.

Physical: Desktop workstation with dual-monitor arm/stand assembly.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, IEC 62304 (medical software), IEC 80001-1 (medical IT network risk management), IEC 62366, ISO 13485, ISO 14971, HL7, CE MDR/FDA. Cybersecurity per FDA/IMDRF guidance and IEC 80001-1.

## 11.3 Standard Accessories

- High-Performance Central Processing Workstation Tower — Multi-core processor; ≥ 16 GB RAM; ≥ 500 GB SSD; redundant power supply preferred; medical/hospital-approved OS with vendor-supported cybersecurity updates.

- Dual 24-inch Medical-Grade High-Resolution Display Monitors — LED/LCD; ≥ 1920 × 1080 each; ≥ 250 cd/m² brightness; anti-glare; wide viewing angle; adjustable tilt/swivel; VESA compatible.

- Wireless LAN Router or Network Switch for Patient Monitor Telemetry Connection — Enterprise-grade managed switch/router; minimum 24-port Gigabit Ethernet for network switch; IEEE 802.11 a/b/g/n/ac for wireless; VLAN and QoS capable; WPA2/WPA3 security.

- Central Station Monitoring Software Licenses for Full Bed Capacity — Perpetual or subscription license for the full bed count required per station; includes real-time monitoring, alarm management, trend and full-disclosure review, ST/arrhythmia review, HL7 gateway, and report generation.

- Alphanumeric Keyboard and Optical Mouse — Standard USB-connected medical/wipe-clean compatible keyboard and mouse; latex-free.

- Laser Printer for Alarm and Waveform Reporting — A4-capable monochrome laser printer; minimum 20 ppm; network-capable; automatic duplex preferred.

- Uninterruptible Power Supply (UPS) with Surge Protection — Online/line-interactive UPS sized for workstation, monitors, and printer; minimum 30 minutes runtime under full load; surge and overvoltage protection; audible/visual status indication.

## 11.4 Installation Requirements

Dedicated 100—240 VAC 50/60 Hz medical-grade outlets at each central station; UPS backup mandatory; structured network cabling (Cat 6/6A minimum) between central station and each bedside monitor; dedicated VLAN for patient monitoring network; equipotential grounding; ambient 18—27°C, RH 30—70%; adequate desk space with dual-monitor mounting.

## 11.5 Documentation

User and administrator manuals, system architecture documentation, network configuration documentation, software installation media and license keys, cybersecurity risk assessment, service manuals, parts catalogue, electrical safety and EMC test reports, PM schedule, installation and commissioning reports.

## 11.6 Training

Nursing/clinical user training, IT/network administrator training, BME/PM training, cybersecurity awareness, and application training with certificates.

## 11.7 Warranty

Minimum 2 years comprehensive including software updates and cybersecurity patches, hardware parts, labor, PM, breakdown maintenance, and 24/7 technical support.

# 12. BIS Monitor

Quantity: 1

## 12.1 Clinical & Functional Specifications

Bispectral Index (BIS) or equivalent processed-EEG depth-of-anesthesia monitor for real-time monitoring of the hypnotic component of anesthesia in adult and pediatric patients. Shall provide a numeric index (0—100 scale or equivalent) indicating level of consciousness, along with EEG raw waveform, suppression ratio, signal quality index, and electromyographic (EMG) activity indicator. Trend display for the duration of anesthesia case and case archiving. Standalone bedside or anesthesia-machine-mounted operation.

## 12.2 Engineering Technical Specifications

Display: ≥ 5-inch color LCD/TFT (integrated in module) or displayed on host monitor; capable of showing numeric depth-of-hypnosis index, EEG waveform, suppression ratio, signal quality, and EMG bar.

Index Output: Numeric processed-EEG index on 0—100 scale (or equivalent internationally recognized scale); update rate typically ≤ 5 seconds.

EEG: Single or dual channel; frequency response typically 0.5—70 Hz; sampling ≥ 128 samples/s/channel; sensor impedance monitoring with real-time indication.

Suppression Ratio: 0—100%.

EMG Indicator: dB scale with real-time bar/graph.

Signal Quality Index: 0—100%, real-time.

Data Storage: Case-level trend storage of numeric index and processed parameters for duration of case; case archive minimum 20 cases or export via USB.

Alarms: Three-level per IEC 60601-1-8.

Power: 100—240 VAC 50/60 Hz via AC adapter; low-voltage DC operation from module; internal battery optional.

Connectivity: Serial/USB output for host monitor integration; capable of forwarding data to compatible multi-parameter monitors.

Environmental: 5—40°C; RH 15—95% NC.

Physical: Module ≈ 1—3 kg; portable/pole-mountable.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, IEC 60601-2-26 (electroencephalographs), IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 12.3 Standard Accessories

- Bispectral Index (BIS) Patient Interface Cable (PIC-Plus or Equivalent) — Reusable shielded patient interface cable with impedance-check capability, compatible with standard disposable BIS sensors; defib-protected; ≥ 2.5 m.

- Reusable Monitor Processing Module — Digital signal-processing module housing the processed-EEG algorithm; connects between patient interface cable and host display; latex-free wipe-clean housing.

- Rolling Mobile Stand with Accessory Basket or Pole Mounting Clamp — Antimicrobial finish; ≥ 2 locking casters (if mobile stand) or standard IV-pole clamp with quick-release mechanism.

- AC Power Adapter and Main Power Cord — Medical-grade external AC/DC adapter, 100—240 VAC 50/60 Hz input; region-appropriate hospital-grade mains cord.

- Test Sensor/Simulator for System Diagnostic Check — Reusable test/simulator plug enabling verification of module and cable integrity per manufacturer diagnostic procedure.

## 12.4 Installation Requirements

Medical-grade 100—240 VAC outlet; equipotential grounding; ambient OT conditions 18—24°C; RH 40—60%; connection to compatible multi-parameter monitor (if integrated).

## 12.5 Documentation

Full documentation set including sensor placement and interpretation guide.

## 12.6 Training

Clinical (anesthesia), BME, PM, and troubleshooting training with certificates.

## 12.7 Warranty

Minimum 2 years comprehensive.

# 13. Neuro Monitoring (IONM)

Quantity: 1

## 13.1 Clinical & Functional Specifications

Multichannel intraoperative neurophysiological monitoring (IONM) system for real-time monitoring of neural function during high-risk neurosurgical, orthopedic (spine), ENT, vascular, and cardiothoracic procedures. Shall support the full range of standard IONM modalities: somatosensory evoked potentials (SSEP), motor evoked potentials (MEP) with transcranial electrical stimulation, brainstem auditory evoked potentials (BAEP), visual evoked potentials (VEP), free-run and triggered electromyography (EMG), electroencephalography (EEG), and nerve conduction studies. Real-time waveform acquisition, averaging, and display; automatic threshold monitoring and alarming; case archiving and reporting; and networked data export.

## 13.2 Engineering Technical Specifications

Amplifier: Minimum 16 channels (32 channels preferred and upgradable); differential input; input impedance ≥ 100 MΩ; CMRR ≥ 100 dB; input noise typically ≤ 1 μV RMS.

Frequency Response / Filters: Configurable high-pass 0.1—3000 Hz; low-pass 30—10000 Hz; notch 50/60 Hz selectable; bandwidth suitable for full range of modalities.

Sampling Rate: ≥ 10 000 samples/second/channel (or higher, matching modality requirements).

A/D Resolution: Minimum 16-bit (24-bit preferred).

Stimulator (Electrical): Constant-current output; current range 0—100 mA (SSEP/EMG), high-voltage/high-current transcranial capability for MEP up to internationally standard levels; pulse duration adjustable typically 50—1000 μs; pulse rate configurable 0.1—500 Hz; single, train, and paired-pulse modes.

Modalities Supported: SSEP, MEP (TES), BAEP, VEP, spontaneous and triggered EMG, EEG, and nerve conduction studies.

Display: ≥ 15-inch high-resolution color LCD on medical-grade laptop/workstation; ≥ 1920 × 1080; simultaneous multi-modality waveform display.

Data Storage: Solid-state storage ≥ 500 GB; case archive with full raw and averaged waveforms; export in standard formats (EDF, DICOM, PDF).

Audio Feedback: Real-time EMG audio output via speakers/headphones.

Alarms: Configurable automatic threshold-based warnings on latency and amplitude changes.

Isolation & Safety: Patient isolation compliant with IEC 60601-1; defibrillator-proof inputs where clinically appropriate; medical isolation transformer for full system.

Power: 100—240 VAC 50/60 Hz via isolation transformer; internal UPS or laptop battery ≥ 1 h.

Connectivity: USB, Ethernet, wireless optional; DICOM/HL7 export.

Environmental: 5—40°C; RH 15—95% NC.

Physical: Portable cart-based or rigid-case transportable configuration.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, IEC 60601-2-26 (EEG), IEC 60601-2-40 (electromyographs and evoked response equipment), IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 13.3 Standard Accessories

- Multichannel Intraoperative Neuromonitoring Amplifier Unit — Isolated differential amplifier ≥ 16 channels; input impedance ≥ 100 MΩ; CMRR ≥ 100 dB; 16-bit or better A/D; defib-protected inputs.

- Subdermal Needle Electrode Starter Kit — Sterile single-use stainless-steel subdermal needle electrodes with color-coded insulated lead wires; assorted lengths and configurations (paired, twisted-pair) sufficient for a starter case load.

- Reusable Stimulating and Recording Cables — Shielded reusable cables and headboxes for stimulating and recording channels; defib-protected; latex-free; wipe-clean/sterilization compatible per manufacturer.

- Constant Current Stimulator Probe Handle with Interchangeable Tips — Sterile-compatible/autoclavable or high-level-disinfection-compatible handheld stimulator probe with interchangeable monopolar and bipolar tips.

- High-Performance Laptop/Workstation Controller with Medical-Grade Screen — Medical/industrial-grade portable computer; multi-core processor; ≥ 16 GB RAM; ≥ 500 GB SSD; ≥ 15-inch anti-glare medical-grade display; USB, Ethernet, wireless.

- Audio Feedback Speakers or Headphones — Powered speakers and/or high-quality headphones for real-time EMG audio; medical-environment compatible.

- Isolation Transformer and Main Power Assembly — Medical-grade isolation transformer compliant with IEC 60601-1; VA rating sized for the full IONM system; equipotential earthing terminal.

- Rigid Transport Case — Impact-resistant wheeled case with custom foam interior for safe transport and storage of the complete IONM system.

## 13.4 Installation Requirements

Medical-grade 100—240 VAC 50/60 Hz outlet(s) sized for isolation transformer load; equipotential grounding per IEC 60601-1; OT ambient conditions 18—24°C, RH 40—60%; EMI/ESU-tolerant setup; network drop for data export.

## 13.5 Documentation

Complete documentation set including modality-specific application manuals, safety limits for stimulation output, and case archive management.

## 13.6 Training

Clinical (neurophysiologist/surgeon/neuromonitorist), BME, PM, troubleshooting, and application training with certificates. Extended application training strongly recommended given specialty nature.

## 13.7 Warranty

Minimum 2 years comprehensive including software updates.

# 14. Pulse Oximeter (Handheld, Portable)

Quantity: 18

## 14.1 Clinical & Functional Specifications

Portable handheld pulse oximeter for spot-check and continuous measurement of functional oxygen saturation (SpO₂) and pulse rate on adult, pediatric, and (where sensor supports) neonatal patients. Rugged construction for ward, ambulatory, and transport use. Motion-tolerant, low-perfusion-capable measurement algorithm; three-level alarms with adjustable limits; trend memory; USB data export; rechargeable battery operation with desktop charging.

## 14.2 Engineering Technical Specifications

Display: ≥ 2.4-inch color LCD/OLED; sufficient resolution for numeric SpO₂, PR, plethysmograph, and signal quality indication; readable in typical clinical lighting.

SpO₂: Range 0—100%; accuracy ±2% (70—100%, adult/pediatric); ±3% (neonatal); resolution 1%; motion-tolerant, low-perfusion capable algorithm.

Pulse Rate: Range 25—250 bpm; accuracy ±2 bpm or ±2%, whichever greater.

Plethysmograph: Real-time waveform with auto/manual gain.

Signal Quality / Perfusion: Perfusion index display and/or signal quality indicator.

Trend Memory: ≥ 24 hours continuous trend storage or ≥ 100 spot-check records; USB export.

Alarms: Three-level visual/audible per IEC 60601-1-8; adjustable limits.

Power: Rechargeable NiMH or lithium battery; ≥ 8 hours continuous runtime under typical operation; charging via AC adapter or desktop cradle ≤ 6 hours; visual battery indicator.

Connectivity: USB for data export; optional Bluetooth.

Environmental: Operating 5—40°C; RH 15—95% NC; drop-resistant construction; ingress protection IP22 minimum (IPX4 preferred).

Physical: Handheld, weight approximately 0.3—0.6 kg with battery; ergonomic design.

Applicable Standards: IEC 60601-1, 60601-1-2, 60601-1-8, ISO 80601-2-61, IEC 62304, ISO 13485, ISO 14971, CE MDR/FDA.

## 14.3 Standard Accessories

- Reusable Adult SpO₂ Finger Probe — Soft-shell/silicone reusable adult finger clip; standard dual-wavelength (\~660/940 nm); latex-free; low-level disinfectant compatible; cable ≥ 1.0 m.

- Protective Rubber Boot/Cover with Integrated Table Stand — Impact-absorbing elastomer sleeve providing drop protection; integrated fold-out stand for tabletop use.

- Rechargeable NiMH or Lithium Battery Set — Rated for ≥ 8 hours continuous operation; ≥ 500 cycles at 80% capacity retention; over-charge/thermal protection.

- Desktop Charging Cradle with AC Power Adapter — Docking-style cradle with charge status LED; medical-grade AC/DC adapter, 100—240 VAC 50/60 Hz input.

- Carrying Lanyard or Wrist Strap — Nylon lanyard/strap with quick-release safety mechanism; latex-free.

## 14.4 Installation Requirements

100—240 VAC 50/60 Hz standard outlet for charging cradle; no special installation; storage/charging area with ambient 5—40°C.

## 14.5 Documentation

User manual, service manual, spare parts catalogue, factory test/calibration certificate, IEC 62353 electrical safety test report, EMC compliance, Declaration of Conformity, PM schedule.

## 14.6 Training

Clinical/user training and BME/PM familiarization with certificates.

## 14.7 Warranty

Minimum 2 years comprehensive.

# 15. Glucometer

Quantity: 12

## 15.1 Clinical & Functional Specifications

Portable handheld blood glucose meter for capillary whole-blood glucose measurement at point-of-care for hospital ward, ICU, ED, and outpatient use. Compatible with commercially available disposable test strips; internal quality-control and calibration verification; memory of previous readings with date/time; alarm/flag for hypo- and hyperglycemia; optional data export for EMR integration.

## 15.2 Engineering Technical Specifications

Measurement Principle: Electrochemical (glucose oxidase/dehydrogenase) or photometric, using disposable single-use test strips.

Measurement Range: Minimum 20—600 mg/dL (approximately 1.1—33.3 mmol/L).

Accuracy: Compliant with ISO 15197 latest edition (typically ≥ 95% of results within ±15 mg/dL for glucose \< 100 mg/dL and ±15% for glucose ≥ 100 mg/dL, when compared with laboratory reference).

Sample Volume: Typically ≤ 1.0 μL capillary whole blood.

Measurement Time: Typically ≤ 10 seconds per test.

Hematocrit Range: Typically 20—70% (per manufacturer specification).

Display: Backlit LCD with readable numeric result, date, time, and unit (mg/dL or mmol/L) selectable.

Memory: ≥ 500 test results with date/time.

Connectivity: USB or Bluetooth data export (optional) for EMR/data management.

Coding / Calibration: Auto-coding preferred, or single-code lot calibration; internal quality-control checks each test.

Power: Rechargeable or long-life coin-cell batteries; typical life ≥ 1000 tests per battery set.

Environmental: Operating 5—40°C; storage per manufacturer; RH 10—90% NC.

Physical: Handheld, weight approximately 50—150 g.

Applicable Standards: ISO 15197 (in vitro diagnostic test systems—blood glucose monitoring), IEC 61010 (safety for laboratory equipment) where applicable, IEC 60601-1-2 (EMC where relevant), ISO 13485, ISO 14971, CE IVDR/FDA.

## 15.3 Standard Accessories

- Rigid Plastic Carrying and Storage Case — Impact-resistant molded case with compartments for meter, lancing device, test strips, control solutions, and lancets; latex-free.

- Lancing Device — Reusable spring-loaded lancing pen with adjustable puncture depth (multiple settings); compatible with standard disposable single-use lancets; ergonomic design.

- Check Strip or Optical Verification Strip for Calibration — Reusable check/verification strip enabling routine functional verification of the meter per manufacturer's quality-control procedure.

- Rechargeable or Long-Life Coin-Cell Batteries — Manufacturer-specified rechargeable pack or coin-cell battery set rated for typical usage of ≥ 1000 tests; safety-protected.

## 15.4 Installation Requirements

No special installation required; storage area with ambient conditions per manufacturer specification (typically 5—30°C).

## 15.5 Documentation

User manual, quick reference guide, factory calibration certificate, ISO 15197 accuracy compliance statement, quality-control procedure, EMC declaration where applicable.

## 15.6 Training

Clinical/user training on operation, quality control, and cleaning; BME familiarization; certificates issued.

## 15.7 Warranty

Minimum 2 years comprehensive on meter; consumables (strips, lancets, control solution) excluded from warranty.