Hospital Medical Equipment — Technical Tender Specifications

Technical Tender Document

Table of Contents

1. ARTERIAL BLOOD GAS (ABG) ANALYZER	4

1.1 Clinical & Functional Specifications	4

1.2 Engineering Technical Specifications	4

1.3 Standard Accessories	7

1.4 Installation Requirements	8

1.5 Documentation	8

1.6 Training	8

1.7 Warranty	9

2. ACTIVATED CLOTTING TIME (ACT) MACHINE	10

2.1 Clinical & Functional Specifications	10

2.2 Engineering Technical Specifications	10

2.3 Standard Accessories	12

2.4 Installation Requirements	13

2.5 Documentation	13

2.6 Training	13

2.7 Warranty	13

3. COAGULATION ANALYZER (PT/INR)	15

3.1 Clinical & Functional Specifications	15

3.2 Engineering Technical Specifications	15

3.3 Standard Accessories	17

3.4 Installation Requirements	17

3.5 Documentation	18

3.6 Training	18

3.7 Warranty	18

4. CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) MACHINE	19

4.1 Clinical & Functional Specifications	19

4.2 Engineering Technical Specifications	19

4.3 Standard Accessories	22

4.4 Installation Requirements	24

4.5 Documentation	24

4.6 Training	24

4.7 Warranty	25

# 1. ARTERIAL BLOOD GAS (ABG) ANALYZER

Quantity: 3

## 1.1 Clinical & Functional Specifications

The Arterial Blood Gas Analyzer shall be a fully automated, microprocessor-controlled, benchtop point-of-care in-vitro diagnostic system designed for rapid quantitative measurement of blood gases, electrolytes, metabolites, and co-oximetry parameters from whole-blood samples in critical care units, operating theatres, emergency departments, neonatal intensive care units, and cardiac catheterization laboratories.

The system shall be capable of analyzing arterial, venous, capillary, and mixed-venous whole-blood samples with minimal sample volume requirement, providing accurate quantitative results within a clinically acceptable turnaround time. The analyzer shall support both syringe and capillary sample inlet modes and shall automatically aspirate the sample, perform measurement, calculate derived parameters, print/transmit the report, and prepare itself for the next sample without operator intervention.

The system shall provide, as a minimum, the following measured parameters: pH, partial pressure of oxygen (pO₂), partial pressure of carbon dioxide (pCO₂), sodium (Na⁺), potassium (K⁺), chloride (Cl⁻), ionized calcium (iCa²⁺), glucose, lactate, hematocrit (Hct), and total hemoglobin (tHb). The system shall additionally compute derived/calculated parameters including bicarbonate (HCO₃⁻), base excess (BE), oxygen saturation (SO₂ calc), anion gap, total CO₂, P/F ratio, A-a gradient, and temperature-corrected values.

The analyzer shall provide fully automatic one-point, two-point, and multi-point calibration at scheduled intervals with automatic quality control (QC) management, automatic reagent management, automatic sample handling, and automatic clot detection with self-clearing capability. The system shall be operable by trained clinical users with minimal biomedical intervention and shall provide clear on-screen error messages, corrective action prompts, and audible alarms.

The analyzer shall support bidirectional connectivity with the Hospital Information System (HIS), Laboratory Information System (LIS), and Electronic Medical Record (EMR) using standard healthcare communication protocols, and shall enable centralized data management, remote monitoring, user access control, and audit trail.

## 1.2 Engineering Technical Specifications

Measurement Technology

- Electrochemical sensors (potentiometric, amperometric) and/or optical/spectrophotometric technology for co-oximetry, or equivalent internationally recognized measurement principle

- Maintenance-free or long-life sensor cartridge / cassette-based measurement platform

- Sensor operational life: minimum 4 weeks or manufacturer-specified equivalent

Sample Handling

- Sample volume requirement: minimum practical volume, typically ≤ 200 µL for full-panel analysis

- Sample types: arterial, venous, capillary, and mixed-venous whole blood

- Sample inlet: automatic aspiration compatible with syringe and capillary tubes

- Automatic clot detection and clot management system

- Analysis time (result turnaround): typically ≤ 90 seconds per sample

Measurement Ranges and Accuracy (minimum acceptable)

- pH range: 6.500 — 8.000; resolution 0.001; accuracy within ±0.02

- pCO₂ range: 5 — 250 mmHg; accuracy within ±3% or better

- pO₂ range: 0 — 750 mmHg; accuracy within ±3% or better

- Na⁺ range: 100 — 200 mmol/L; accuracy within ±2 mmol/L

- K⁺ range: 0.5 — 15 mmol/L; accuracy within ±0.1 mmol/L

- Cl⁻ range: 50 — 200 mmol/L

- iCa²⁺ range: 0.2 — 5.0 mmol/L

- Glucose range: 20 — 700 mg/dL

- Lactate range: 0.3 — 25 mmol/L

- Hematocrit range: 10 — 75%

- Total Hb range: 3 — 25 g/dL

- SO₂ (measured) range: 0 — 100%

Display

- Color TFT LCD or equivalent display technology

- Minimum 8-inch diagonal viewing area

- Capacitive touchscreen operation

- Minimum resolution 800 × 480 pixels or better

- Multi-language user interface (English mandatory; additional languages selectable)

Processor, Memory & Data Storage

- Embedded medical-grade microprocessor of adequate performance

- Internal memory sufficient to store minimum 10,000 patient results and minimum 5,000 QC records with full audit trail

- Non-volatile storage with power-fail data protection

Calibration & Quality Control

- Fully automatic 1-point and 2-point calibration at programmable intervals

- Automatic QC scheduling with multi-level liquid or electronic QC

- Westgard rules or equivalent statistical QC evaluation

- Levey-Jennings chart display capability

User Interface & Software

- Icon-driven graphical user interface

- Configurable operator access levels with password protection

- Multi-user profile management

- Configurable reporting formats and units

- Full audit trail compliant with regulatory requirements

Connectivity

- Ethernet (RJ-45) 10/100/1000 Mbps

- Minimum 2 × USB ports

- Wireless connectivity (Wi-Fi) optional/configurable

- HL7 and POCT1-A2 protocols for LIS/HIS/EMR integration

- Barcode input for patient/operator/sample identification

Cybersecurity

- Password-protected multi-level user access

- Data encryption during transmission

- Audit trail for all user actions

- Compliance with IEC 80001 and manufacturer's cybersecurity documentation

Electrical Specifications

- Input voltage: 100 — 240 VAC, 50/60 Hz, single-phase

- Power consumption: typically ≤ 250 VA

- Internal rechargeable battery backup: minimum 30 minutes standby / short-run operational continuity

- Compatible with online UPS

Environmental Conditions

- Operating temperature: 15 — 32 °C

- Operating humidity: 20 — 85% RH, non-condensing

- Altitude: up to 3000 m

Dimensions and Weight (generalized)

- Compact benchtop configuration suitable for point-of-care placement

- Approximate weight: 10 — 25 kg depending on configuration

Safety & Compliance Standards

- IEC 60601-1 (General requirements for basic safety and essential performance)

- IEC 60601-1-2 (Electromagnetic compatibility)

- IEC 61010-1 / IEC 61010-2-101 (In-vitro diagnostic equipment safety)

- IEC 62304 (Medical device software life-cycle)

- IEC 62366 (Usability engineering)

- ISO 13485 (Quality management system — manufacturer)

- ISO 14971 (Risk management)

- CE MDR / IVDR marking and/or US FDA clearance

- HL7 / POCT1-A2 connectivity compliance

## 1.3 Standard Accessories

a) Integrated Barcode Scanner

- Built-in or externally tethered 1D/2D barcode scanner

- Capable of reading Code 39, Code 128, EAN, QR, and Data Matrix symbologies

- Used for patient ID, operator ID, sample ID, and reagent lot data entry

- Scan distance: minimum 5 — 20 cm; scan rate suitable for clinical use

- Compatible with the analyzer's data management software

b) Integrated Thermal Printer Module

- Built-in thermal printer within the analyzer chassis

- Print width suitable for standard patient result report

- Print speed: minimum 25 mm/second

- Compatible with standard thermal paper roll

- Supply: 2 rolls of thermal paper included at delivery

c) Calibration Verification Pack / Electronic QC Simulator

- Electronic quality control simulator or equivalent verification tool

- Capable of verifying sensor and electronic measurement integrity

- Compatible with the offered analyzer platform

- Storage life at recommended conditions clearly labeled

- Traceable to internationally recognized reference standards

d) External Waste Container Bottle

- Rigid, chemically resistant, biohazard-compatible plastic construction

- Adequate capacity for typical clinical workload between disposal cycles

- Automatic liquid level detection with overflow prevention sensor

- Sealed connection preventing leakage or aerosol release

- Compliant with infection-control and biohazard-waste handling requirements

e) Rechargeable Internal Battery Pack

- Sealed rechargeable lithium-ion or equivalent battery technology

- Backup runtime: minimum 30 minutes for orderly shutdown / short standby

- Charging time: maximum 8 hours from fully discharged

- Automatic switchover on mains failure

- Battery health monitoring displayed on user interface

f) Main Power Supply Adapter & Medical-Grade Power Cord

- Medical-grade power cord compliant with IEC 60601-1

- Input: 100 — 240 VAC, 50/60 Hz

- Country-specific plug configuration as required by purchaser

- Cord length: minimum 2.5 m

- Adequate insulation and strain relief

## 1.4 Installation Requirements

- Dedicated electrical outlet: 100 — 240 VAC, 50/60 Hz, single-phase, minimum 15 A

- Proper protective earthing with earth-to-neutral voltage \< 1 V

- Online UPS compatibility (recommended minimum 1 kVA)

- Bench space with vibration-free, level surface

- Ambient temperature 15 — 32 °C, humidity 20 — 85% RH non-condensing

- Adequate ventilation clearance around the unit as per manufacturer's guidance

- Network point (RJ-45) for LIS/HIS/EMR connectivity

- Waste disposal arrangement compliant with biohazard regulations

- No direct sunlight, no proximity to strong EM sources

## 1.5 Documentation

Supplier shall provide (in English, hard copy and soft copy):

- User / Operator Manual

- Service / Technical Manual

- Spare Parts Catalogue

- Factory Test Certificate

- Calibration Certificate

- Electrical Safety Test Report (per IEC 62353)

- Certificate of Origin

- CE / FDA regulatory certificates

- Preventive Maintenance Schedule

- Installation and Commissioning Report

## 1.6 Training

Supplier shall provide on-site training including:

- Clinical / user application training for laboratory and ICU staff

- Biomedical engineering training on system architecture

- Preventive maintenance training

- Basic troubleshooting and first-level fault correction

- QC management and data management training

- Minimum training duration adequate to achieve user competency

## 1.7 Warranty

Minimum comprehensive warranty: two (2) years from date of successful commissioning, covering:

- Labour and service visits

- All spare parts (excluding consumables and reagents)

- Scheduled preventive maintenance visits (minimum twice per year)

- Breakdown/corrective maintenance with defined response and resolution time

- All software updates and patches

- Remote and on-site technical support

- Guaranteed availability of spare parts for a minimum of 7 years post-installation

# 2. ACTIVATED CLOTTING TIME (ACT) MACHINE

Quantity: 2

## 2.1 Clinical & Functional Specifications

The Activated Clotting Time (ACT) Analyzer shall be a portable, microprocessor-controlled, point-of-care coagulation testing device intended for the rapid quantitative measurement of activated clotting time in whole-blood samples for monitoring high-dose and low-dose heparin anticoagulation therapy in cardiac catheterization laboratories, cardiac operating theatres (cardiopulmonary bypass), interventional radiology, hemodialysis, ECMO, and intensive care units.

The system shall be capable of performing multiple coagulation tests including, as a minimum, high-range ACT, low-range ACT, and additional cartridge/channel-based tests as offered by the platform. It shall support simultaneous dual-channel testing enabling parallel patient sample runs or duplicate testing on the same sample.

The instrument shall automatically detect clot formation using a validated mechanical, electromechanical, or optical detection principle, automatically time the reaction from sample introduction to clot endpoint, display the result on-screen, print the result, and store the result with full patient and operator identification.

The device shall provide temperature-controlled incubation blocks maintained at physiological temperature with continuous self-monitoring. It shall support on-screen operator prompts, quality control programs, electronic verification, and data export to the hospital information network.

## 2.2 Engineering Technical Specifications

Measurement Principle

- Mechanical, electromechanical (magnetic), or optical clot-endpoint detection technology, or equivalent internationally recognized principle

- Dual independent test channels operable simultaneously and independently

Measurement Range and Performance

- ACT measurement range: minimum 50 — 1000 seconds (covering low-range and high-range clinical needs)

- Resolution: 1 second

- Reproducibility (CV): typically ≤ 10% within the clinical range

- Result turnaround time: dependent on patient sample; endpoint detection automatic

Heating Block

- Dual test wells / dual-slot heating block with independent channels

- Temperature control: 37 °C ± 0.5 °C (physiological temperature)

- Digital temperature calibration feedback with continuous self-monitoring

- Alarm on out-of-range temperature

Display

- Color or monochrome LCD/TFT display

- Minimum 3.5-inch diagonal viewing area

- Adequate resolution for clear reading of numeric results, menus, and prompts

- Integrated menu selection buttons and/or touchscreen navigation

Processor, Memory & Data Storage

- Embedded microprocessor of adequate performance

- Internal memory sufficient to store minimum 500 patient results with time-stamp, patient ID, operator ID, and QC records

Test Compatibility

- Compatible with standard disposable ACT test cartridges/tubes for high-range and low-range applications

- Automatic test identification (via cartridge coding or manual selection)

User Interface & Software

- Menu-driven interface with multi-level password-protected access

- Operator lockout and QC-lockout features

- Configurable reporting format

- Full audit trail

Connectivity

- Minimum 1 × USB port for data export / firmware update

- Ethernet or wireless (Wi-Fi/Bluetooth) connectivity for LIS/HIS/EMR integration

- HL7 / POCT1-A2 protocol support

- Barcode input option for patient/operator ID

Electrical Specifications

- Input voltage: 100 — 240 VAC, 50/60 Hz via external medical-grade adapter

- Rechargeable internal battery

- Battery runtime: minimum 4 hours continuous point-of-care operation

- Charging time: maximum 6 — 8 hours

- Compatible with online UPS

Environmental Conditions

- Operating temperature: 15 — 30 °C

- Operating humidity: 20 — 80% RH, non-condensing

Dimensions and Weight (generalized)

- Portable, handheld or compact benchtop configuration

- Approximate weight: 1 — 4 kg depending on configuration

- Suitable for mobile bedside use

Safety & Compliance Standards

- IEC 60601-1 (Basic safety and essential performance)

- IEC 60601-1-2 (EMC)

- IEC 61010-1 / 61010-2-101 (IVD safety)

- IEC 62304 (Software life-cycle)

- IEC 62366 (Usability)

- ISO 13485 (Manufacturer QMS)

- ISO 14971 (Risk management)

- CE MDR / IVDR and/or US FDA clearance

## 2.3 Standard Accessories

a) Dual-Slot Heating Block Tracking Module

- Dual independent temperature-controlled test wells

- Digital temperature calibration feedback with continuous monitoring

- Temperature setpoint: 37 °C ± 0.5 °C

- Compatible with the offered ACT analyzer platform

- Self-diagnostic alarm on out-of-range temperature

b) Electronic Calibration Verification Tool / Test Plug

- Electronic verification device / calibration standard plug

- Used to verify electronic and thermal integrity of the measurement channel

- Compatible with the offered analyzer

- Storage case included

c) High-Resolution Digital Display Screen with Integrated Menu Selection Buttons

- Color or monochrome LCD/TFT display, minimum 3.5-inch diagonal

- Adequate resolution for clear numeric and text readability under ambient lighting

- Tactile, sealed menu selection buttons resistant to cleaning agents

- Backlit display for low-light environments (cath-lab conditions)

d) Built-in Thermal Printer Module

- Integrated thermal printer within the analyzer chassis

- Prints patient result, operator ID, timestamp, and QC information

- Compatible with standard thermal paper roll

- Supply: 2 rolls of thermal chart paper included at delivery

e) Rechargeable Internal Battery Pack

- Sealed rechargeable lithium-ion or equivalent battery

- Runtime: minimum 4 hours continuous point-of-care use

- Charging time: maximum 6 — 8 hours

- Automatic charge management and health indication on user interface

f) AC Power Supply Adapter and Main Electrical Power Cord

- Medical-grade external power adapter compliant with IEC 60601-1

- Input: 100 — 240 VAC, 50/60 Hz

- Country-specific plug configuration

- Cord length: minimum 2 m

## 2.4 Installation Requirements

- Dedicated electrical outlet: 100 — 240 VAC, 50/60 Hz

- Protective earthing per IEC standards

- UPS compatibility recommended for bedside use

- Bench or trolley placement, or handheld/mobile use

- Ambient temperature 15 — 30 °C, humidity 20 — 80% RH

- Optional network point for LIS integration

## 2.5 Documentation

Supplier shall provide:

- User / Operator Manual

- Service Manual

- Spare Parts Catalogue

- Factory Test Certificate

- Calibration Certificate

- Electrical Safety Test Report (IEC 62353)

- CE / FDA certificates

- Preventive Maintenance Schedule

- Installation and Commissioning Report

## 2.6 Training

- Clinical / user training for cath-lab, OT, ICU, and dialysis staff

- Biomedical engineering training

- Preventive maintenance training

- Basic troubleshooting

- QC and data management training

## 2.7 Warranty

Minimum comprehensive warranty: two (2) years from commissioning, covering labour, spare parts (excluding consumables), scheduled PM, breakdown maintenance, software updates, and technical support. Guaranteed spare-parts availability for minimum 7 years post-installation.

# 3. COAGULATION ANALYZER (PT/INR)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The PT/INR Coagulation Analyzer shall be a portable, microprocessor-controlled, point-of-care in-vitro diagnostic system designed for rapid quantitative measurement of prothrombin time (PT) and international normalized ratio (INR) from capillary or venous whole-blood samples for monitoring oral anticoagulation therapy (Vitamin K antagonists) in outpatient clinics, anticoagulation clinics, wards, ICUs, emergency departments, and home-care/point-of-care settings.

The analyzer shall accept minimal blood-sample volume, apply the sample onto a disposable single-use reagent test strip, automatically initiate the coagulation reaction, detect the clot endpoint, calculate PT and INR values with ISI correction, display and print the result, and store the record with patient/operator identification.

The system shall support electronic quality control, on-strip quality control, operator lockout, patient database management, and bidirectional connectivity to LIS/HIS/EMR.

## 3.2 Engineering Technical Specifications

Measurement Principle

- Electrochemical, optical, or mechanical clot endpoint detection using disposable reagent test strips

- Automatic ISI correction based on strip lot data (barcode-encoded)

Measurement Range and Performance

- INR range: minimum 0.8 — 8.0 or wider

- PT range: adequate to cover clinical monitoring needs

- Sample volume: minimum practical volume, typically ≤ 20 µL

- Sample types: capillary or venous whole blood

- Result turnaround time: typically ≤ 2 minutes

Display

- Color TFT LCD or equivalent

- High-resolution color touchscreen

- Minimum 3.5-inch diagonal viewing area

- Minimum resolution 320 × 240 pixels or better

- Backlit for low-light environments

Processor, Memory & Data Storage

- Embedded microprocessor

- Internal memory: minimum 1000 patient results with timestamp, patient/operator ID

- Non-volatile storage with battery backup

Reagent Strip Handling

- Compatible with manufacturer-specified single-use disposable strips

- Automatic strip identification via integrated barcode reader

- Automatic lot/expiry management

- On-strip quality control detection

User Interface & Software

- Icon-driven touchscreen interface

- Multi-level operator access with password protection

- Operator lockout, QC lockout, patient lockout features

- Configurable reporting format and units

Connectivity

- Wireless (Wi-Fi and/or Bluetooth) connectivity

- USB port

- Optional docking cradle with Ethernet

- HL7 / POCT1-A2 protocol support

- Barcode input for patient/operator/strip lot identification

Electrical Specifications

- Input: 100 — 240 VAC, 50/60 Hz via medical-grade external adapter

- Rechargeable internal lithium-ion battery

- Battery runtime: minimum 50 tests per full charge

- Charging time: maximum 4 — 6 hours

- Desktop charging cradle included

Environmental Conditions

- Operating temperature: 15 — 32 °C

- Operating humidity: 20 — 85% RH, non-condensing

Dimensions and Weight (generalized)

- Handheld, portable configuration

- Approximate weight: ≤ 1 kg including battery

- Ergonomic single-hand operation

Safety & Compliance Standards

- IEC 60601-1, IEC 60601-1-2

- IEC 61010-1 / 61010-2-101

- IEC 62304, IEC 62366

- ISO 13485, ISO 14971

- CE MDR / IVDR and/or US FDA clearance

## 3.3 Standard Accessories

a) High-Resolution Color Touchscreen User Interface

- Color TFT LCD touchscreen, minimum 3.5-inch diagonal

- Capacitive or resistive touch technology

- Icon-driven configuration menu

- Backlit for wide ambient conditions

- Cleaning-agent-resistant surface

b) Integrated Barcode Reader

- Built-in 1D/2D barcode reader

- Rapid reagent strip lot identification, patient ID, operator ID

- Compatible with standard barcode symbologies (Code 39, Code 128, QR, Data Matrix)

- Integrated within the device housing

c) Electronic Quality Control (EQC) Testing Strip / Verification Standard

- Electronic verification standard block or QC strip

- Validates measurement channel integrity

- Compatible with the offered analyzer

- Traceable to internationally recognized reference

d) Rigid Plastic Transport and Protection Case

- Impact-resistant, rigid plastic carry case

- Fitted foam layout with slots for analyzer, cradle, adapter, cables, and consumables

- Suitable for safe transport in point-of-care and mobile clinical use

- Handle for portability

e) Desktop Charging Cradle Stand with Universal AC Power Adapter

- Desktop docking cradle for stable seating and charging of the analyzer

- Charging contacts with automatic engagement

- Universal AC power adapter, medical-grade, IEC 60601-1 compliant

- Input: 100 — 240 VAC, 50/60 Hz

- Optional Ethernet/USB pass-through for data transfer

f) Rechargeable Lithium-ion Battery Pack (Pre-installed)

- Sealed rechargeable Li-ion battery installed within device handle

- Runtime: minimum 50 tests per full charge

- Charging time: maximum 4 — 6 hours

- Automatic charge management and health indication

## 3.4 Installation Requirements

- Electrical outlet: 100 — 240 VAC, 50/60 Hz

- Protective earthing per IEC standards

- Bench or desk space for cradle

- Ambient temperature 15 — 32 °C, humidity 20 — 85% RH

- Optional Wi-Fi network access for LIS/HIS connectivity

## 3.5 Documentation

- User / Operator Manual

- Service Manual

- Spare Parts Catalogue

- Factory Test & Calibration Certificate

- Electrical Safety Test Report (IEC 62353)

- CE / FDA certificates

- Preventive Maintenance Schedule

- Installation and Commissioning Report

## 3.6 Training

- Clinical / user training for anticoagulation clinic, ward, and OPD staff

- Biomedical engineering training

- Preventive maintenance training

- Basic troubleshooting

- QC and data management training

## 3.7 Warranty

Minimum comprehensive warranty: two (2) years from commissioning, covering labour, spare parts (excluding consumable strips), scheduled PM, breakdown maintenance, software updates, and technical support. Guaranteed spare-parts availability for minimum 7 years.

# 4. CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) MACHINE

Quantity: 2

## 4.1 Clinical & Functional Specifications

The Continuous Renal Replacement Therapy (CRRT) Machine shall be a fully automated, microprocessor-controlled, mobile extracorporeal blood-purification system designed for continuous, slow, and controlled renal replacement therapy in critically ill adult and pediatric patients with acute kidney injury (AKI), fluid overload, sepsis, hepatic failure, drug intoxication, and multi-organ dysfunction managed in intensive care units, high-dependency units, cardiac ICUs, and specialized nephrology units.

The system shall support, as a minimum, the following therapy modalities: Slow Continuous Ultrafiltration (SCUF), Continuous Veno-Venous Hemofiltration (CVVH), Continuous Veno-Venous Hemodialysis (CVVHD), Continuous Veno-Venous Hemodiafiltration (CVVHDF), and Therapeutic Plasma Exchange (TPE). Additional modalities such as hemoperfusion, coupled plasma filtration adsorption (CPFA), and continuous cytokine removal shall be supported where offered.

The machine shall provide four independent, high-precision peristaltic pumps for blood, pre/post-dilution replacement fluid, dialysate, and effluent, with fully independent gravimetric fluid balance control using four independent weight scale platforms. It shall provide integrated syringe pump for anticoagulant (heparin or citrate) delivery, integrated blood/fluid warming, comprehensive pressure and safety monitoring, and full data logging.

The user interface shall be a touchscreen with intuitive workflow guidance for setup, priming, running, alarm management, and de-installation. The system shall include automatic self-test, guided priming, comprehensive alarms, and internal backup battery for uninterrupted therapy during transient power loss or intra-hospital transfer.

## 4.2 Engineering Technical Specifications

Therapy Modalities

- SCUF, CVVH (pre- and post-dilution), CVVHD, CVVHDF, TPE

- Adult and pediatric therapy support

- Additional modalities configurable/expandable as offered

Pump System

- Four independent peristaltic roller pumps: blood, pre/post-dilution, dialysate, effluent

- Blood flow rate range: minimum 10 — 500 mL/min, adjustable in fine increments (typically 10 mL/min)

- Replacement fluid rate: minimum 0 — 8000 mL/hour

- Dialysate flow rate: minimum 0 — 8000 mL/hour

- Effluent flow rate: minimum 0 — 10000 mL/hour

- Net ultrafiltration rate: adjustable minimum 0 — 2000 mL/hour

- Flow accuracy: within ±10% or better across the clinical range

Weight/Fluid Balance System

- Four independent electronic weight scale platforms with auto-calibration

- Each scale capacity: minimum 10 kg per bag/container

- Fluid balance accuracy: within ±30 g/hour or manufacturer-specified equivalent

- Automatic detection of empty/full bag conditions

Anticoagulation System

- Integrated syringe pump for systemic heparin anticoagulation

- Syringe compatibility with standard 20 — 50 mL medical syringes

- Delivery rate range: adjustable minimum 0 — 10 mL/hour with fine resolution

- Regional citrate anticoagulation (RCA) protocol support (configurable)

Safety Monitoring

- Non-invasive ultrasonic air/bubble detector on venous line

- Automatic line clamp on air detection or safety violation

- Blood leak detector on effluent line

- Access pressure sensor (arterial line): typical range −250 to +50 mmHg

- Return pressure sensor (venous line): typical range −50 to +350 mmHg

- Effluent pressure sensor

- Transmembrane pressure (TMP) monitoring and calculation

- Filter pressure drop monitoring

- Sensor accuracy: within ±10 mmHg or better

Blood/Fluid Warming

- Integrated blood/fluid warming jacket or in-line warmer

- Target patient-return temperature: 35 — 39 °C, adjustable

- Independent safety thermostat with over-temperature cut-off

- Continuous temperature monitoring with alarm

Display

- Color TFT LCD touchscreen

- Minimum 12-inch diagonal viewing area

- Minimum resolution 1024 × 768 pixels or better

- Icon-driven multi-language interface (English mandatory)

- On-screen trend display and workflow guidance

Processor, Memory & Data Storage

- Embedded medical-grade microprocessor

- Internal memory sufficient for minimum 72 hours of continuous therapy data logging per patient

- Non-volatile storage for treatment history, alarms, and event log

- Data export capability

Alarm System

- Multi-level audible and visual alarms per IEC 60601-1-8

- Alarm categories: high/medium/low priority

- Configurable alarm limits within clinically safe bounds

- Alarm history with timestamp

Circuit / Consumables

- Compatible with the manufacturer's disposable extracorporeal circuit sets (hemofilters, tubing, bags)

- Automatic circuit recognition where offered

- Guided priming and rinseback sequences

Connectivity

- Ethernet (RJ-45) for HIS/EMR/ICU-CIS integration

- USB port(s) for data export and service

- HL7 protocol support

Cybersecurity

- Password-protected multi-level user access

- Audit trail

- Compliance with IEC 80001 principles

Electrical Specifications

- Input voltage: 100 — 240 VAC, 50/60 Hz, single-phase

- Power consumption: typically ≤ 500 VA

- Internal emergency backup battery: minimum 30 minutes uninterrupted therapy continuation on mains failure or during intra-hospital transfer

- Compatible with online UPS (recommended)

Environmental Conditions

- Operating temperature: 15 — 30 °C

- Operating humidity: 20 — 80% RH, non-condensing

Mechanical Configuration

- Heavy-duty mobile trolley chassis

- Anti-static, locking, swivel casters

- Steering handle for maneuvering

- Articulated support arm for extracorporeal circuit organization

- IV pole for fluid bags

Dimensions and Weight (generalized)

- Compact mobile configuration suitable for bedside ICU use

- Approximate weight: 70 — 120 kg depending on configuration

Infection Control

- External surfaces resistant to standard hospital disinfectants

- Smooth, sealed, easily cleanable enclosure

- Compatible with routine infection-control protocols

Safety & Compliance Standards

- IEC 60601-1 (Basic safety and essential performance)

- IEC 60601-1-2 (EMC)

- IEC 60601-1-6 (Usability)

- IEC 60601-1-8 (Alarm systems)

- IEC 60601-2-16 (Particular requirements for hemodialysis, hemodiafiltration, and hemofiltration equipment)

- IEC 62304 (Software life-cycle)

- IEC 62366 (Usability engineering)

- ISO 13485 (Manufacturer QMS)

- ISO 14971 (Risk management)

- CE MDR and/or US FDA clearance

## 4.3 Standard Accessories

a) Four High-Precision Automated Peristaltic Roller Pumps

- Four independent peristaltic roller pumps integrated within the main housing

- Dedicated pumps for blood, pre/post-dilution replacement, dialysate, and effluent

- Flow accuracy within ±10% or better

- Occlusion detection and automatic stop on abnormal loading

- Compatible with standard extracorporeal tubing sets of the offered platform

b) Four Independent Weight Scale Platforms

- Four independent electronic gravimetric scales

- Auto-calibration verification sensors

- Capacity: minimum 10 kg per scale

- Fluid balance resolution and accuracy adequate for critical-care CRRT

- Individual scale identification and self-diagnostic alarm

c) Integrated Syringe Pump for Anticoagulant Delivery

- Integrated syringe pump driver mounted on the machine

- Compatible with standard 20 — 50 mL medical syringes (multiple brands)

- Delivery rate: adjustable 0 — 10 mL/hour minimum, with fine resolution

- Bolus, occlusion, and end-of-syringe alarms

- Suitable for heparin and equivalent anticoagulants

d) Ultrasonic Air Bubble Detector and Automated Line Clamp

- Non-invasive ultrasonic air/bubble detector on venous return line

- Sensitivity adequate to detect clinically significant air volumes

- Automated electromechanical fluid-line clamp with immediate closure on detection

- Automatic pump stop coupled with clamp closure

- Self-test at each therapy setup

e) Dual Blood Line Pressure Sensors

- Multiple non-invasive pressure sensors monitoring access, return, effluent, and filter (transmembrane) pressures

- Continuous display and trending on user interface

- Configurable high/low alarm limits within clinically safe bounds

- Accuracy: within ±10 mmHg or better

f) Integrated Blood/Fluid Warming System with Safety Thermostat

- Integrated in-line blood/fluid warming device (jacket, dry-heat plate, or equivalent)

- Adjustable target temperature 35 — 39 °C

- Independent hardware safety thermostat with over-temperature cut-off

- Continuous temperature monitoring with alarm and automatic shutdown

g) Articulated Support Arm for Extracorporeal Tubing

- Articulated adjustable arm bracket mounted on the chassis

- Organizes and supports extracorporeal blood and fluid lines

- Prevents kinking and tension on the circuit

- Cleanable, corrosion-resistant material

h) High-Capacity Internal Emergency Backup Battery

- Sealed rechargeable internal battery pack

- Runtime: minimum 30 minutes uninterrupted therapy on mains failure

- Automatic switchover on power interruption

- Charging time: maximum 8 hours

- Battery health monitoring on user interface

- Adequate for intra-hospital patient transfer

i) Heavy-Duty Equipment Trolley Chassis

- Rigid metal or reinforced polymer chassis

- Four anti-static, swivel, locking casters (minimum two lockable)

- Steering handle with ergonomic grip

- Stable footprint preventing tip-over

- Integrated IV pole for fluid bags

- Compatible with elevator/hospital corridor dimensions

## 4.4 Installation Requirements

- Dedicated grounded electrical outlet: 100 — 240 VAC, 50/60 Hz, minimum 15 A

- Protective earthing per IEC 60364 / equivalent

- Online UPS compatibility (recommended minimum 2 kVA per unit)

- ICU bedside space with unobstructed access on all sides

- Ambient temperature 15 — 30 °C, humidity 20 — 80% RH

- Network point (RJ-45) for CIS/HIS integration

- Adequate storage for consumables and fluid bags

- No requirement for hospital water supply (CRRT uses pre-packaged sterile fluids)

- Waste-fluid disposal arrangement

## 4.5 Documentation

Supplier shall provide:

- User / Operator Manual

- Service / Technical Manual with circuit diagrams

- Spare Parts Catalogue

- Factory Test Certificate

- Calibration Certificate

- Electrical Safety Test Report (IEC 62353)

- CE MDR / FDA regulatory certificates

- Certificate of Origin

- Preventive Maintenance Schedule

- Installation and Commissioning Report

- Compatible consumables list

## 4.6 Training

Supplier shall provide comprehensive on-site training:

- Clinical / user training for ICU nursing and medical staff (multiple sessions per shift)

- Advanced clinical application training on all therapy modalities

- Biomedical engineering training on machine architecture, service, and calibration

- Preventive maintenance training

- Troubleshooting and first-level fault correction

- Anticoagulation protocol training

- Data management and reporting training

- Refresher training within 6 — 12 months post-commissioning

## 4.7 Warranty

Minimum comprehensive warranty: two (2) years from date of successful commissioning, covering:

- All labour and unlimited service visits

- All spare parts (excluding disposable consumables)

- Scheduled preventive maintenance visits (minimum twice per year)

- Breakdown/corrective maintenance with defined response (≤ 24 hours) and resolution (≤ 72 hours) time

- All software updates and cybersecurity patches

- Remote and on-site technical support

- Battery replacement if capacity falls below manufacturer's specification during warranty

- Guaranteed availability of spare parts and consumables for minimum 10 years post-installation