TECHNICAL TENDER SPECIFICATIONS — REHABILITATION EQUIPMENT

Technical Tender Document

Table of Contents

1. Therapeutic Ultrasound Unit	22

1.1 Clinical & Functional Specifications	22

1.2 Engineering Technical Specifications	22

1.3 Standard Accessories	23

1.4 Installation Requirements	23

1.5 Documentation	23

1.6 Training	23

1.7 Warranty	23

2. Short-Wave Diathermy Unit	24

2.1 Clinical & Functional Specifications	24

2.2 Engineering Technical Specifications	24

2.3 Standard Accessories	24

2.4 Installation Requirements	25

2.5 Documentation	25

2.6 Training	25

2.7 Warranty	25

3. TENS Unit	26

3.1 Clinical & Functional Specifications	26

3.2 Engineering Technical Specifications	26

3.3 Standard Accessories	26

3.4 Installation Requirements	27

3.5 Documentation	27

3.6 Training	27

3.7 Warranty	27

4. Interferential Current (IFC) Unit	28

4.1 Clinical & Functional Specifications	28

4.2 Engineering Technical Specifications	28

4.3 Standard Accessories	28

4.4 Installation Requirements	28

4.5 Documentation	29

4.6 Training	29

4.7 Warranty	29

5. Laser Therapy Unit	30

5.1 Clinical & Functional Specifications	30

5.2 Engineering Technical Specifications	30

5.3 Standard Accessories	30

5.4 Installation Requirements	30

5.5 Documentation	31

5.6 Training	31

5.7 Warranty	31

6. Electrical Muscle Stimulator (EMS)	32

6.1 Clinical & Functional Specifications	32

6.2 Engineering Technical Specifications	32

6.3 Standard Accessories	32

6.4 Installation Requirements	32

6.5 Documentation	33

6.6 Training	33

6.7 Warranty	33

7. General Electrical Stimulator (Faradic / Galvanic)	34

7.1 Clinical & Functional Specifications	34

7.2 Engineering Technical Specifications	34

7.3 Standard Accessories	34

7.4 Installation Requirements	34

7.5 Documentation	34

7.6 Training	35

7.7 Warranty	35

8. Shockwave Therapy Unit (Radial)	36

8.1 Clinical & Functional Specifications	36

8.2 Engineering Technical Specifications	36

8.3 Standard Accessories	36

8.4 Installation Requirements	36

8.5 Documentation	37

8.6 Training	37

8.7 Warranty	37

9. Transcranial Direct Current Stimulation (tDCS) Unit	38

9.1 Clinical & Functional Specifications	38

9.2 Engineering Technical Specifications	38

9.3 Standard Accessories	38

9.4 Installation Requirements	38

9.5 Documentation	38

9.6 Training	39

9.7 Warranty	39

10. Biofeedback Unit	41

10.1 Clinical & Functional Specifications	41

10.2 Engineering Technical Specifications	41

10.3 Standard Accessories	41

10.4 Installation Requirements	42

10.5 Documentation	42

10.6 Training	42

10.7 Warranty	42

11. Cervical Traction Unit	43

11.1 Clinical & Functional Specifications	43

11.2 Engineering Technical Specifications	43

11.3 Standard Accessories	43

11.4 Installation Requirements	44

11.5 Documentation	44

11.6 Training	44

11.7 Warranty	44

12. Lumbar Traction Unit	45

12.1 Clinical & Functional Specifications	45

12.2 Engineering Technical Specifications	45

12.3 Standard Accessories	45

12.4 Installation Requirements	45

12.5 Documentation	46

12.6 Training	46

12.7 Warranty	46

13. Paraffin Wax Bath Unit	47

13.1 Clinical & Functional Specifications	47

13.2 Engineering Technical Specifications	47

13.3 Standard Accessories	47

13.4 Installation Requirements	47

13.5 Documentation	48

13.6 Training	48

13.7 Warranty	48

14. Infrared Radiator	49

14.1 Clinical & Functional Specifications	49

14.2 Engineering Technical Specifications	49

14.3 Standard Accessories	49

14.4 Installation Requirements	49

14.5 Documentation	50

14.6 Training	50

14.7 Warranty	50

15. Electrical Heating Pad	51

15.1 Clinical & Functional Specifications	51

15.2 Engineering Technical Specifications	51

15.3 Standard Accessories	51

15.4 Installation Requirements	51

15.5 Documentation	51

15.6 Training	52

15.7 Warranty	52

16. Moist Heating Pad (Hydrocollator-Style Pack)	53

16.1 Clinical & Functional Specifications	53

16.2 Engineering Technical Specifications	53

16.3 Standard Accessories	53

16.4 Installation Requirements	53

16.5 Documentation	53

16.6 Training	53

16.7 Warranty	54

17. Cold Pack	55

17.1 Clinical & Functional Specifications	55

17.2 Engineering Technical Specifications	55

17.3 Standard Accessories	55

17.4 Installation Requirements	55

17.5 Documentation	55

17.6 Training	55

17.7 Warranty	56

18. Electrical Tilt Table	58

18.1 Clinical & Functional Specifications	58

18.2 Engineering Technical Specifications	58

18.3 Standard Accessories	58

18.4 Installation Requirements	59

18.5 Documentation	59

18.6 Training	59

18.7 Warranty	59

19. Rehabilitation Treadmill	60

19.1 Clinical & Functional Specifications	60

19.2 Engineering Technical Specifications	60

19.3 Standard Accessories	60

19.4 Installation Requirements	61

19.5 Documentation	61

19.6 Training	61

19.7 Warranty	61

20. Recumbent Bike	62

20.1 Clinical & Functional Specifications	62

20.2 Engineering Technical Specifications	62

20.3 Standard Accessories	62

20.4 Installation Requirements	63

20.5 Documentation	63

20.6 Training	63

20.7 Warranty	63

21. Elliptical / Cross Trainer	64

21.1 Clinical & Functional Specifications	64

21.2 Engineering Technical Specifications	64

21.3 Standard Accessories	64

21.4 Installation Requirements	65

21.5 Documentation	65

21.6 Training	65

21.7 Warranty	65

22. Arm Ergometer (Upper Body Cycle)	66

22.1 Clinical & Functional Specifications	66

22.2 Engineering Technical Specifications	66

22.3 Standard Accessories	66

22.4 Installation Requirements	66

22.5 Documentation	67

22.6 Training	67

22.7 Warranty	67

23. Stepper for Lower-Limb Re-education	68

23.1 Clinical & Functional Specifications	68

23.2 Engineering Technical Specifications	68

23.3 Standard Accessories	68

23.4 Installation Requirements	68

23.5 Documentation	69

23.6 Training	69

23.7 Warranty	69

24. Multi-Station Gym System	70

24.1 Clinical & Functional Specifications	70

24.2 Engineering Technical Specifications	70

24.3 Standard Accessories	70

24.4 Installation Requirements	71

24.5 Documentation	71

24.6 Training	71

24.7 Warranty	71

25. Parallel Bars (Adult)	73

25.1 Clinical & Functional Specifications	73

25.2 Engineering Technical Specifications	73

25.3 Standard Accessories	73

25.4 Installation Requirements	74

25.5 Documentation	74

25.6 Training	74

25.7 Warranty	74

26. Paediatric Parallel Bars	75

26.1 Clinical & Functional Specifications	75

26.2 Engineering Technical Specifications	75

26.3 Standard Accessories	75

26.4 Installation Requirements	76

26.5 Documentation	76

26.6 Training	76

26.7 Warranty	76

27. Exercise Staircase	77

27.1 Clinical & Functional Specifications	77

27.2 Engineering Technical Specifications	77

27.3 Standard Accessories	77

27.4 Installation Requirements	78

27.5 Documentation	78

27.6 Training	78

27.7 Warranty	78

28. Para-Podium (Standing Frame)	79

28.1 Clinical & Functional Specifications	79

28.2 Engineering Technical Specifications	79

28.3 Standard Accessories	79

28.4 Installation Requirements	80

28.5 Documentation	80

28.6 Training	80

28.7 Warranty	80

29. Universal Exercise Unit with Suit (Rocher Cage / Spider Cage System)	81

29.1 Clinical & Functional Specifications	81

29.2 Engineering Technical Specifications	81

29.3 Standard Accessories	81

29.4 Installation Requirements	82

29.5 Documentation	82

29.6 Training	82

29.7 Warranty	82

30. Robotic Hand Glove	84

30.1 Clinical & Functional Specifications	84

30.2 Engineering Technical Specifications	84

30.3 Standard Accessories	84

30.4 Installation Requirements	85

30.5 Documentation	85

30.6 Training	85

30.7 Warranty	85

31. Stroke Rehabilitation Glove (Passive)	86

31.1 Clinical & Functional Specifications	86

31.2 Engineering Technical Specifications	86

31.3 Standard Accessories	86

31.4 Installation Requirements	86

31.5 Documentation	86

31.6 Training	86

31.7 Warranty	87

32. Balance / Wobble Board	88

32.1 Clinical & Functional Specifications	88

32.2 Engineering Technical Specifications	88

32.3 Standard Accessories	88

32.4 Installation Requirements	88

32.5 Documentation	88

32.6 Training	88

32.7 Warranty	89

33. Bosu Ball (Balance Trainer)	90

33.1 Clinical & Functional Specifications	90

33.2 Engineering Technical Specifications	90

33.3 Standard Accessories	90

33.4 Installation Requirements	90

33.5 Documentation	90

33.6 Training	90

33.7 Warranty	91

34. Mini-Trampoline (Rebounder)	92

34.1 Clinical & Functional Specifications	92

34.2 Engineering Technical Specifications	92

34.3 Standard Accessories	92

34.4 Installation Requirements	92

34.5 Documentation	92

34.6 Training	93

34.7 Warranty	93

35. Platform Swing (Sensory Integration)	94

35.1 Clinical & Functional Specifications	94

35.2 Engineering Technical Specifications	94

35.3 Standard Accessories	94

35.4 Installation Requirements	94

35.5 Documentation	94

35.6 Training	95

35.7 Warranty	95

36. Disc Swing & Sensory-Integration Swings	96

36.1 Clinical & Functional Specifications	96

36.2 Engineering Technical Specifications	96

36.3 Standard Accessories	96

36.4 Installation Requirements	96

36.5 Documentation	96

36.6 Training	97

36.7 Warranty	97

37. Peanut Ball	100

37.1 Clinical & Functional Specifications	100

37.2 Engineering Technical Specifications	100

37.3 Standard Accessories	100

37.4 Installation Requirements	100

37.5 Documentation	100

37.6 Training	100

37.7 Warranty	101

38. Gym Ball Set	102

38.1 Clinical & Functional Specifications	102

38.2 Engineering Technical Specifications	102

38.3 Standard Accessories	102

38.4 Installation Requirements	102

38.5 Documentation	102

38.6 Training	102

38.7 Warranty	103

39. Medicine Ball Set	104

39.1 Clinical & Functional Specifications	104

39.2 Engineering Technical Specifications	104

39.3 Standard Accessories	104

39.4 Installation Requirements	104

39.5 Documentation	104

39.6 Training	104

39.7 Warranty	104

40. Resistance Tubing Set	105

40.1 Clinical & Functional Specifications	105

40.2 Engineering Technical Specifications	105

40.3 Standard Accessories	105

40.4 Installation Requirements	105

40.5 Documentation	105

40.6 Training	105

40.7 Warranty	106

41. TheraBand (Continuous-Roll Elastic Resistance Bands)	107

41.1 Clinical & Functional Specifications	107

41.2 Engineering Technical Specifications	107

41.3 Standard Accessories	107

41.4 Installation Requirements	107

41.5 Documentation	107

41.6 Training	107

41.7 Warranty	107

42. Foam Roller	109

42.1 Clinical & Functional Specifications	109

42.2 Engineering Technical Specifications	109

42.3 Standard Accessories	109

42.4 Installation Requirements	109

42.5 Documentation	109

42.6 Training	109

42.7 Warranty	109

43. Foam Wedge Set	111

43.1 Clinical & Functional Specifications	111

43.2 Engineering Technical Specifications	111

43.3 Standard Accessories	111

43.4 Installation Requirements	111

43.5 Documentation	111

43.6 Training	111

43.7 Warranty	111

44. Dumbbell Set with Hand Straps (1 kg)	112

44.1 Clinical & Functional Specifications	112

44.2 Engineering Technical Specifications	112

44.3 Standard Accessories	112

44.4 Installation Requirements	112

44.5 Documentation	112

44.6 Training	112

44.7 Warranty	112

45. Wrist Weight (1 kg)	113

45.1 Clinical & Functional Specifications	113

45.2 Engineering Technical Specifications	113

45.3 Standard Accessories	113

45.4 Installation Requirements	113

45.5 Documentation	113

45.6 Training	113

45.7 Warranty	113

46. Ankle Weight (1 kg)	115

46.1 Clinical & Functional Specifications	115

46.2 Engineering Technical Specifications	115

46.3 Standard Accessories	115

46.4 Installation Requirements	115

46.5 Documentation	115

46.6 Training	115

46.7 Warranty	115

47. Mobilization Belt	116

47.1 Clinical & Functional Specifications	116

47.2 Engineering Technical Specifications	116

47.3 Standard Accessories	116

47.4 Installation Requirements	116

47.5 Documentation	116

47.6 Training	116

47.7 Warranty	116

48. Finger Wall Ladder (Shoulder Finger Ladder)	118

48.1 Clinical & Functional Specifications	118

48.2 Engineering Technical Specifications	118

48.3 Standard Accessories	118

48.4 Installation Requirements	118

48.5 Documentation	118

48.6 Training	118

48.7 Warranty	118

49. Chiropractic / Percussion Massage Gun	121

49.1 Clinical & Functional Specifications	121

49.2 Engineering Technical Specifications	121

49.3 Standard Accessories	121

49.4 Installation Requirements	121

49.5 Documentation	122

49.6 Training	122

49.7 Warranty	122

50. Goniometry Set	123

50.1 Clinical & Functional Specifications	123

50.2 Engineering Technical Specifications	123

50.3 Standard Accessories	123

50.4 Installation Requirements	123

50.5 Documentation	123

50.6 Training	123

50.7 Warranty	123

51. Inclinometer	124

51.1 Clinical & Functional Specifications	124

51.2 Engineering Technical Specifications	124

51.3 Standard Accessories	124

51.4 Installation Requirements	124

51.5 Documentation	124

51.6 Training	124

51.7 Warranty	124

52. Dynamometer (Hand-Grip)	125

52.1 Clinical & Functional Specifications	125

52.2 Engineering Technical Specifications	125

52.3 Standard Accessories	125

52.4 Installation Requirements	125

52.5 Documentation	125

52.6 Training	125

52.7 Warranty	125

53. Algometer (Pressure Pain Threshold)	126

53.1 Clinical & Functional Specifications	126

53.2 Engineering Technical Specifications	126

53.3 Standard Accessories	126

53.4 Installation Requirements	126

53.5 Documentation	126

53.6 Training	126

53.7 Warranty	126

54. Scoliometer	127

54.1 Clinical & Functional Specifications	127

54.2 Engineering Technical Specifications	127

54.3 Standard Accessories	127

54.4 Installation Requirements	127

54.5 Documentation	127

54.6 Training	127

54.7 Warranty	127

55. Stopwatch	128

55.1 Clinical & Functional Specifications	128

55.2 Engineering Technical Specifications	128

55.3 Standard Accessories	128

55.4 Installation Requirements	128

55.5 Documentation	128

55.6 Training	128

55.7 Warranty	128

56. Measuring Tape (Anthropometric)	129

56.1 Clinical & Functional Specifications	129

56.2 Engineering Technical Specifications	129

56.3 Standard Accessories	129

56.4 Installation Requirements	129

56.5 Documentation	129

56.6 Training	129

56.7 Warranty	129

57. Reflex Hammer Kit	130

57.1 Clinical & Functional Specifications	130

57.2 Engineering Technical Specifications	130

57.3 Standard Accessories	130

57.4 Installation Requirements	130

57.5 Documentation	130

57.6 Training	130

57.7 Warranty	130

58. Tuning Fork Kit	131

58.1 Clinical & Functional Specifications	131

58.2 Engineering Technical Specifications	131

58.3 Standard Accessories	131

58.4 Installation Requirements	131

58.5 Documentation	131

58.6 Training	131

58.7 Warranty	131

59. Sensory Examination Kit	132

59.1 Clinical & Functional Specifications	132

59.2 Engineering Technical Specifications	132

59.3 Standard Accessories	132

59.4 Installation Requirements	132

59.5 Documentation	132

59.6 Training	132

59.7 Warranty	132

60. Monofilaments (Semmes-Weinstein Sensory Testing)	133

60.1 Clinical & Functional Specifications	133

60.2 Engineering Technical Specifications	133

60.3 Standard Accessories	133

60.4 Installation Requirements	133

60.5 Documentation	133

60.6 Training	133

60.7 Warranty	133

61. Cotton Balls (Sensory Kit)	134

61.1 Clinical & Functional Specifications	134

61.2 Engineering Technical Specifications	134

61.3 Standard Accessories	134

61.4 Installation Requirements	134

61.5 Documentation	134

61.6 Training	134

61.7 Warranty	134

62. Stereognosis Kit	135

62.1 Clinical & Functional Specifications	135

62.2 Engineering Technical Specifications	135

62.3 Standard Accessories	135

62.4 Installation Requirements	135

62.5 Documentation	135

62.6 Training	135

62.7 Warranty	135

63. Barognosis Kit	136

63.1 Clinical & Functional Specifications	136

63.2 Engineering Technical Specifications	136

63.3 Standard Accessories	136

63.4 Installation Requirements	136

63.5 Documentation	136

63.6 Training	136

63.7 Warranty	136

64. Cranial Nerve Examination Kit	137

64.1 Clinical & Functional Specifications	137

64.2 Engineering Technical Specifications	137

64.3 Standard Accessories	137

64.4 Installation Requirements	137

64.5 Documentation	137

64.6 Training	137

64.7 Warranty	137

65. Sensory Integration Kit	140

65.1 Clinical & Functional Specifications	140

65.2 Engineering Technical Specifications	140

65.3 Standard Accessories	140

65.4 Installation Requirements	140

65.5 Documentation	140

65.6 Training	140

65.7 Warranty	141

66. Fine Motor Equipment	142

66.1 Clinical & Functional Specifications	142

66.2 Engineering Technical Specifications	142

66.3 Standard Accessories	142

66.4 Installation Requirements	142

66.5 Documentation	142

66.6 Training	142

66.7 Warranty	142

67. Finger Exerciser	144

67.1 Clinical & Functional Specifications	144

67.2 Engineering Technical Specifications	144

67.3 Standard Accessories	144

67.4 Installation Requirements	144

67.5 Documentation	144

67.6 Training	144

67.7 Warranty	144

68. Hand Exercise Ball	145

68.1 Clinical & Functional Specifications	145

68.2 Engineering Technical Specifications	145

68.3 Standard Accessories	145

68.4 Installation Requirements	145

68.5 Documentation	145

68.6 Training	145

68.7 Warranty	145

69. Hand Gripper	146

69.1 Clinical & Functional Specifications	146

69.2 Engineering Technical Specifications	146

69.3 Standard Accessories	146

69.4 Installation Requirements	146

69.5 Documentation	146

69.6 Training	146

69.7 Warranty	146

70. Stack and Sort Board	147

70.1 Clinical & Functional Specifications	147

70.2 Engineering Technical Specifications	147

70.3 Standard Accessories	147

70.4 Installation Requirements	147

70.5 Documentation	147

70.6 Training	147

70.7 Warranty	147

71. Velcro Board	148

71.1 Clinical & Functional Specifications	148

71.2 Engineering Technical Specifications	148

71.3 Standard Accessories	148

71.4 Installation Requirements	148

71.5 Documentation	148

71.6 Training	148

71.7 Warranty	148

72. ADL Board (Activities of Daily Living Training Board)	149

72.1 Clinical & Functional Specifications	149

72.2 Engineering Technical Specifications	149

72.3 Standard Accessories	149

72.4 Installation Requirements	149

72.5 Documentation	149

72.6 Training	149

72.7 Warranty	149

73. Ring Tower & Ring Toss	151

73.1 Clinical & Functional Specifications	151

73.2 Engineering Technical Specifications	151

73.3 Standard Accessories	151

73.4 Installation Requirements	151

73.5 Documentation	151

73.6 Training	151

73.7 Warranty	151

74. Blocks, Cones & Musical Toys	152

74.1 Clinical & Functional Specifications	152

74.2 Engineering Technical Specifications	152

74.3 Standard Accessories	152

74.4 Installation Requirements	152

74.5 Documentation	152

74.6 Training	152

74.7 Warranty	152

75. Squeezing Ball for Hand Activities	154

75.1 Clinical & Functional Specifications	154

75.2 Engineering Technical Specifications	154

75.3 Standard Accessories	154

75.4 Installation Requirements	154

75.5 Documentation	154

75.6 Training	154

75.7 Warranty	154

76. Needle Board (Threading Board)	155

76.1 Clinical & Functional Specifications	155

76.2 Engineering Technical Specifications	155

76.3 Standard Accessories	155

76.4 Installation Requirements	155

76.5 Documentation	155

76.6 Training	155

76.7 Warranty	155

77. Beads and Cotton Balls	156

77.1 Clinical & Functional Specifications	156

77.2 Engineering Technical Specifications	156

77.3 Standard Accessories	156

77.4 Installation Requirements	156

77.5 Documentation	156

77.6 Training	156

77.7 Warranty	156

78. Mat Set (Therapy Mats)	157

78.1 Clinical & Functional Specifications	157

78.2 Engineering Technical Specifications	157

78.3 Standard Accessories	157

78.4 Installation Requirements	157

78.5 Documentation	157

78.6 Training	157

78.7 Warranty	157

79. Exercise Mats — Floor-Based (Yoga/Pilates Style)	159

79.1 Clinical & Functional Specifications	159

79.2 Engineering Technical Specifications	159

79.3 Standard Accessories	159

79.4 Installation Requirements	159

79.5 Documentation	159

79.6 Training	159

79.7 Warranty	159

ELECTROTHERAPY & ELECTROPHYSICAL AGENTS

# 1. Therapeutic Ultrasound Unit

Quantity: 2

## 1.1 Clinical & Functional Specifications

The equipment shall be a microprocessor-controlled therapeutic ultrasound unit intended for the deep and superficial thermal and non-thermal treatment of musculoskeletal disorders, soft-tissue injuries, tendinopathies, joint contractures, scar tissue, and post-operative rehabilitation. The unit shall deliver continuous and pulsed ultrasound output through interchangeable applicators. It shall permit therapist selection of treatment frequency, intensity, duty cycle, and treatment duration, with automatic timer control and acoustic-coupling detection to inhibit output when coupling is inadequate. The unit shall be suitable for use in physiotherapy departments, outpatient rehabilitation clinics, and sports-medicine facilities.

## 1.2 Engineering Technical Specifications

- Display: Color LCD/TFT or equivalent backlit digital display, minimum 3.5-inch diagonal, showing all treatment parameters simultaneously; membrane keypad or capacitive touch interface.

- Operating Frequencies: Dual selectable ultrasound frequencies of 1 MHz and 3 MHz (±10%).

- Output Modes: Continuous and Pulsed modes; selectable duty cycles including at least 10%, 20%, 50%, and 100%; pulse repetition frequency selectable within 16—150 Hz range or equivalent.

- Output Intensity: Adjustable from 0.1 W/cm² up to a minimum of 2.0 W/cm² (SATA), in increments of 0.1 W/cm² or finer.

- Beam Non-Uniformity Ratio (BNR): Maximum 6:1 or better, compliant with IEC 60601-2-5.

- Effective Radiating Area (ERA): Standard applicator minimum 5 cm² (±20%); small applicator minimum 1 cm² (±20%).

- Timer: Selectable treatment duration 1—30 minutes minimum, in 1-minute increments; automatic shut-off at end of cycle.

- Acoustic Coupling Detection: Automatic pause of output when applicator loses contact with treatment surface.

- Memory/Presets: Minimum 10 user-configurable clinical protocol presets.

- Power Supply: 100—240 VAC, 50/60 Hz, single-phase.

- Power Consumption: Approximately 50—100 VA.

- Weight: Approximately 2—5 kg (tabletop configuration).

- Dimensions: Compact tabletop footprint suitable for physiotherapy trolley.

- Environmental: Operating temperature 10—40 °C; relative humidity 30—75% non-condensing.

- Applicable Standards: IEC 60601-1 (general safety), IEC 60601-1-2 (EMC), IEC 60601-2-5 (particular requirements for ultrasound therapy equipment), ISO 13485, CE MDR or FDA clearance.

## 1.3 Standard Accessories

- Multi-frequency ultrasound applicator probe (1 MHz / 3 MHz): Sealed waterproof construction (minimum IPX7), ERA ≥ 5 cm², BNR ≤ 6:1, ergonomic housing, hospital-grade connector, cable length minimum 2 m.

- Micro-ultrasound applicator probe: ERA ≥ 1 cm² for small anatomical regions (hand, face, TMJ); waterproof; matching connector.

- Waterproof probe holder cradle: Integrated into chassis; corrosion-resistant material; compatible with all supplied applicators.

- Main control console panel: Fluid-resistant membrane or touch interface; digital display of frequency, intensity, mode, duty cycle, elapsed/remaining time.

- Acoustic coupling gel: Hypoallergenic, water-soluble, non-staining, non-sensitizing, minimum 250 mL bottle; compatible with reusable and single-use applications.

- AC power supply adapter with medical-grade cord: IEC 60601-1 compliant; country-specific hospital-grade plug; cable length minimum 2 m.

## 1.4 Installation Requirements

Single-phase 100—240 VAC, 50/60 Hz outlet with proper earthing; hospital-grade socket; UPS-compatible; ambient temperature 10—40 °C; tabletop or trolley placement; no special HVAC or medical gases required.

## 1.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate, Electrical Safety Test Report (IEC 62353), Certificate of Conformity, Warranty Certificate, PM Schedule — all in English.

## 1.6 Training

On-site clinical application training for physiotherapists, biomedical engineering orientation, PM and basic troubleshooting training.

## 1.7 Warranty

Minimum two (2) years comprehensive warranty covering labour, spare parts, PM visits, breakdown maintenance, software updates, and technical support.

# 2. Short-Wave Diathermy Unit

Quantity: 1

## 2.1 Clinical & Functional Specifications

The equipment shall be a continuous and pulsed short-wave diathermy unit for deep heating of subcutaneous tissues, muscles, joints, and periarticular structures for treatment of chronic musculoskeletal conditions, subacute inflammation, joint stiffness, and deep-tissue pain. The unit shall operate at the internationally allocated ISM frequency of 27.12 MHz and shall support both inductive (coil/drum) and capacitive electrode application methods.

## 2.2 Engineering Technical Specifications

- Operating Frequency: 27.12 MHz (ISM band), ±0.6 MHz tolerance per ITU allocation.

- Modes: Continuous Short-Wave (CSWD) and Pulsed Short-Wave (PSWD).

- Output Power (Continuous): Adjustable from 0 up to minimum 400 W.

- Output Power (Pulsed, mean): Adjustable up to minimum 200 W mean power.

- Pulse Duration: Selectable within 65—400 µs range or equivalent.

- Pulse Repetition Rate: Selectable within 15—200 Hz range or equivalent.

- Timer: 1—30 minutes minimum, automatic shut-off.

- Display: Backlit digital display of frequency, power, mode, and time; minimum 3.5-inch diagonal.

- Tuning: Automatic or manual patient-circuit tuning with visual indicator.

- Presets: Minimum 10 configurable clinical programs.

- Power Supply: 220—240 VAC, 50/60 Hz.

- Power Consumption: Approximately 800—1500 VA.

- Configuration: Mobile console with lockable casters.

- Environmental: Operating temperature 10—40 °C; humidity 30—75%.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-3 (short-wave therapy), ITU frequency allocation, ISO 13485, CE MDR/FDA.

## 2.3 Standard Accessories

- Inductive rubber plate electrodes (Schliephake style), 1 pair: Flexible insulated rubber, medium/large sizes, high-voltage insulation compatible with 27.12 MHz output.

- Capacitive flexible pad electrodes, 1 pair: With integrated high-frequency cables and safety plugs.

- Multi-joint articulated electrode support arms, 1 pair: Non-conductive dielectric material at electrode end; mechanical friction locks; multi-axis positioning; mounted on console.

- High-frequency insulated cables, 1 pair: Rated for full continuous output power; minimum 1.5 m; shielded per IEC EMC requirements.

- Neon tuning indicator lamp: For manual circuit verification.

- Mobile console cabinet: Heavy-duty steel frame; lockable anti-static casters; integrated electrode arm mounts.

## 2.4 Installation Requirements

Dedicated 220—240 VAC, 50/60 Hz socket with proper earthing; EMC-compliant room preferably away from sensitive electronic monitoring equipment; minimum 3 m clearance from other electromedical devices during operation; non-metallic patient couch; ambient temperature 10—40 °C.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test & Calibration Certificate, Electrical Safety Report (IEC 62353), EMC Test Report, PM Schedule.

## 2.6 Training

Clinical application training, biomedical engineering training, PM and troubleshooting training.

## 2.7 Warranty

Minimum two (2) years comprehensive warranty (labour, spares, PM, breakdown, software updates).

# 3. TENS Unit

Quantity: 4

## 3.1 Clinical & Functional Specifications

Portable, battery-operated dual-channel Transcutaneous Electrical Nerve Stimulator for symptomatic relief of acute and chronic pain through non-invasive surface electrode stimulation. Shall support conventional (high-frequency), acupuncture-like (low-frequency), burst, and modulated modes, with independent channel control.

## 3.2 Engineering Technical Specifications

- Channels: Minimum 2 independent isolated output channels.

- Waveform: Symmetric or asymmetric biphasic rectangular pulse, zero net DC.

- Pulse Amplitude: Adjustable 0—100 mA into 500 Ω load (minimum).

- Pulse Width: Selectable within 50—300 µs range.

- Pulse Frequency: Selectable within 1—150 Hz range.

- Modes: Continuous, Burst, Modulated (frequency and/or amplitude modulation), and preset pain-management programs.

- Preset Programs: Minimum 10 configurable clinical protocols.

- Timer: 1—60 minutes selectable, or continuous.

- Display: Backlit LCD showing intensity, mode, frequency, pulse width, timer, battery status.

- Battery: Rechargeable NiMH or Li-ion; minimum 10 hours continuous therapy operation per charge; charging time maximum 8 hours.

- Power: Battery-operated; charger 100—240 VAC, 50/60 Hz.

- Weight: Approximately 150—400 g (handheld).

- Ingress Protection: Minimum IP22.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-10, ISO 13485, CE MDR/FDA.

## 3.3 Standard Accessories

- Dual-channel patient lead wires, 2 pairs: Medical-grade shielded cables, minimum 1.2 m, 2 mm safety pin connectors, color-coded.

- Reusable self-adhesive surface electrodes, 50×50 mm, pack of 4: Hypoallergenic conductive hydrogel, minimum 20 applications per electrode, latex-free.

- Protective silicone rubber boot with belt clip: Impact-resistant, autoclave-incompatible but wipe-cleanable with hospital disinfectants.

- Rechargeable NiMH battery set with desktop dual charging station: Overcharge and short-circuit protected; universal input adapter.

- Rigid plastic transport pouch: Foam-lined; accommodates unit, leads, electrodes, and accessories.

## 3.4 Installation Requirements

No fixed installation; charger requires 100—240 VAC, 50/60 Hz outlet.

## 3.5 Documentation

User Manual, Quick Reference Guide, Factory Test Certificate, Electrical Safety Report, Warranty Certificate.

## 3.6 Training

Clinical application training and patient-education orientation.

## 3.7 Warranty

Minimum two (2) years comprehensive warranty.

# 4. Interferential Current (IFC) Unit

Quantity: 1

## 4.1 Clinical & Functional Specifications

Mains-operated four-pole (quadripolar) interferential current therapy unit for deep pain relief, muscle stimulation, edema reduction, and improvement of local circulation using medium-frequency crossed-current interference. Shall support bipolar (pre-modulated) and quadripolar (true interferential) modes with vector/sweep functions.

## 4.2 Engineering Technical Specifications

- Channels: Two independent isolated output channels supporting bipolar and quadripolar operation.

- Carrier Frequency: Selectable, typically 2 kHz, 4 kHz, 8 kHz, or equivalent medium-frequency range.

- Beat/AMF Frequency: Adjustable within 1—200 Hz range.

- Sweep (Spectrum): Selectable sweep patterns 1/1 s, 6/6 s, 1/30/1/30 s, or equivalent.

- Vector Rotation: Manual and automatic modes.

- Output Current: 0—100 mA per channel into 500 Ω.

- Modes: Bipolar, Quadripolar Fixed, Quadripolar with Vector, Pre-modulated.

- Timer: 1—30 minutes minimum.

- Presets: Minimum 20 configurable clinical programs.

- Vacuum Unit: Integrated low-noise pump with adjustable continuous and pulsed suction, negative pressure range approximately 0 to −0.6 bar.

- Display: Color LCD minimum 4.3-inch; simultaneous display of all parameters.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-10, ISO 13485, CE MDR/FDA.

## 4.3 Standard Accessories

- Four-pole patient lead wire cables: Color-coded safety plugs, shielded, minimum 1.5 m.

- Reusable vacuum suction cup electrodes with sponges, set of 4: Conductive rubber cup electrodes, silicone tubing, reusable washable sponges.

- Integrated vacuum pump module: Continuously adjustable suction, low-noise operation (\< 55 dBA typical), fluid-trap protection.

- Flexible rubber plate electrodes with hook-and-loop straps, set of 4: Carbonized silicone/rubber, various sizes, reusable, wipe-cleanable.

- Medical-grade main power cord: Country-specific hospital plug, minimum 2 m.

## 4.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; tabletop or trolley placement; ambient 10—40 °C.

## 4.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Electrical Safety Report (IEC 62353), PM Schedule.

## 4.6 Training

Clinical, biomedical, PM, and troubleshooting training.

## 4.7 Warranty

Minimum two (2) years comprehensive warranty.

# 5. Laser Therapy Unit

Quantity: 1

## 5.1 Clinical & Functional Specifications

Low-Level Laser Therapy (LLLT) / Photobiomodulation unit for the treatment of musculoskeletal pain, soft-tissue injuries, wound healing, and trigger points. Shall support both cluster (multi-diode) and single-diode probes with red and infrared wavelengths, continuous and pulsed emission modes.

## 5.2 Engineering Technical Specifications

- Laser Classification: Class 3B per IEC 60825-1.

- Wavelengths: Red (approximately 630—680 nm) and Infrared (approximately 780—905 nm), selectable.

- Emission Modes: Continuous Wave (CW) and Pulsed, with adjustable pulse frequency 1 Hz—5000 Hz or equivalent.

- Output Power (single-diode): Adjustable up to minimum 500 mW.

- Cluster Probe Total Output: Multi-diode combined output typically 1—2 W or equivalent.

- Dose Setting: Adjustable joules per point or per cm² (0.1—100 J/cm² or equivalent), with automatic time calculation.

- Timer: Automatic based on dose and area; manual override 1—60 minutes.

- Presets: Minimum 20 clinical protocols.

- Display: Color LCD minimum 4-inch; touchscreen or membrane keypad.

- Safety: Remote interlock connector, key-switch activation, emission-warning indicators.

- Power Supply: 100—240 VAC, 50/60 Hz, or rechargeable battery pack.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-22 (laser equipment), IEC 60825-1 (laser safety), ISO 13485, CE MDR/FDA.

## 5.3 Standard Accessories

- Multi-diode cluster laser applicator probe: Combined red + infrared array; ergonomic housing; integrated emission indicator; cable minimum 2 m.

- Single-diode laser pointer probe: For trigger point/acupuncture application; matched wavelength; integrated safety button.

- Patient safety laser goggles, 1 pair: Optical density appropriate to output wavelengths; CE/EN 207 certified.

- Therapist safety laser goggles, 1 pair: OD ≥ 5 at operating wavelengths; EN 207 certified.

- Remote interlock connector plug: For door/room safety interlock integration.

- Rigid padded protective transit case: Foam-cut compartments for all probes, goggles, and unit.

## 5.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; treatment room with restricted-access laser warning signage compliant with IEC 60825-1; window covering if applicable; door interlock optional.

## 5.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate for output power, Electrical Safety Report (IEC 62353), Laser Safety Report, PM Schedule.

## 5.6 Training

Clinical application training, laser safety training for all users, biomedical PM and troubleshooting training.

## 5.7 Warranty

Minimum two (2) years comprehensive warranty including laser diode module.

# 6. Electrical Muscle Stimulator (EMS)

Quantity: 2

## 6.1 Clinical & Functional Specifications

Neuromuscular electrical stimulator (NMES) for muscle re-education, prevention of disuse atrophy, spasticity management, and strengthening of denervated or partially denervated muscles. Shall provide symmetric biphasic and monophasic stimulation with adjustable ramp-up/down and on/off cycling.

## 6.2 Engineering Technical Specifications

- Channels: Minimum 2 independent isolated channels.

- Waveform: Symmetric/asymmetric biphasic rectangular; Russian (2500 Hz burst-modulated) mode selectable.

- Output Amplitude: 0—100 mA into 500 Ω per channel.

- Pulse Width: Adjustable 50—400 µs.

- Frequency: 1—120 Hz adjustable.

- On/Off Cycling: Selectable within 1—60 s ON and 1—60 s OFF.

- Ramp Up/Down: 0—10 s adjustable.

- Presets: Minimum 10 NMES protocols.

- Timer: 1—60 minutes.

- Display: Backlit LCD showing all parameters.

- Battery: Rechargeable Li-ion or NiMH; minimum 6 hours therapy per charge; charging time maximum 8 hours.

- Power: AC adapter 100—240 VAC, 50/60 Hz + battery.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-10, ISO 13485, CE MDR/FDA.

## 6.3 Standard Accessories

- Dual-channel output lead cables, 2 pairs: Insulated snap-terminal connectors, shielded, minimum 1.2 m.

- Reusable carbonized rubber electrodes, 1 pair: Flexible, various sizes (typically 40×60 mm and 50×90 mm), long-life reusable.

- Absorbent sponge jackets, 1 pair: Washable, sized to match rubber electrodes.

- Elastic fixation bands, set of 2: Hook-and-loop adjustable, latex-free, various sizes.

- AC adapter + rechargeable battery pack: Medical-grade, IEC 60601-1 compliant.

## 6.4 Installation Requirements

Standard mains outlet 100—240 VAC, 50/60 Hz for charging; no special utilities.

## 6.5 Documentation

User Manual, Service Manual, Factory Test Certificate, Electrical Safety Report, PM Schedule.

## 6.6 Training

Clinical, biomedical, PM training.

## 6.7 Warranty

Minimum two (2) years comprehensive warranty.

# 7. General Electrical Stimulator (Faradic / Galvanic)

Quantity: 1

## 7.1 Clinical & Functional Specifications

Multi-mode electrical stimulator supporting Faradic (surged, alternating) and Galvanic (direct current, interrupted and uninterrupted) stimulation for treatment of denervated muscles, iontophoresis, muscle testing, and manual point stimulation. Shall support both surface plate electrode and handheld point-probe application.

## 7.2 Engineering Technical Specifications

- Modes: Galvanic (continuous DC), Interrupted Galvanic, Faradic (surged), Sinusoidal Surged, and manual/interrupted stimulation.

- Output Current: 0—80 mA adjustable into 500 Ω.

- Pulse Duration (Galvanic interrupted): Adjustable 1—1000 ms.

- Faradic Frequency: Approximately 50—100 Hz surged.

- Surge Rate: Adjustable, typically 6—30 surges/minute.

- Iontophoresis Mode: Continuous DC with polarity reversal.

- Timer: 1—30 minutes minimum.

- Display: Digital LCD showing mode, current, timer.

- Power: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-10, ISO 13485, CE MDR/FDA.

## 7.3 Standard Accessories

- Multi-functional patient cables: Compatible with Faradic and Galvanic outputs; safety-plug terminated; minimum 1.5 m.

- Handheld point-stimulation probe with manual interrupt button: Ergonomic; insulated; interchangeable metal tip; cable minimum 1.5 m.

- Flexible conductive metal plate electrodes, set of 4: Tin or stainless steel; various sizes (small, medium, large).

- Moist fabric electrode envelopes / sponge sleeves: Washable, sized to match plate electrodes.

- Adjustable stretch fixation straps set: Latex-free elastic with hook-and-loop closure.

## 7.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; tabletop placement.

## 7.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Electrical Safety Report, PM Schedule.

## 7.6 Training

Clinical, biomedical, PM and troubleshooting.

## 7.7 Warranty

Minimum two (2) years comprehensive warranty.

# 8. Shockwave Therapy Unit (Radial)

Quantity: 1

## 8.1 Clinical & Functional Specifications

Radial extracorporeal shockwave therapy (rESWT) unit generating pneumatically-driven mechanical pressure waves through a projectile-transmitter mechanism, for treatment of chronic tendinopathies (plantar fasciitis, tennis/golfer's elbow, calcific shoulder tendinopathy, patellar tendinopathy), trigger points, and myofascial pain syndromes.

## 8.2 Engineering Technical Specifications

- Wave Generation: Ballistic/pneumatic (radial) principle.

- Pressure/Energy Range: Adjustable 1.0—5.0 bar minimum (equivalent to approximately 0.05—0.55 mJ/mm² energy flux density).

- Frequency: Adjustable 1—21 Hz.

- Shot Count: Programmable up to minimum 10,000 shots per session; total handpiece life minimum 2,000,000 shots between rebuilds.

- Presets: Minimum 15 pre-programmed anatomical/indication-based protocols; user-editable.

- Compressor: Integrated low-noise (\< 65 dBA typical), oil-free.

- Display: Color touchscreen minimum 7-inch diagonal, resolution minimum 800 × 480.

- Data: Session logging with patient ID, protocol, shot count, and energy.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Configuration: Mobile cart with lockable casters.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR/FDA.

## 8.3 Standard Accessories

- Radial shockwave handpiece applicator: Ergonomic, weight approximately 500—900 g; user-rebuildable projectile assembly; standard connector to console.

- Interchangeable transmitter tips (9 mm, 15 mm, 25 mm): Hardened stainless steel or equivalent; heat-treated; user-replaceable; autoclave or chemical disinfectant compatible.

- Internal projectile rebuild kit: Replacement guide tubes, projectiles, and O-rings sufficient for one full maintenance cycle.

- Ergonomic waterproof foot switch (IPX7 minimum): For hands-free shot triggering.

- Heavy-duty mobile cart: Steel frame, four locking casters, integrated handpiece holster.

- Acoustic transmission gel starter bottle: Water-soluble, hypoallergenic, minimum 250 mL.

## 8.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; ambient 10—40 °C; room size sufficient for mobile cart maneuvering; no compressed air or medical gases required.

## 8.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Electrical Safety Report (IEC 62353), Factory Test Report, PM Schedule.

## 8.6 Training

Clinical application training (indications, protocols, safety), biomedical PM training (handpiece rebuild), troubleshooting.

## 8.7 Warranty

Minimum two (2) years comprehensive warranty including one handpiece rebuild service.

# 9. Transcranial Direct Current Stimulation (tDCS) Unit

Quantity: 1

## 9.1 Clinical & Functional Specifications

Constant-current transcranial direct current stimulation device for non-invasive neuromodulation used in research and clinical rehabilitation of stroke, depression, chronic pain, and cognitive rehabilitation. Shall deliver low-intensity DC current between scalp electrodes with precise current control, ramp-up/ramp-down, and impedance monitoring.

## 9.2 Engineering Technical Specifications

- Output Current: Constant-current mode, adjustable 0.5—2.0 mA (up to 4 mA in some configurations), accuracy ±5%.

- Session Duration: Programmable 1—30 minutes.

- Ramp-Up / Ramp-Down: Adjustable 5—60 s.

- Impedance Monitoring: Continuous real-time; automatic shutdown if impedance exceeds safe threshold (typically \> 20 kΩ).

- Sham Mode: Available for double-blind protocols.

- Display: LCD showing current, time, impedance, battery status.

- Data Log: Session storage with export via USB.

- Power: Internal rechargeable battery, minimum 8 hours operation per charge; charging via medical-grade adapter.

- Weight: Approximately 200—500 g (portable).

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-10, ISO 13485, CE MDR/FDA.

## 9.3 Standard Accessories

- Adjustable silicone rubber head straps, set of 2: Cross-hatched configuration, latex-free, adjustable for adult head sizes.

- Conductive rubber electrode plates, set of 2: Typical size 25—35 cm²; conductive carbon-rubber; reusable.

- Absorbent saline sponge pads, starter set of 10: Compatible with electrode dimensions; reusable.

- Precision current-output cables with safety plugs: Shielded; minimum 1.5 m.

- Standard multimeter calibration verification tool: For output current verification.

- Rigid plastic padded carrying case: Foam-cut, holds all components.

## 9.4 Installation Requirements

No fixed installation; charger 100—240 VAC, 50/60 Hz.

## 9.5 Documentation

User Manual, Service Manual, Factory Calibration Certificate (output current), Electrical Safety Report, PM Schedule.

## 9.6 Training

Clinical application training (electrode montage, protocols, safety), biomedical PM training.

## 9.7 Warranty

Minimum two (2) years comprehensive warranty.

BIOFEEDBACK, TRACTION & THERMAL THERAPY

# 10. Biofeedback Unit

Quantity: 1

## 10.1 Clinical & Functional Specifications

The equipment shall be a multi-modality computerized biofeedback system for physiological signal acquisition, real-time visualization, and patient-directed self-regulation training. The unit shall support surface electromyography (sEMG), peripheral skin temperature, and galvanic skin response (GSR) measurements for use in neuromuscular re-education, pelvic floor rehabilitation, stress management, relaxation therapy, and psychophysiological training. The system shall present feedback through graphical, auditory, and interactive game-based displays, and shall permit session recording, patient database management, and progress reporting.

## 10.2 Engineering Technical Specifications

- Channels: Minimum 2 sEMG channels + 1 skin temperature + 1 GSR channel, simultaneously acquired.

- sEMG Input Range: 0—2000 µV RMS minimum; resolution ≤ 0.1 µV.

- sEMG Bandwidth: Minimum 20—500 Hz, with selectable notch filter (50/60 Hz).

- sEMG CMRR: Minimum 100 dB.

- Sampling Rate: Minimum 1000 samples/second/channel.

- A/D Resolution: Minimum 16-bit.

- Skin Temperature Range: 15—45 °C, accuracy ±0.1 °C.

- GSR Range: 0.1—100 µS, resolution ≤ 0.01 µS.

- Isolation: Patient-applied parts isolated per IEC 60601-1 Type BF.

- Interface: USB 2.0 or higher to host PC; wireless connectivity optional.

- Software: Multi-user database, real-time graphs/bars/games, threshold-triggered feedback, session recording and export (PDF, CSV), configurable protocols.

- Host PC Requirements (if separate): Windows-compatible; minimum specifications documented.

- Power: Battery-operated patient unit (minimum 8 hours) + USB power for interface.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-40 (EMG equipment), IEC 62304 (software), ISO 13485, CE MDR/FDA.

## 10.3 Standard Accessories

- sEMG sensor cables with snap adapters, 1 pair: Shielded low-noise cables, minimum 1.5 m, standard snap connectors compatible with disposable/reusable Ag/AgCl electrodes.

- Skin temperature sensor probe: NTC thermistor or equivalent; response time ≤ 3 s; cable minimum 1.5 m; hypoallergenic tip.

- GSR finger band sensors, 1 pair: Reusable, hook-and-loop adjustable, dry-contact silver electrodes.

- Biofeedback training software license: Multi-user, graphical feedback and gaming interface, therapist configuration, session archiving, report generation.

- USB data interface cable: Shielded, minimum 2 m, USB 2.0 or higher.

## 10.4 Installation Requirements

Host PC or laptop (supplied or user-provided per tender); 100—240 VAC, 50/60 Hz mains for PC/charger; quiet clinical room with low EMI environment; ambient 10—40 °C.

## 10.5 Documentation

User Manual, Software Manual, Service Manual, Factory Calibration Certificate, Electrical Safety Report (IEC 62353), Software Validation Certificate, PM Schedule.

## 10.6 Training

Clinical application training (protocol design, interpretation), software training, biomedical PM training.

## 10.7 Warranty

Minimum two (2) years comprehensive warranty including software updates.

# 11. Cervical Traction Unit

Quantity: 1

## 11.1 Clinical & Functional Specifications

Microprocessor-controlled motorized cervical traction unit for the mechanical decompression of cervical spine structures in the management of cervical radiculopathy, disc herniation, cervical spondylosis, muscle spasm, and postural cervical pain. Shall deliver continuous (static) and intermittent (cyclic) traction with programmable force, hold/rest times, and treatment duration. Patient-held emergency stop shall immediately release traction force.

## 11.2 Engineering Technical Specifications

- Traction Force Range: 0—25 kgf minimum (0—55 lbf), adjustable in increments ≤ 0.5 kgf.

- Force Accuracy: ±1 kgf or ±5%, whichever is greater.

- Modes: Continuous (static) and Intermittent (cyclic).

- Hold Time: Adjustable 1—99 s.

- Rest Time: Adjustable 1—99 s.

- Rest Force: Adjustable 0—50% of peak force.

- Ramp-Up / Ramp-Down: Programmable.

- Treatment Duration: 1—60 minutes.

- Presets: Minimum 10 configurable clinical protocols.

- Display: Backlit LCD minimum 4-inch showing force, time, mode, cycle.

- Safety: Patient-held emergency stop; mechanical overload protection; automatic force release on power failure.

- Power: 100—240 VAC, 50/60 Hz.

- Mounting: Wall-bracket or heavy-duty stand configuration.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR/FDA.

## 11.3 Standard Accessories

- Motorized digital traction head module: Precision servo/stepper motor drive with load-cell feedback; enclosed housing.

- Padded cervical halter harness with adjustable chin-straps: Latex-free, wipe-cleanable, multi-size adjustable (adult and adolescent).

- Universal spreader bar hanger: Corrosion-resistant metal, standard hook attachment.

- Handheld emergency-stop switch with coiled cable: Patient-operated, red mushroom button, cable minimum 1.5 m coiled.

- Wall-mounting bracket / heavy-duty table mount: Steel construction, load-tested to minimum 3× working load.

## 11.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; structural wall capable of supporting mount (concrete or reinforced), or heavy-duty stand; adjacent treatment chair/couch of adequate stability; ambient 10—40 °C.

## 11.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Calibration Certificate (force output), Electrical Safety Report (IEC 62353), Installation Report, PM Schedule.

## 11.6 Training

Clinical training (protocols, contraindications, patient positioning), biomedical PM and calibration training.

## 11.7 Warranty

Minimum two (2) years comprehensive warranty including motor and load cell.

# 12. Lumbar Traction Unit

Quantity: 1

## 12.1 Clinical & Functional Specifications

Microprocessor-controlled motorized lumbar traction unit integrated with split traction table for the mechanical decompression of lumbar spine structures in the management of lumbar disc herniation, radiculopathy, facet joint dysfunction, and chronic low back pain. Shall provide continuous and intermittent traction with programmable force, hold/rest cycling, and patient safety controls.

## 12.2 Engineering Technical Specifications

- Traction Force Range: 0—90 kgf minimum (0—200 lbf), adjustable in increments ≤ 1 kgf.

- Force Accuracy: ±2 kgf or ±5%.

- Modes: Continuous and Intermittent.

- Hold Time: 1—99 s adjustable.

- Rest Time: 1—99 s adjustable.

- Rest Force: 0—50% of peak.

- Ramp-Up / Ramp-Down: Programmable.

- Duration: 1—60 minutes.

- Presets: Minimum 10 configurable protocols.

- Table: Three-section (thoracic fixed, lumbar sliding, leg section) with low-friction sliding mechanism; patient weight capacity minimum 180 kg.

- Display: Backlit LCD minimum 4-inch.

- Safety: Patient-held emergency stop, mechanical over-force cut-out, auto-release on power loss.

- Power: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, ISO 13485, CE MDR/FDA.

## 12.3 Standard Accessories

- Motorized microprocessor traction head console: Load-cell feedback, precision drive, digital control.

- Padded lumbar/pelvic harness belt with rapid securing buckles: Heavy-duty webbing, latex-free padding, adjustable for varied patient sizes, wipe-cleanable.

- Thoracic stabilization chest harness belt: Heavy-duty, adjustable, latex-free padding.

- Patient emergency-stop remote switch: Handheld, red mushroom button, coiled cable minimum 1.5 m.

- Three-section sliding traction table: Padded, upholstered with wipe-clean fluid-proof vinyl, adjustable split sections, integrated harness anchor points.

## 12.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz outlet with earthing; floor space minimum 2.5 m × 2 m; level floor, load-bearing capacity for equipment + patient (minimum 300 kg total); ambient 10—40 °C.

## 12.5 Documentation

User Manual, Service Manual, Parts Catalogue, Calibration Certificate (force), Electrical Safety Report, Installation & Commissioning Report, PM Schedule.

## 12.6 Training

Clinical, biomedical PM and calibration training.

## 12.7 Warranty

Minimum two (2) years comprehensive warranty including motor, load cell, and table mechanism.

# 13. Paraffin Wax Bath Unit

Quantity: 1

## 13.1 Clinical & Functional Specifications

Thermostatically-controlled paraffin wax heating and holding bath for superficial heat therapy of the hands, wrists, feet, and small joints, particularly in the management of rheumatoid arthritis, osteoarthritis, scleroderma, joint stiffness, and pre-exercise tissue conditioning. Shall maintain wax within a therapeutically safe and clinically effective temperature range with automatic thermostatic regulation and over-temperature protection.

## 13.2 Engineering Technical Specifications

- Wax Capacity: Minimum 3—5 kg (approximately 4—7 litres liquefied volume).

- Temperature Range: Adjustable within 45—60 °C.

- Thermostatic Accuracy: ±1.5 °C.

- Heat-Up Time (from cold to operating): Approximately 90—180 minutes for full load.

- Construction: Double-walled insulated stainless steel tank; external polymer or coated steel housing.

- Safety: Over-temperature cut-off (independent thermostat), safety lid, thermal insulation of external surfaces to prevent burns.

- Controls: Thermostat dial or digital control; power-on indicator; ready indicator.

- Power: 220—240 VAC, 50/60 Hz.

- Power Consumption: Approximately 200—500 W.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-35 (heating devices using blankets, pads and mattresses — where applicable) or equivalent thermal therapy standard, ISO 13485, CE MDR/FDA.

## 13.3 Standard Accessories

- Insulated stainless steel double-walled tank with safety lid: Food-grade stainless steel interior, corrosion-resistant, easy-clean.

- Protective plastic bottom grid: Perforated, prevents direct skin contact with heating element base.

- Insulated terry-cloth mitts, 1 pair: Multi-layer thermal-insulating cotton terry, hand-size, washable.

- Insulated terry-cloth booties, 1 pair: Multi-layer thermal-insulating, foot-size, washable.

- Plastic liner bags, 50 pieces starter pack: Food-grade polyethylene, disposable, sized to cover wax-dipped hand/foot.

## 13.4 Installation Requirements

220—240 VAC, 50/60 Hz outlet with earthing; stable non-flammable countertop; adequate ventilation; ambient 10—35 °C; away from flammable materials; fire extinguisher (Class B) in vicinity per general clinical safety practice.

## 13.5 Documentation

User Manual, Service Manual, Factory Test Certificate, Electrical Safety Report (IEC 62353), MSDS for supplied paraffin (if included), PM Schedule.

## 13.6 Training

Clinical application training (dipping technique, contraindications, hygiene), biomedical PM training.

## 13.7 Warranty

Minimum two (2) years comprehensive warranty.

# 14. Infrared Radiator

Quantity: 2

## 14.1 Clinical & Functional Specifications

Mobile infrared radiant heat therapy unit for superficial heat application in the management of muscle spasm, subacute and chronic musculoskeletal pain, and pre-exercise warming. Shall provide adjustable-height, multi-axis positioning of the infrared source with timer-controlled operation and protective safety mesh over the emitter.

## 14.2 Engineering Technical Specifications

- Emitter Type: Near/mid-infrared incandescent or halogen bulb; wavelength range approximately 780 nm—3000 nm.

- Bulb Power Rating: Selectable 150 W or 250 W (or equivalent), with replacement bulb compatibility.

- Reflector: Polished aluminum or equivalent parabolic reflector for directional beam.

- Arm Reach: Spring-balanced articulated arm; minimum 1 m reach from stand column; multi-axis (minimum 3 degrees of freedom).

- Height Adjustment: Column-mounted; adjustable working height range approximately 90—170 cm.

- Timer: Mechanical or digital countdown timer, 1—60 minutes, with automatic shut-off.

- Safety: Protective wire mesh guard over bulb face, thermal cut-off on housing, tip-resistant base.

- Base: Heavy-duty 5-caster mobile base with minimum 2 lockable casters.

- Power: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-57 (non-laser light sources for therapeutic use), ISO 13485, CE MDR/FDA.

## 14.3 Standard Accessories

- Mobile base stand with lockable casters: Heavy-duty five-star base, anti-tip design, smooth-rolling casters, minimum 2 with locks.

- Spring-balanced articulated extension arm: Multi-axis, friction-lock joints, minimum 1 m reach.

- Infrared heating lamp bulb (150 W or 250 W): Manufacturer-standard replaceable bulb.

- Wire mesh safety guard: Chrome-plated or stainless steel, covers full emitter aperture.

- Countdown timer switch: Mechanical or electronic, 1—60 min, auto-off.

## 14.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with earthing; adequate floor space and clearance for arm movement; ambient 10—40 °C; away from flammable materials and oxygen-enriched atmospheres.

## 14.5 Documentation

User Manual, Service Manual, Factory Test Certificate, Electrical Safety Report (IEC 62353), PM Schedule.

## 14.6 Training

Clinical application training (distance, dosage, safety), biomedical PM training (bulb replacement).

## 14.7 Warranty

Minimum two (2) years comprehensive warranty (bulbs typically covered for defects, not consumption).

# 15. Electrical Heating Pad

Quantity: 4

## 15.1 Clinical & Functional Specifications

Flexible electrically-heated therapeutic pad for superficial local heat application in the management of muscle spasm, subacute musculoskeletal pain, menstrual pain, and pre-exercise conditioning. Shall provide multiple selectable temperature levels with automatic timed shut-off, thermostatic control, and washable protective cover.

## 15.2 Engineering Technical Specifications

- Heating Element: Flexible resistive wire grid distributed evenly across pad surface.

- Temperature Levels: Minimum 3 selectable settings (approximately low \~40 °C, medium \~50 °C, high \~60 °C).

- Thermostatic Accuracy: ±3 °C at set point.

- Safety Thermostat: Independent over-temperature cut-off (typically maximum 65 °C surface).

- Automatic Shut-Off: Timer-controlled, selectable 15/30/60/90 minutes.

- Pad Size: Standard therapeutic size, approximately 30 × 40 cm to 40 × 60 cm.

- Cover: Removable, washable soft fleece jacket.

- Cable: Minimum 2 m detachable or fixed cord with in-line controller.

- Power: 100—240 VAC, 50/60 Hz.

- Power Consumption: Approximately 60—120 W.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-35 (heating pads and blankets), ISO 13485, CE MDR/FDA.

## 15.3 Standard Accessories

- Flexible internal heating wire element with safety thermostat: Uniformly distributed resistive grid, dual-thermostat safety.

- Washable soft fleece cover jacket: Removable, machine-washable, hypoallergenic.

- Handheld controller switch: Digital or analog, 3 temperature levels, timer, indicator LEDs.

- Automated shut-off mechanism: Timer-controlled auto-off safety.

- Main power cord with surge protector plug: Country-specific hospital-grade plug, minimum 2 m.

## 15.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet; no fixed installation.

## 15.5 Documentation

User Manual, Factory Test Certificate, Electrical Safety Report (IEC 62353), Warranty Certificate.

## 15.6 Training

Clinical use, cleaning, and safety orientation.

## 15.7 Warranty

Minimum two (2) years comprehensive warranty.

# 16. Moist Heating Pad (Hydrocollator-Style Pack)

Quantity: 4

## 16.1 Clinical & Functional Specifications

Reusable hydrocollator-type moist heat pack for the delivery of moist superficial heat therapy in the management of muscle spasm, myofascial pain, joint stiffness, and preparation for manual therapy or exercise. The pack shall retain heat when removed from a heated water bath and deliver moist heat to the treatment area through insulated terry-cloth covers.

## 16.2 Engineering Technical Specifications

- Filler: Bentonite or equivalent hydrophilic thermal-retention silicate gel, sealed in canvas envelope.

- Envelope: Heavy-duty multi-layer canvas fabric, seam-reinforced, colorfast.

- Heat Retention: Minimum 20—30 minutes of therapeutic warmth after removal from 70—80 °C bath.

- Standard Sizes: Cervical (approximately 15 × 60 cm), Standard (approximately 25 × 30 cm), Oversize (approximately 40 × 60 cm) or equivalent range.

- Handles: Integrated woven fabric hanging tabs for hydrocollator storage/extraction.

- Compatibility: Compatible with standard thermostatically-controlled water-bath heaters (70—80 °C).

- Applicable Standards: ISO 13485, CE MDR/FDA general medical device requirements.

## 16.3 Standard Accessories

- Canvas envelope with bentonite gel filler: Heavy-duty stitched construction, standard therapeutic size.

- Insulated terry-cloth wrap covers, 2 pieces per pad: Multi-layered thermal-insulating cotton terry, washable, prevents skin burns.

- Overhead storage hanging hook tabs: Reinforced woven fabric loops.

- Color-coded handles for extraction safety: High-visibility tabs.

## 16.4 Installation Requirements

No electrical connection; requires hydrocollator water-bath heater (separate item if not already available). Shelf/hanging storage adjacent to heater.

## 16.5 Documentation

Product Information, Cleaning Instructions, Warranty Certificate.

## 16.6 Training

Clinical application and safety orientation (extraction, layering, skin monitoring).

## 16.7 Warranty

Minimum one (1) year against manufacturing defects; typical service life 12—24 months depending on usage.

# 17. Cold Pack

Quantity: 3

## 17.1 Clinical & Functional Specifications

Reusable freeze-resistant gel cold pack for cryotherapy application in the management of acute soft-tissue injury, post-exercise inflammation, post-operative swelling, and acute musculoskeletal pain. Shall remain flexible at freezer temperatures for anatomical contouring and shall be provided with insulated cover barriers to prevent skin damage.

## 17.2 Engineering Technical Specifications

- Filler: Non-toxic hypoallergenic freeze-resistant gel (typically water/glycol/CMC or equivalent).

- Envelope: Heavy-duty flexible vinyl or polyurethane, puncture and leak resistant.

- Flexibility at −18 °C: Remains pliable, conforms to anatomy.

- Standard Sizes: Cervical (approximately 15 × 50 cm), Standard (approximately 25 × 30 cm), Oversize (approximately 30 × 45 cm) or equivalent range.

- Cold Retention: Minimum 20 minutes of therapeutic cooling after removal from freezer (−15 to −20 °C).

- Cleaning: Wipe-clean with hospital disinfectant.

- Applicable Standards: ISO 13485, CE MDR/FDA general medical device requirements; RoHS-compliant materials.

## 17.3 Standard Accessories

- Flexible vinyl exterior pouch with non-toxic gel filler: Puncture-resistant, leak-proof, hypoallergenic.

- Insulated textile protective barrier covers, 2 pieces: Cotton or terry, washable, prevents direct skin contact and frostbite.

- Adjustable elastic compression wrap straps with hook-and-loop tabs: Latex-free, adjustable circumference.

## 17.4 Installation Requirements

Requires standard therapeutic freezer (−15 to −20 °C, separate item); storage compartment sized for pack quantity.

## 17.5 Documentation

Product Information, Cleaning Instructions, MSDS for gel filler, Warranty Certificate.

## 17.6 Training

Clinical use and safety orientation (barrier use, exposure time, skin monitoring, contraindications).

## 17.7 Warranty

Minimum one (1) year against leakage/manufacturing defects.

MOBILITY & CARDIO REHABILITATION EQUIPMENT

# 18. Electrical Tilt Table

Quantity: 1

## 18.1 Clinical & Functional Specifications

The equipment shall be a motorized electrically-operated tilt table for progressive verticalization and orthostatic tolerance training of patients with prolonged immobilization, spinal cord injury, stroke, traumatic brain injury, and cardiovascular deconditioning. The table shall permit smooth continuous adjustment of tilt angle from horizontal (0°) to full vertical (minimum 85°) with patient safety restraints, footplate weight-bearing support, and handheld remote inclination control. An emergency-stop function shall be provided. The unit shall support progressive standing training, cardiovascular conditioning, and prevention of contractures and osteoporosis.

## 18.2 Engineering Technical Specifications

- Tilt Angle Range: Continuously adjustable from 0° (horizontal) to minimum 85° (near-vertical).

- Tilt Angle Accuracy: ±2°.

- Drive Mechanism: Electric linear actuator or motorized screw drive with self-locking gearbox (no drop-back on power failure).

- Tilt Speed: Smooth continuous motion, approximately 0.5—2°/second.

- Patient Weight Capacity: Minimum 150 kg (safe working load).

- Table Length: Approximately 190—210 cm.

- Table Width: Approximately 60—75 cm.

- Height (horizontal position): Approximately 60—80 cm from floor.

- Table Surface: Upholstered with high-density foam padding, wipe-clean fluid-proof antimicrobial vinyl cover with stitched or welded seams.

- Footplate: Adjustable, non-slip surface, angle-adjustable, capable of full patient weight-bearing.

- Frame: Heavy-duty welded steel with corrosion-resistant powder-coat or equivalent finish.

- Base: Anti-static locking casters (minimum 4, at least 2 with brakes) or fixed floor base per tender preference.

- Control: Handheld wired remote pendant with up/down buttons and emergency stop; angle indicator display.

- Safety: Emergency stop button, anti-crush safety cut-out, low-voltage patient-side controls, mechanical over-travel limits.

- Power: 100—240 VAC, 50/60 Hz; optional battery backup for controlled descent on power failure.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-52 (medical beds — where applicable), ISO 13485, CE MDR/FDA.

## 18.3 Standard Accessories

- Heavy-duty steel frame chassis with anti-static locking casters: Powder-coated finish, minimum 4 casters with at least 2 locking, tip-resistant footprint.

- Motorized continuous vertical tilt adjustment drive mechanism (0° to minimum 85°): Self-locking linear actuator, smooth continuous motion, low-noise operation.

- Handheld wired remote control pendant: Ergonomic housing, up/down buttons, emergency stop, integrated angle indicator or digital display, coiled cable minimum 1.5 m.

- Adjustable angle metal footplate platform with non-slip rubber padding: Full weight-bearing capacity, tool-free angle adjustment.

- Wide patient restraint straps with quick-release buckles, set of 3: Chest, pelvis, and knee positions; heavy-duty webbing; padded contact surfaces; latex-free; wipe-cleanable.

## 18.4 Installation Requirements

Dedicated 100—240 VAC, 50/60 Hz outlet with proper earthing; floor space minimum 3 m × 2 m; level floor with adequate load-bearing capacity (minimum 350 kg total); ambient temperature 10—40 °C; UPS-compatible.

## 18.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Report (IEC 62353), Load-Test Certificate, Installation & Commissioning Report, PM Schedule.

## 18.6 Training

Clinical training (verticalization protocols, patient monitoring, contraindications), biomedical PM training, troubleshooting.

## 18.7 Warranty

Minimum two (2) years comprehensive warranty including actuator, control system, and upholstery workmanship.

# 19. Rehabilitation Treadmill

Quantity: 1

## 19.1 Clinical & Functional Specifications

Medical-grade rehabilitation treadmill for supervised gait training, cardiovascular conditioning, and functional locomotor rehabilitation of neurological, orthopedic, cardiac, and geriatric patients. Shall provide very-low starting speeds suitable for rehabilitation, fine speed increments, adjustable incline, full-length parallel safety handrails, telemetry heart-rate monitoring, and immediate emergency stop via magnetic safety clip lanyard. Shock-absorbing running deck shall reduce joint impact loads during gait training.

## 19.2 Engineering Technical Specifications

- Speed Range: Minimum 0.1—20 km/h, adjustable in increments of 0.1 km/h.

- Motor: DC or AC servo drive, continuous-duty rating minimum 2.5 HP (approximately 1.85 kW).

- Incline Range: Motorized, minimum 0—15% gradient, adjustable in increments of 0.5% or finer.

- Running Belt Dimensions: Minimum 150 × 50 cm effective running surface.

- Deck: Multi-layer shock-absorbing orthopedic construction with elastomer cushioning.

- Patient Weight Capacity: Minimum 180 kg.

- Handrails: Full-length parallel safety handrails extending the full length of the running belt on both sides, height adjustable or fixed at ergonomic gait-training height.

- Console Display: Backlit color LCD/TFT minimum 7-inch showing speed, incline, time, distance, heart rate, calories, and program status.

- Programs: Minimum 10 pre-programmed rehabilitation and cardiovascular protocols, plus user-programmable modes.

- Heart Rate Monitoring: Integrated telemetry receiver (Polar or equivalent 5 kHz / Bluetooth) plus hand-grip contact sensors.

- Emergency Stop: Magnetic safety clip lanyard cord, plus manual emergency stop button on console.

- Data Export: USB or equivalent for session logging (optional).

- Power Supply: 220—240 VAC, 50/60 Hz, dedicated circuit.

- Power Consumption: Approximately 2000—3000 VA.

- Dimensions (deployed): Approximately 200 × 90 × 140 cm.

- Weight: Approximately 130—200 kg.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-1-6 (usability), EN 957 (stationary training equipment) or equivalent, ISO 13485, CE MDR/FDA.

## 19.3 Standard Accessories

- Full-length parallel safety handrails: Steel construction, ergonomic diameter, non-slip surface finish, extending the full length of the running belt on both sides.

- Low-speed motorized drive (0.1 km/h increment): Servo drive with fine speed resolution suitable for early gait rehabilitation.

- Magnetic emergency safety-stop clip lanyard: Clip-to-patient magnetic key, coiled cord, instantaneous belt stop when detached.

- Integrated telemetry heart-rate receiver: Compatible with standard chest-strap transmitters (5 kHz or Bluetooth Smart).

- Shock-absorbing orthopedic running belt: Multi-layer belt with cushioned deck substrate, low-friction underside, replaceable.

## 19.4 Installation Requirements

Dedicated 220—240 VAC, 50/60 Hz circuit with proper earthing and residual-current protection; floor space minimum 3 m × 2 m plus 1 m clearance behind unit for safety; level floor of adequate load-bearing capacity; ambient 10—35 °C, humidity 30—75% non-condensing; UPS-compatible for control electronics.

## 19.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Report (IEC 62353), Installation & Commissioning Report, PM Schedule.

## 19.6 Training

Clinical application training (gait training protocols, patient safety, harness compatibility), biomedical PM training (belt tension, lubrication, motor service), troubleshooting.

## 19.7 Warranty

Minimum two (2) years comprehensive warranty including motor, drive electronics, and running belt (excluding wear-out consumption).

# 20. Recumbent Bike

Quantity: 1

## 20.1 Clinical & Functional Specifications

Rehabilitation-grade recumbent stationary exercise bike for cardiovascular conditioning, lower-limb strengthening, and functional re-training of patients with cardiac, pulmonary, orthopedic, neurological, and geriatric conditions. The step-through open-frame design shall allow easy patient transfer; the ergonomic seat with lumbar support shall accommodate patients with limited mobility. Magnetic eddy-current resistance shall provide smooth, quiet, and maintenance-free operation.

## 20.2 Engineering Technical Specifications

- Frame: Step-through open frame with low step-over height (approximately 20—30 cm), heavy-duty welded steel.

- Resistance System: Magnetic eddy-current braking, computer-controlled or manually adjustable.

- Resistance Levels: Minimum 16 selectable levels, or continuously adjustable up to minimum 300 W workload.

- Seat: Multi-position sliding adjustable padded bucket seat with contoured lumbar backrest.

- Seat Adjustment: Fore/aft sliding rail with locking mechanism, tool-free adjustment.

- Pedals: Self-leveling wide platform pedals with adjustable hook-and-loop safety foot straps.

- User Weight Capacity: Minimum 150 kg.

- User Height Range: Approximately 150—200 cm.

- Display Console: Backlit LCD showing speed, RPM, time, distance, resistance level, heart rate, calories, workload (Watts).

- Programs: Minimum 10 pre-set + user-programmable.

- Heart Rate Monitoring: Hand-grip contact sensors integrated in side handles + telemetry receiver.

- Power: Self-generating or 100—240 VAC, 50/60 Hz.

- Dimensions: Approximately 150 × 65 × 130 cm.

- Weight: Approximately 50—100 kg.

- Applicable Standards: EN 957 / EN ISO 20957 (stationary training equipment) Class SA (professional/institutional use), IEC 60601-1 (if line-powered), ISO 13485, CE MDR/FDA (where applicable).

## 20.3 Standard Accessories

- Step-through open frame configuration: Low step-over height, heavy-duty steel, powder-coated finish.

- Multi-position adjustable padded bucket seat with lumbar backrest: Fore/aft adjustment, contoured foam, wipe-cleanable upholstery.

- Self-leveling wide foot pedals with safety foot straps: Adjustable hook-and-loop retention, non-slip surface.

- Hand-grip heart rate pulse sensors: Embedded in side handles, contact-type.

- Digital graphic display console: Backlit LCD, electronic resistance controls, standard exercise metrics.

## 20.4 Installation Requirements

Standard 100—240 VAC, 50/60 Hz outlet (if line-powered); floor space minimum 2 × 1.5 m; level, load-bearing floor; ambient 10—35 °C.

## 20.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, PM Schedule.

## 20.6 Training

Clinical application training (cardiac/pulmonary rehab protocols, patient positioning), biomedical PM training.

## 20.7 Warranty

Minimum two (2) years comprehensive warranty.

# 21. Elliptical / Cross Trainer

Quantity: 1

## 21.1 Clinical & Functional Specifications

Low-impact elliptical cross-trainer for cardiovascular conditioning and simultaneous upper and lower body exercise in rehabilitation, cardiac reconditioning, and general fitness training. The articulating pedal motion shall reduce joint impact compared with running, making the device suitable for patients with mild to moderate lower-limb joint conditions. Dual-action handlebars shall provide combined upper- and lower-body workout capability.

## 21.2 Engineering Technical Specifications

- Resistance System: Magnetic eddy-current braking, electronically controlled.

- Resistance Levels: Minimum 16 selectable levels, or continuously adjustable up to minimum 350 W workload.

- Stride Length: Fixed or adjustable within approximately 40—55 cm range.

- Flywheel: Heavy inertial flywheel, approximately 10—20 kg, for smooth motion.

- Foot Pedals: Articulating low-impact platforms with non-slip rubber surfaces.

- Handlebars: Dual-action moving handlebars for upper-body engagement + fixed central handlebars for stability.

- User Weight Capacity: Minimum 150 kg.

- User Height Range: Approximately 150—200 cm.

- Console Display: Backlit LCD/TFT showing speed, RPM, time, distance, resistance, heart rate, calories, workload.

- Programs: Minimum 10 pre-set + user-programmable.

- Heart Rate Monitoring: Hand-grip contact sensors + telemetry receiver.

- Rear Stabilization: Leveling feet for uneven floor.

- Power: Self-generating or 100—240 VAC, 50/60 Hz.

- Dimensions: Approximately 180 × 70 × 170 cm.

- Weight: Approximately 80—130 kg.

- Applicable Standards: EN 957 / EN ISO 20957 Class SA, IEC 60601-1 (if line-powered), ISO 13485, CE MDR/FDA (where applicable).

## 21.3 Standard Accessories

- Low-impact articulating foot pedal platforms: Non-slip rubber linings, adjustable or fixed stride.

- Dual-action upper body exercise handlebars and fixed central handlebars: Ergonomic foam grips, dual functionality.

- Integrated telemetry heart-rate console display: Compatible with standard chest-strap transmitters.

- Magnetic resistance eddy-current braking system: Maintenance-free, silent operation.

- Rear stabilization leveling feet: Adjustable for uneven flooring.

## 21.4 Installation Requirements

Standard 100—240 VAC, 50/60 Hz outlet (if line-powered); floor space minimum 2.5 × 1.5 m; level, load-bearing floor; ambient 10—35 °C.

## 21.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, PM Schedule.

## 21.6 Training

Clinical application training, biomedical PM training.

## 21.7 Warranty

Minimum two (2) years comprehensive warranty.

# 22. Arm Ergometer (Upper Body Cycle)

Quantity: 1

## 22.1 Clinical & Functional Specifications

Upper-extremity ergometer for cardiovascular conditioning, upper-limb strengthening, endurance training, and rehabilitation of patients with lower-limb impairment, paraplegia, cardiac conditions, or post-orthopedic upper-limb rehabilitation. Shall permit bidirectional (forward and reverse) rotation with adjustable resistance, and shall be mountable on a tabletop, wall-mount, or floor stand configuration.

## 22.2 Engineering Technical Specifications

- Resistance System: Mechanical friction brake or magnetic resistance, continuously adjustable.

- Workload Range: Approximately 0—200 W adjustable.

- Rotation: Bidirectional (forward and reverse).

- Crank Length: Adjustable or fixed, typically 15—20 cm.

- Mounting: Tabletop clamping base with wall-mount adaptability, or provided stand.

- Display Console: Digital display of RPM, time, distance, calories, workload.

- Heart Rate: Optional telemetry receiver compatibility.

- Power: Battery-operated console or self-generating; no mains required for basic operation.

- User Weight Capacity: Minimum 130 kg.

- Dimensions: Compact tabletop configuration; footprint approximately 40 × 40 cm.

- Weight: Approximately 8—20 kg.

- Applicable Standards: EN ISO 20957 (where applicable), ISO 13485, CE MDR/FDA (where applicable).

## 22.3 Standard Accessories

- Tabletop/wall-mount adjustable frame base structure with clamping knobs: Heavy-duty steel, corrosion-resistant, non-marring pads for surface protection.

- Bidirectional rotating crank arm assembly with comfortable hand grips: Ergonomic foam or rubber grips, ball-bearing rotation, smooth operation.

- Digital performance monitoring console: RPM, time, workload, calories, battery-operated.

- Adjustable resistance control turn-knob: Mechanical brake with fine adjustment.

- Non-slip rubber counter pads: Protect underlying surface, prevent movement during use.

## 22.4 Installation Requirements

No electrical connection required for basic mechanical unit; ambient 10—40 °C; requires stable tabletop, sturdy chair, or wheelchair-accessible station.

## 22.5 Documentation

User Manual, Parts List, Factory Test Certificate, Warranty Certificate.

## 22.6 Training

Clinical application orientation.

## 22.7 Warranty

Minimum two (2) years comprehensive warranty.

# 23. Stepper for Lower-Limb Re-education

Quantity: 1

## 23.1 Clinical & Functional Specifications

Hydraulic dual-cylinder stepper for lower-limb strengthening, gait pattern re-training, cardiovascular conditioning, and functional weight-bearing exercise. Shall provide independent left/right step motion with adjustable hydraulic resistance, height-adjustable hand-support rails on both sides for patient stability, and digital step-count display.

## 23.2 Engineering Technical Specifications

- Step Mechanism: Independent left/right hydraulic cylinders, alternating motion.

- Resistance: Manually adjustable hydraulic dial, continuously variable.

- Step Height (travel): Approximately 15—25 cm.

- Hand Support Rails: Height-adjustable towers on both sides, ergonomic grip diameter, non-slip surface.

- Step Surfaces: Non-slip corrugated rubber, wide platform (minimum 25 × 15 cm each).

- Display: Digital step counter with time, calories, and auto-reset.

- Power: Battery-operated display (long-life button cell).

- User Weight Capacity: Minimum 130 kg.

- Frame: Heavy-duty welded steel, powder-coated.

- Dimensions: Approximately 60 × 60 × 140 cm.

- Weight: Approximately 20—35 kg.

- Applicable Standards: EN ISO 20957 (stationary training equipment), ISO 13485.

## 23.3 Standard Accessories

- Independent left/right hydraulic step pedal mechanisms: Sealed cylinders, smooth operation, low-maintenance.

- Height-adjustable hand support rail towers: Both sides, tool-free height adjustment, ergonomic grips.

- Manual hydraulic resistance regulation dials: Continuously variable, calibrated indicator.

- Digital step counter display with auto-reset: Battery-operated LCD.

- Non-slip corrugated rubber step surfaces: High-friction, replaceable.

## 23.4 Installation Requirements

No electrical connection required; floor space approximately 1 × 1 m plus lateral clearance; level, non-slip floor; ambient 10—35 °C.

## 23.5 Documentation

User Manual, Parts List, Factory Test Certificate, PM Schedule.

## 23.6 Training

Clinical application orientation.

## 23.7 Warranty

Minimum two (2) years comprehensive warranty including hydraulic cylinders.

# 24. Multi-Station Gym System

Quantity: 1

## 24.1 Clinical & Functional Specifications

Institutional-grade multi-station selectorized gym system for progressive resistance training of major muscle groups (chest, back, shoulders, arms, legs, and core) during musculoskeletal rehabilitation and general strengthening programs. Shall provide multiple exercise stations sharing selectorized weight stacks with magnetic or pin-locking weight selection, protective weight-stack shrouds, and interchangeable exercise attachments.

## 24.2 Engineering Technical Specifications

- Number of Stations: Minimum 3 exercise stations (typical configuration includes lat pull-down, low row, chest press, leg extension, or equivalent).

- Weight Stacks: Selectorized cast-iron or steel plate stacks with magnetic or pin locking selector; minimum stack weight 80 kg per station, in 5 kg increments (or 10-lb increments).

- Frame: Heavy-duty welded steel cage structure, minimum 60 × 60 mm tubing or equivalent, powder-coated corrosion-resistant finish.

- Cables: High-tensile aircraft-grade steel cables, nylon-coated, minimum 5 mm diameter, rated breaking load minimum 10× working load.

- Pulleys: Fibreglass-reinforced nylon or equivalent, sealed ball bearings, self-aligning.

- Weight-Stack Shrouds: Full-height protective enclosures preventing access to moving weights.

- Exercise Bench: Upholstered multi-position adjustable bench, high-density foam padding, wipe-cleanable vinyl.

- User Weight Capacity: Minimum 150 kg.

- Overall Dimensions: Approximately 300 × 200 × 220 cm (footprint depends on configuration).

- Weight (equipment): Approximately 300—600 kg total.

- Applicable Standards: EN ISO 20957 Class SA (professional/institutional use), ISO 13485 where applicable.

## 24.3 Standard Accessories

- Heavy-duty steel frame cage structure with protective weight stack shrouds: Welded steel, powder-coated, full-enclosure shrouds.

- Selectorized weight stack blocks with magnetic or pin selector locking: Cast-iron or steel plates, minimum 80 kg per stack, precise selector engagement.

- Interchangeable pull-down lat bar, rowing handle, and ankle cuffs: Chrome or rubber-coated, ergonomic grips, quick-connect carabiners.

- High-tensile steel aircraft cables with nylon-coated pulleys: Rated 10× working load minimum, sealed bearing pulleys.

- Upholstered multi-position adjustable exercise bench: Foam-padded, wipe-cleanable vinyl, multi-angle adjustment, integrated.

## 24.4 Installation Requirements

No electrical connection required; floor space minimum 4 × 3 m plus 1 m clearance around perimeter; level reinforced floor with load-bearing capacity minimum 800 kg; ceiling height minimum 2.5 m; ambient 10—35 °C. Professional installation and anchoring recommended per manufacturer guidance.

## 24.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Installation Report, PM Schedule.

## 24.6 Training

Clinical application training (exercise prescription, patient safety, spotting), biomedical/facility PM training (cable inspection, lubrication, upholstery care).

## 24.7 Warranty

Minimum two (2) years comprehensive warranty on frame, weight stacks, cables, and pulleys; upholstery minimum 1 year.

GAIT & STANDING REHABILITATION EQUIPMENT

# 25. Parallel Bars (Adult)

Quantity: 1

## 25.1 Clinical & Functional Specifications

Floor-mounted adult parallel bars for supervised gait training, weight-bearing progression, balance re-education, and functional ambulation training of patients recovering from stroke, spinal cord injury, lower-limb orthopedic surgery, amputation with prosthesis, and neurological gait disorders. The bars shall provide bilateral upper-extremity support during walking practice, with height and width adjustment to accommodate patients of varying stature and clinical requirement. Tapered entry ramps at both ends shall permit wheelchair-accessible entry and exit.

## 25.2 Engineering Technical Specifications

- Overall Length: Minimum 3.0 m (working length between end supports).

- Handrail Height Range: Adjustable, approximately 70—110 cm from platform surface.

- Handrail Width (between bars): Adjustable, approximately 45—85 cm.

- Height/Width Adjustment: Tool-free manual turn-knob or hand-lever mechanism with positive locking.

- Handrail Material: Stainless steel or chrome-plated steel tubing, external diameter approximately 35—40 mm, smooth polished finish.

- Handrail Load Capacity: Minimum 150 kg static load per bar at mid-span.

- Base Platform: Continuous non-slip wooden or steel platform between uprights; alternately, floor-mount uprights only with anti-slip floor covering.

- Platform Surface: Anti-slip texturized surface, wipe-cleanable.

- Entry Ramps: Tapered wooden or metal run-up ramps at both ends, minimum 30 cm length, non-slip surface, integrated or removable.

- Frame: Heavy-duty welded steel uprights, powder-coated corrosion-resistant finish.

- User Weight Capacity: Minimum 150 kg.

- Applicable Standards: EN ISO 20957 (institutional training equipment), ISO 13485 (where applicable), general medical device requirements.

## 25.3 Standard Accessories

- Heavy-duty floor-mount non-slip wooden or steel base platform: Continuous platform between uprights, anti-slip texturized surface, wipe-cleanable, reinforced load-bearing construction.

- Height-adjustable stainless steel handrail tubes with manual turn-knobs: Corrosion-resistant tubing, tool-free positive-locking height adjustment, smooth polished finish.

- Width-adjustable handrail mechanism with sliding brackets: Tool-free width adjustment, positive locking, symmetric adjustment.

- Tapered run-up entry ramp thresholds at both ends: Non-slip surface, wheelchair-accessible slope, integrated or removable.

## 25.4 Installation Requirements

Floor space minimum 4 × 2 m; level reinforced floor with anchor points for uprights (concrete or reinforced sub-floor); ceiling clearance minimum 2.5 m; ambient 10—35 °C. Professional installation and anchoring per manufacturer guidance. No electrical connection required.

## 25.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Load-Test Certificate, Installation Report, PM Schedule.

## 25.6 Training

Clinical application training (gait training progressions, patient safety, guarding techniques), facility PM training (bolt torque check, handrail alignment).

## 25.7 Warranty

Minimum two (2) years comprehensive warranty on frame and hardware.

# 26. Paediatric Parallel Bars

Quantity: 1

## 26.1 Clinical & Functional Specifications

Floor-mounted paediatric parallel bars for gait training, weight-bearing progression, balance re-education, and functional ambulation training of children with cerebral palsy, developmental delay, post-orthopedic surgery, spina bifida, and other paediatric neurological or musculoskeletal conditions. The bars shall be sized for the paediatric population with narrow-diameter handrails, low-profile floor base, and micro-adjustable height and width mechanisms. Brightly colored finishes and removable path divider obstacles shall enhance engagement and support therapeutic obstacle-navigation training.

## 26.2 Engineering Technical Specifications

- Overall Length: Minimum 2.0 m (working length between end supports).

- Handrail Height Range: Adjustable, approximately 40—75 cm from platform surface (paediatric range).

- Handrail Width (between bars): Adjustable, approximately 25—55 cm.

- Height/Width Adjustment: Micro-adjustable tool-free turn-knob or hand-lever mechanism with fine increments and positive locking.

- Handrail Material: Stainless steel or chrome-plated steel tubing, narrow paediatric diameter approximately 25—30 mm.

- Handrail Load Capacity: Minimum 75 kg static load per bar at mid-span.

- Base Platform: Low-profile continuous non-slip platform between uprights; brightly colored child-friendly finish.

- Platform Surface: Anti-slip texturized surface, wipe-cleanable, non-toxic finish per EN 71-3 (safety of toys — migration of certain elements) or equivalent paediatric safety standard.

- Frame: Heavy-duty welded steel uprights with brightly colored powder-coat finish, rounded edges and corners, no pinch points accessible to children.

- User Weight Capacity: Minimum 75 kg.

- Applicable Standards: EN 71 series (paediatric safety), EN ISO 20957, ISO 13485 where applicable.

## 26.3 Standard Accessories

- Child-sized low-profile floor base structure with brightly colored finishes: Non-toxic paint per paediatric safety standards, rounded edges, wipe-cleanable.

- Micro-adjustable height and width handrail locks: Fine-increment adjustment optimized for paediatric anthropometry, tool-free, positive locking.

- Ergonomic narrow-diameter handrail tubes: Sized for small hand grips (25—30 mm diameter), smooth polished stainless steel or chrome-plated finish.

- Removable colored walking path divider obstacles set: Non-toxic materials, rounded edges, various shapes and heights for obstacle-navigation training.

## 26.4 Installation Requirements

Floor space minimum 3 × 2 m; level reinforced floor with anchor points; ceiling clearance minimum 2.5 m; ambient 10—35 °C; paediatric-safe environment (no sharp edges, secure anchoring). Professional installation per manufacturer guidance. No electrical connection required.

## 26.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Load-Test Certificate, Paediatric Safety Certificate (materials/finish), Installation Report, PM Schedule.

## 26.6 Training

Clinical application training (paediatric gait training, engagement strategies), facility PM training.

## 26.7 Warranty

Minimum two (2) years comprehensive warranty.

# 27. Exercise Staircase

Quantity: 1

## 27.1 Clinical & Functional Specifications

Corner-configured therapeutic exercise staircase for stair-climbing rehabilitation, functional gait training, quadriceps and hip strengthening, and confidence-building in patients recovering from lower-limb injury, joint replacement, stroke, or neurological gait impairment. The unit shall provide both short-step (low-rise) and tall-step (standard-rise) sections in a corner configuration, with continuous double-tier handrails at both adult and paediatric heights, and anti-slip tread surfaces on all steps.

## 27.2 Engineering Technical Specifications

- Configuration: Corner (L-shaped) design with two adjoining flights: short-step section and tall-step section.

- Short-Step Section: Minimum 3 steps, riser height approximately 8—12 cm.

- Tall-Step Section: Minimum 3 steps, riser height approximately 15—18 cm (standard building-code rise).

- Tread Depth: Minimum 28 cm per step.

- Tread Width: Minimum 60 cm.

- Top Platform: Minimum 80 × 80 cm intermediate landing at corner.

- Handrails: Continuous double-tier wooden or chrome-plated steel handrails on both sides, at adult height (approximately 90 cm) and paediatric height (approximately 60 cm).

- Tread Surface: Anti-slip high-friction safety mats bonded to all treads.

- Frame: Reinforced wooden or steel load-bearing structure with structural support brackets.

- User Weight Capacity: Minimum 150 kg.

- Overall Footprint: Approximately 2 × 2 m (L-shape).

- Applicable Standards: ISO 13485 where applicable, general institutional safety requirements, applicable national building safety codes.

## 27.3 Standard Accessories

- Corner-configured wooden structure with short and tall step sections: L-shaped configuration, integrated corner landing, reinforced wooden construction with clear protective finish.

- Continuous double-tier wood handrail tracking (adult and paediatric heights): Both sides, smooth polished finish, ergonomic diameter, positive attachment.

- Anti-slip high-friction tread safety mats: Bonded to all step surfaces, high-friction texturized, replaceable.

- Reinforced load-bearing structural support brackets: Steel or heavy-duty timber, load-tested to minimum 3× working load.

## 27.4 Installation Requirements

Floor space minimum 3 × 3 m plus 1 m clearance on approach sides; level, load-bearing floor; ceiling clearance minimum 2.8 m; ambient 10—35 °C. Professional installation and anchoring recommended.

## 27.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Load-Test Certificate, Installation Report, PM Schedule.

## 27.6 Training

Clinical application training (stair-training progression, safety guarding), facility PM training.

## 27.7 Warranty

Minimum two (2) years comprehensive warranty on frame and hardware.

# 28. Para-Podium (Standing Frame)

Quantity: 1

## 28.1 Clinical & Functional Specifications

Dynamic upright standing frame (para-podium) for supported passive and active standing of patients with paraplegia, spinal cord injury, severe neurological impairment, muscular dystrophy, and prolonged wheelchair dependence. Shall provide safe hands-free upright positioning with anterior chest support, posterior lumbar/pelvic support, knee stabilizers, and wide anti-tip foot platforms. Enables prevention of contractures, osteoporosis, urinary and pulmonary complications, and improvement of cardiovascular tolerance to the upright position.

## 28.2 Engineering Technical Specifications

- Frame: Dynamic upright metal cage structure, heavy-duty welded steel with corrosion-resistant powder-coat finish.

- Standing Configuration: Full weight-bearing upright standing with hands-free support.

- Anterior Chest Support: Padded upholstered chest rest panel, adjustable height and depth.

- Posterior Support: Upholstered pelvic/lumbar back brace plates, adjustable height.

- Knee Stabilizers: Padded adjustable knee blocks with safety lock brackets, bilateral, adjustable width and height.

- Foot Platform: Wide counter-balanced anti-tip platforms with adjustable shoe straps, minimum 40 × 30 cm each.

- Adjustability: Tool-free adjustment of chest, pelvis, knee, and foot components for various patient heights (approximately 150—200 cm).

- User Weight Capacity: Minimum 130 kg.

- Base: Wide anti-tip footprint with lockable casters (optional) or fixed base per tender preference.

- Upholstery: Wipe-clean fluid-proof antimicrobial vinyl over high-density foam padding.

- Applicable Standards: ISO 13485, EN 12182 (assistive products for persons with disability — general requirements and test methods), CE MDR/FDA where applicable.

## 28.3 Standard Accessories

- Dynamic upright standing metal frame cage structure: Heavy-duty steel, powder-coated, corrosion-resistant, anti-tip wide-base footprint.

- Padded adjustable knee stabilizer supports with safety lock brackets: Bilateral, tool-free width and height adjustment, positive locking, wipe-cleanable padding.

- Upholstered pelvic/lumbar structural back brace plates: Height-adjustable, contoured foam padding, wipe-cleanable vinyl.

- Wide counter-balanced anti-tip foot platforms with shoe straps: Adjustable hook-and-loop shoe retention, non-slip surface, wide stable footprint.

- Padded anterior chest rest support panel: Height and depth adjustable, contoured foam padding, wipe-cleanable vinyl.

## 28.4 Installation Requirements

Floor space minimum 1.5 × 1.5 m; level, load-bearing floor; ambient 10—35 °C. No electrical connection required. Assistance of two staff members typically required for patient transfer.

## 28.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Load-Test Certificate, PM Schedule.

## 28.6 Training

Clinical application training (transfer techniques, positioning, contraindications, monitoring for orthostatic intolerance), biomedical/facility PM training.

## 28.7 Warranty

Minimum two (2) years comprehensive warranty on frame and upholstery workmanship.

# 29. Universal Exercise Unit with Suit (Rocher Cage / Spider Cage System)

Quantity: 1

## 29.1 Clinical & Functional Specifications

Multi-purpose universal exercise unit consisting of a heavy-duty steel wire-mesh suspension cage (Rocher-style / spider cage) with associated dynamic orthotic compression suit, bungee/pulley suspension system, slings, and stabilization anchors. The system shall support intensive neurodevelopmental rehabilitation of paediatric and adult patients with cerebral palsy, stroke, traumatic brain injury, spinal cord injury, and neuromotor disorders through suspension therapy, dynamic proprioceptive resistance training, and assisted or resisted functional movement patterns. The compression suit shall provide dynamic proprioceptive input and postural support during exercise.

## 29.2 Engineering Technical Specifications

- Cage Structure: Heavy-duty welded steel wire-mesh cage, minimum 2 × 2 × 2 m internal working volume, powder-coated corrosion-resistant finish.

- Wire Mesh: Structural steel grid, spacing approximately 10 × 10 cm, load-rated for multiple simultaneous suspension attachment points.

- Attachment Points: Multiple S-hook or carabiner anchor points on ceiling, walls, and floor of cage, minimum 100 usable attachment locations.

- Load Capacity per Attachment Point: Minimum 100 kg working load.

- Compression Suit: Dynamic orthotic garment with adjustable elastic tensioning bands, provided in Small, Medium, and Large sizes (or equivalent size range), latex-free, wipe-cleanable/washable.

- Suspension Components: Elastic bungee tension cords of graded resistances, weighted pulley suspension lines, padded slings, canvas straps, adjustable webbing straps, metal S-hooks and carabiners.

- Foot-Anchoring System: Fitted stabilizer shoe tracks/anchors within cage base for lower-limb fixation.

- Frame Material: Heavy-duty steel, minimum 40 × 40 mm tubing or equivalent structural section.

- User Weight Capacity: Minimum 100 kg.

- Applicable Standards: EN 12182, ISO 13485, general institutional load-bearing safety requirements.

## 29.3 Standard Accessories

- Heavy-duty steel wire-mesh suspension cage (Rocher cage style): Welded steel construction, multi-point attachment grid, powder-coated finish.

- Dynamic orthotic compression suit garments (Small, Medium, Large): Latex-free elasticated material, adjustable tension, washable, sized for paediatric through adult use.

- Elastic bungee tension cords and weighted pulley suspension lines: Multiple graded resistances, safety-rated attachment terminations.

- Padded suspension slings, canvas straps, and metal S-hooks set: Load-rated hardware, comfort-padded slings, adjustable webbing.

- Fitted stabilizer foot-anchoring shoes / tracks inside cage base: Adjustable, secure lower-limb fixation.

## 29.4 Installation Requirements

Dedicated therapy room; floor space minimum 3 × 3 m; ceiling height minimum 2.5 m; reinforced level floor with load-bearing capacity minimum 500 kg; wall or floor anchoring per manufacturer specification; ambient 10—35 °C. Professional installation, structural anchoring, and post-installation load testing required.

## 29.5 Documentation

User Manual, Assembly Manual, Parts Catalogue, Factory Test Certificate, Load-Test Certificate (per attachment point), Installation & Commissioning Report, PM Schedule.

## 29.6 Training

Clinical application training (neurodevelopmental therapy protocols, suspension setup, patient safety), biomedical/facility PM training (hardware inspection, cable/bungee replacement intervals, cage integrity).

## 29.7 Warranty

Minimum two (2) years comprehensive warranty on cage structure and hardware; consumable elastic bungee cords and suit components typically warranted for defects only.

ROBOTICS & BALANCE REHABILITATION EQUIPMENT

# 30. Robotic Hand Glove

Quantity: 1

## 30.1 Clinical & Functional Specifications

The equipment shall be a motorized pneumatic (or equivalent actuation) robotic exoskeleton hand glove system for active-assisted, passive, and mirror-therapy rehabilitation of the hand and fingers in patients with stroke, traumatic brain injury, spinal cord injury, and peripheral nerve injury. The system shall provide programmable finger flexion and extension exercises with adjustable range of motion, speed, and repetitions. Mirror-therapy operation shall permit the affected hand to mirror the movements of the healthy contralateral hand through a data-sensing mirror glove. The device shall support multiple therapy modes including continuous passive motion, active-assisted, task-oriented grasp-release training, and bilateral symmetric training.

## 30.2 Engineering Technical Specifications

- Actuation: Pneumatic soft-robotic actuators or equivalent (cable-driven, motor-driven) exoskeleton mechanism.

- Configurations: Left and Right hand units, adult sizing (S/M/L or equivalent).

- Actuated Digits: Minimum 4 fingers (index through little finger); thumb actuation preferred.

- Range of Motion: Full physiological flexion/extension of MCP, PIP, and DIP joints or equivalent functional range.

- Modes: Passive (CPM), Active-Assisted, Mirror Therapy, Task-Oriented Grasp/Release, Bilateral Training.

- Speed: Adjustable cycle time, approximately 1—10 seconds per flexion/extension cycle.

- Force / Pressure Control: Adjustable actuation pressure with over-pressure safety cut-off.

- Session Duration: Programmable, minimum 1—60 minutes.

- Repetitions: Programmable, minimum 1—500 per session.

- User Interface: Touchscreen console minimum 7-inch, resolution minimum 800 × 480, with graphical protocol selection, session logging, and progress reporting.

- Data Management: Patient database, session archiving, and export via USB or network.

- Safety: Emergency stop button, pressure/force limit, automatic release on power loss.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-1-6 (usability), IEC 62304 (software), IEC 80601-2-78 (robotic rehabilitation equipment — where applicable), ISO 13485, CE MDR/FDA.

## 30.3 Standard Accessories

- Motorized pneumatic exoskeleton glove garments (Left and Right): Soft-robotic construction, latex-free, hypoallergenic contact surfaces, adjustable strap retention, adult sizing.

- Data-sensing mirror glove: Sensor-instrumented glove for healthy hand, bilateral tracking of finger movements, adjustable fit.

- Main pneumatic pump control console tower: Integrated pump/compressor, digital touchscreen minimum 7-inch, low-noise operation.

- Flexible air line connection tubes with quick-click couplers: Kink-resistant medical-grade tubing, quick-disconnect fittings, minimum 2 m length.

- AC power supply adapter with medical-grade cord: IEC 60601-1 compliant, country-specific plug.

## 30.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet with proper earthing; tabletop or mobile-cart placement; ambient 10—35 °C; humidity 30—75% non-condensing; UPS-compatible.

## 30.5 Documentation

User Manual, Service Manual, Parts Catalogue, Software Manual, Factory Test Certificate, Electrical Safety Report (IEC 62353), Software Validation Certificate, PM Schedule.

## 30.6 Training

Clinical application training (protocols, patient selection, contraindications), software training, biomedical PM training.

## 30.7 Warranty

Minimum two (2) years comprehensive warranty including actuators, pump, and control electronics; software updates included.

# 31. Stroke Rehabilitation Glove (Passive)

Quantity: 1

## 31.1 Clinical & Functional Specifications

Passive spring-assisted mechanical rehabilitation glove for finger extension assistance and functional grasp training in patients with stroke, peripheral nerve injury, spasticity, and chronic hand impairment. The device shall passively assist finger extension against flexor spasticity or weakness, allowing active flexion for grasp and spring-assisted release. Multiple sizes and interchangeable tension elements shall accommodate varied patient hand sizes and impairment severities.

## 31.2 Engineering Technical Specifications

- Actuation: Passive spring-assisted finger extension mechanism (dynamic outrigger or equivalent).

- Sizing: Multiple sizes provided (Small, Medium, Large or equivalent).

- Adjustable Tension: Interchangeable spring rods or elastic bands with graded resistance/assistance levels (minimum 3 levels).

- Wrist Support: Integrated adjustable wrist brace/splint base with hook-and-loop straps.

- Materials: Latex-free, hypoallergenic contact surfaces, breathable fabric.

- Finger Contact: Texturized non-slip finger pad grips for functional object handling.

- Cleaning: Wipe-cleanable, spot-washable.

- Applicable Standards: ISO 13485, EN 12182 (assistive products), CE MDR/FDA where applicable, RoHS-compliant materials.

## 31.3 Standard Accessories

- Passive spring-assisted finger extension glove garments (various sizes): S/M/L range, latex-free breathable fabric.

- Interchangeable tension spring rods / elastic resistance bands: Multiple graded levels, tool-free replacement.

- Adjustable wrist brace splint base with hook-and-loop straps: Anatomical fit, adjustable tension.

- Texturized non-slip finger pad grips: Silicone or equivalent non-slip surface for functional grasp.

## 31.4 Installation Requirements

None. Portable device requiring no electrical connection.

## 31.5 Documentation

User Manual, Fitting Guide, Cleaning Instructions, Warranty Certificate.

## 31.6 Training

Clinical application training (fitting, patient selection, home-use instructions).

## 31.7 Warranty

Minimum one (1) year against manufacturing defects (soft goods).

# 32. Balance / Wobble Board

Quantity: 1

## 32.1 Clinical & Functional Specifications

Circular platform balance/wobble board with integrated fulcrum pivot for proprioceptive training, ankle rehabilitation, balance re-education, and core stabilization. Shall support multi-directional tilt training in the management of ankle instability, post-surgical rehabilitation, fall-prevention programs, sports performance training, and neurological balance retraining.

## 32.2 Engineering Technical Specifications

- Platform Diameter: Approximately 40 cm (standard adult size).

- Platform Material: Heavy-duty wood (birch/plywood) or rigid impact-resistant plastic.

- Fulcrum: Integrated bottom hemisphere pivot dome, height approximately 6—8 cm.

- Tilt Angle: Multi-directional (360°) tilt, maximum tilt angle approximately 15—20° from horizontal.

- Surface Coating: High-friction anti-slip emery texturized top surface.

- Perimeter: Protective rubber edge molding strip.

- User Weight Capacity: Minimum 130 kg.

- Weight: Approximately 1.5—3 kg.

- Applicable Standards: ISO 13485 where applicable; general institutional safety.

## 32.3 Standard Accessories

- Heavy-duty wooden or rigid plastic circular platform disk: Non-toxic finish, warp-resistant, load-tested.

- Integrated bottom fulcrum hemisphere dome pivot: Fixed to platform, non-slip base.

- High-friction anti-slip emery texturized top surface coating: Bonded to top face, replaceable if required.

- Perimeter protective edge rubber molding strip: Impact-absorbing, prevents floor damage.

## 32.4 Installation Requirements

None. Requires level non-slip floor and open training space minimum 2 × 2 m with fall clearance.

## 32.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 32.6 Training

Clinical use orientation.

## 32.7 Warranty

Minimum one (1) year against manufacturing defects.

# 33. Bosu Ball (Balance Trainer)

Quantity: 1

## 33.1 Clinical & Functional Specifications

Hemispheric inflatable balance trainer with rigid flat base platform for balance training, proprioception, core stabilization, and functional dynamic exercise. The dome-up configuration provides an unstable dynamic surface for standing balance training; the dome-down (flat side up) configuration provides an unstable surface for upper-extremity and core exercises.

## 33.2 Engineering Technical Specifications

- Dome: Heavy-duty burst-resistant rubber hemisphere, diameter approximately 60—65 cm.

- Base Platform: Rigid flat fiberglass or reinforced plastic circular ring with non-skid rubber pads.

- Height (inflated): Approximately 25—30 cm.

- Burst Resistance: Anti-burst construction rated for user weight cycling.

- User Weight Capacity: Minimum 130 kg.

- Inflation: Adjustable via dual-action foot pump; recommended firmness verified via measuring tape.

- Cleaning: Wipe-clean with mild soap or hospital disinfectant.

- Applicable Standards: EN ISO 20957 where applicable, general institutional safety.

## 33.3 Standard Accessories

- Inflatable heavy-duty burst-resistant rubber dome hemisphere: Anti-burst construction, textured non-slip surface.

- Rigid flat fiberglass base platform ring with non-skid rubber pads: Corrosion-resistant, load-tested.

- Manual dual-action foot inflation pump assembly: Bi-directional, integrated pressure gauge or measuring tape.

- Pressure verification measuring tape gauge and valve plug replacement sets: For consistent inflation dome height.

## 33.4 Installation Requirements

None. Requires level non-slip floor and open training space minimum 2 × 2 m with fall clearance.

## 33.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 33.6 Training

Clinical use orientation.

## 33.7 Warranty

Minimum one (1) year against manufacturing defects.

# 34. Mini-Trampoline (Rebounder)

Quantity: 1

## 34.1 Clinical & Functional Specifications

Circular mini-trampoline (rebounder) for low-impact cardiovascular conditioning, balance training, proprioception, and functional rebound therapy. Shall provide a stable, spring-suspended jumping surface with detachable stability handlebar for supported rebounding of geriatric, neurological, or deconditioned patients. Intended for supervised institutional use.

## 34.2 Engineering Technical Specifications

- Frame: Heavy-duty tubular steel circular frame, powder-coated corrosion-resistant finish.

- Diameter: Approximately 100—110 cm.

- Height (jumping surface): Approximately 22—28 cm from floor.

- Suspension: High-tensile industrial steel rebound spring coils (minimum 30 springs) or equivalent bungee-band suspension; protective pad covering springs at all times.

- Jumping Mat: Heavy-duty woven nylon or polypropylene mat, UV- and tear-resistant.

- Handlebar: Detachable stability front handlebar, height-adjustable, foam grip padding.

- User Weight Capacity: Minimum 130 kg.

- Legs: Non-slip heavy-duty rubber leg tips (minimum 6 legs).

- Applicable Standards: EN ISO 20957 or equivalent institutional training equipment standard.

## 34.3 Standard Accessories

- Heavy-duty tubular steel circular frame structure: Welded, powder-coated corrosion-resistant.

- High-tensile industrial rebound spring coils with protective pad cover: Minimum 30 springs, spring pad covers entire spring assembly.

- Heavy-duty woven nylon jumping mat sheet: UV- and tear-resistant, replaceable.

- Detachable stability front handlebar attachment with foam grip: Tool-free assembly, height-adjustable, ergonomic grip.

- Non-slip heavy-duty rubber leg tips: Floor-protecting, non-marring.

## 34.4 Installation Requirements

None fixed. Requires level non-slip floor, ceiling clearance minimum 2.5 m, and open training space minimum 3 × 3 m with fall clearance.

## 34.5 Documentation

Assembly Manual, User Manual, Exercise Guide, Warranty Certificate.

## 34.6 Training

Clinical use orientation.

## 34.7 Warranty

Minimum two (2) years on frame; springs and mat minimum one (1) year.

# 35. Platform Swing (Sensory Integration)

Quantity: 1

## 35.1 Clinical & Functional Specifications

Ceiling-suspended rigid rectangular platform swing for sensory integration therapy, vestibular stimulation, proprioceptive training, and paediatric neurodevelopmental therapy. Shall support seated, prone, supine, or standing configurations for children and adults with sensory processing disorders, autism spectrum disorder, cerebral palsy, and other neurodevelopmental conditions.

## 35.2 Engineering Technical Specifications

- Platform: Rigid wooden rectangular base, dimensions approximately 90 × 60 cm, carpeted top surface, padded perimeter bumper.

- Load Capacity: Minimum 100 kg working load.

- Suspension: Heavy-duty dynamic ropes or equivalent, variable height adjustment, four-point corner suspension.

- Ceiling Attachment: Industrial-strength swivel hook brackets with safety carabiners, ceiling load-rated per suspension load × safety factor of minimum 5.

- Swivel: Rotational single-point swivel joint adapter for 360° rotation.

- Height: Adjustable via rope length; typical operating height 15—50 cm above floor.

- Applicable Standards: EN 12182, ISO 13485 where applicable, structural load-testing requirements.

## 35.3 Standard Accessories

- Carpeted rigid wooden rectangular platform base with padded perimeter bumper: Reinforced construction, non-toxic finish, wipe-cleanable padding.

- Heavy-duty dynamic suspension ropes with variable height adjustment links: Load-rated ropes, adjustable coupling links.

- Industrial-strength ceiling-mount swivel hook brackets with safety carabiners: Steel construction, load-tested to minimum 5× working load.

- Rotational single-point swivel joint adapter: Full 360° rotation, sealed bearings.

## 35.4 Installation Requirements

Ceiling anchor point with structural load-bearing capacity minimum 5× working load, professionally installed and certified; floor space minimum 3 × 3 m clear on all sides; ceiling height minimum 2.5 m; padded floor mats recommended below swing; ambient 10—35 °C.

## 35.5 Documentation

Installation Manual, Load-Test Certificate, User Manual, PM Schedule, Warranty Certificate.

## 35.6 Training

Clinical application training (sensory integration protocols, safety), installation and inspection training.

## 35.7 Warranty

Minimum two (2) years on hardware; soft goods (ropes, padding) minimum one (1) year.

# 36. Disc Swing & Sensory-Integration Swings

Quantity: 1

## 36.1 Clinical & Functional Specifications

Set of ceiling-suspended sensory integration swings including disc swing, Lycra compression hammock, and bolster roll swing, for vestibular stimulation, proprioceptive input, deep-pressure sensory input, and paediatric neurodevelopmental therapy. Shall be interchangeable via standard carabiner attachments.

## 36.2 Engineering Technical Specifications

- Disc Swing: Molded rubber circular seat, approximately 40 cm diameter, central vertical rope attachment.

- Lycra Hammock: Compression fabric hammock, dimensions approximately 200 × 150 cm, elasticated for deep-pressure input, dual suspension points.

- Bolster Swing: Cylindrical padded bolster, approximately 20 cm diameter × 90 cm length, dual suspension lines.

- Suspension: Heavy-duty soft-grip vertical rope (disc) or dual suspension lines (hammock/bolster); load-rated to minimum 100 kg.

- Hardware: Locking steel carabiners and swivel connectors, load-tested minimum 5× working load.

- Ceiling Attachment: Compatible with standard sensory-integration ceiling anchor points.

- Applicable Standards: EN 12182, EN 71 (paediatric safety), ISO 13485 where applicable.

## 36.3 Standard Accessories

- Molded rubber circular disc swing seat: Non-toxic material, textured non-slip surface, rounded edges.

- Heavy-duty soft-grip central vertical suspension rope: Load-rated, ergonomic grip.

- Lycra compression hammock swing garment: Latex-free elasticated fabric, washable.

- Bolster swing roll attachment with dual suspension lines: Padded cylindrical bolster, wipe-cleanable, dual suspension.

- Heavy-duty locking steel carabiners and safety swivel connectors: Load-rated, corrosion-resistant.

## 36.4 Installation Requirements

Ceiling anchor points structurally rated per swing load × safety factor of minimum 5; floor space minimum 3 × 3 m clear; ceiling height minimum 2.5 m; padded floor mats recommended; ambient 10—35 °C.

## 36.5 Documentation

Installation Manual, Load-Test Certificate, User Manual, PM Schedule, Warranty Certificate.

## 36.6 Training

Clinical application training (sensory integration protocols, safety, swing selection).

## 36.7 Warranty

Minimum two (2) years on hardware; soft goods minimum one (1) year.

### Wheel Board (Scooter Board) — Quantity: 1

#### 1. Clinical & Functional Specifications

Low-profile rectangular scooter board with four swiveling caster wheels for prone-position paediatric therapy, upper-body strengthening, sensory-motor integration, and functional play-based rehabilitation. Shall support prone/supine positioning of paediatric patients for pushing/pulling activities in occupational and physical therapy.

#### 2. Engineering Technical Specifications

- Platform: Rectangular padded wooden or high-density plastic board, dimensions approximately 40 × 30 cm.

- Padding: Foam padding with wipe-clean vinyl cover.

- Wheels: Four 360°-swiveling ball-bearing caster wheels, non-marring polymer treads, wheel diameter approximately 5 cm.

- Handles: Dual side-mounted grab-hand slots or cutouts.

- Perimeter: Foam bumper strip on all edges for wall impact protection.

- User Weight Capacity: Minimum 75 kg (paediatric/adolescent use).

- Applicable Standards: EN 71 (paediatric safety), ISO 13485 where applicable, non-toxic finish.

#### 3. Standard Accessories

- Low-profile rectangular padded wooden rolling platform (scooter board): Reinforced construction, non-toxic finish, wipe-cleanable padding.

- Four 360°-swiveling ball-bearing caster wheels: Non-marring treads, sealed bearings, smooth operation.

- Dual safety side-mounted handle grab slots: Ergonomic cut-outs or attached handles.

- Perimeter foam bumper strip: Impact-absorbing, prevents wall/floor damage.

#### 4. Installation Requirements

None. Requires level smooth floor and open training space minimum 4 × 4 m with fall clearance.

#### 5. Documentation

Product Information, Exercise Guide, Warranty Certificate.

#### 6. Training

Clinical use orientation.

#### 7. Warranty

Minimum one (1) year against manufacturing defects.

EXERCISE BALLS, BANDS & WEIGHTS

# 37. Peanut Ball

Quantity: 1

## 37.1 Clinical & Functional Specifications

Elongated peanut-shaped inflatable therapy ball for paediatric neurodevelopmental therapy, core stabilization, trunk control training, postural exercises, and adult musculoskeletal rehabilitation. The peanut shape shall provide dual-directional stability (rolls in one axis, stable in the perpendicular axis), permitting controlled therapeutic exercise for patients requiring less challenging balance surfaces than round balls.

## 37.2 Engineering Technical Specifications

- Shape: Elongated peanut/dumbbell configuration.

- Standard Size: Approximately 40 × 80 cm (diameter × length), or equivalent size range.

- Material: Heavy-duty burst-resistant PVC or equivalent, phthalate-free.

- Burst Resistance: Anti-burst construction per EN 71 or equivalent standard.

- User Weight Capacity: Minimum 130 kg (adult) or paediatric-rated per configuration.

- Surface: Non-slip texturized finish.

- Cleaning: Wipe-clean with mild soap or hospital disinfectant.

- Applicable Standards: EN 71 (paediatric safety), ISO 13485 where applicable, phthalate-free/RoHS-compliant materials.

## 37.3 Standard Accessories

- Inflatable heavy-duty burst-resistant PVC elongated peanut shell ball: Anti-burst construction, phthalate-free, non-slip surface.

- Manual dual-action hand inflation pump cylinder: Bi-directional pumping, compatible valve fitting.

- Tapered replacement air valve pins and pin-extraction fork tool: For deflation, storage, and pressure adjustment.

## 37.4 Installation Requirements

None. Storage rack recommended.

## 37.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 37.6 Training

Clinical use orientation.

## 37.7 Warranty

Minimum one (1) year against manufacturing defects.

# 38. Gym Ball Set

Quantity: 1

## 38.1 Clinical & Functional Specifications

Set of anti-burst inflatable therapy balls in multiple sizes for core stabilization, balance training, spinal rehabilitation, postural exercises, paediatric and adult neurodevelopmental therapy, and general therapeutic exercise. Sizes shall accommodate patients of varying stature and clinical requirement.

## 38.2 Engineering Technical Specifications

- Sizes Included: Minimum three sizes: 55 cm, 65 cm, 75 cm diameter (or equivalent metric range).

- Material: Heavy-duty burst-proof PVC or equivalent, phthalate-free.

- Burst Resistance: Slow-deflation anti-burst construction; rated to minimum 300 kg static load per ball (large size).

- Surface: Non-slip texturized finish.

- Storage: Vertical steel ring organizer stand for compact storage of multiple balls.

- Cleaning: Wipe-clean with mild soap or hospital disinfectant.

- Applicable Standards: EN 71 (paediatric safety), phthalate-free/RoHS-compliant materials.

## 38.3 Standard Accessories

- Inflatable burst-proof dynamic exercise balls (55 cm, 65 cm, 75 cm): Anti-burst construction, phthalate-free, non-slip surface, minimum 3 balls per set.

- Vertical steel ring organizer storage tree stand: Powder-coated steel, holds all supplied ball sizes, floor-standing.

- High-capacity foot inflation pump assembly: Bi-directional or single-action high-volume pump, compatible fittings.

- Replacement valve plug starter kit: Assorted valve plugs and extraction tool.

## 38.4 Installation Requirements

Floor space for ball rack approximately 0.8 × 0.5 m; therapy floor space minimum 3 × 3 m; level non-slip floor.

## 38.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 38.6 Training

Clinical use orientation.

## 38.7 Warranty

Minimum one (1) year against manufacturing defects; storage rack minimum two (2) years.

# 39. Medicine Ball Set

Quantity: 1

## 39.1 Clinical & Functional Specifications

Weighted medicine ball set for progressive resistance training, functional strengthening, sports rehabilitation, and dynamic proprioceptive exercise. Graded weights shall permit progressive loading in upper-extremity, core, and functional throw/catch/press exercises.

## 39.2 Engineering Technical Specifications

- Weight Range: 1 kg, 2 kg, 3 kg, 4 kg, 5 kg (or equivalent graded set).

- Ball Diameter: Approximately 15—25 cm (consistent diameter across weights).

- Surface: Texturized heavy-duty rubber with high-grip finish.

- Filler: Internal bounce-reduction sand/composite filling for controlled rebound.

- Weight Marking: Clearly labeled weight on ball surface and color-coded per weight.

- Storage: Multi-tier vertical steel storage rack.

- Applicable Standards: EN ISO 20957 where applicable, RoHS-compliant materials.

## 39.3 Standard Accessories

- Texturized heavy-duty rubber weighted medicine balls (1, 2, 3, 4, 5 kg): Non-slip texturized surface, weight-labeled, color-coded, one ball per weight increment.

- Multi-tier vertical steel storage rack stand with non-scratch coating: Powder-coated finish, holds full set, floor-standing.

- Internal bounce-reduction sand filling: Controlled rebound properties, sealed.

## 39.4 Installation Requirements

Rack footprint approximately 0.6 × 0.4 m; therapy floor space minimum 3 × 3 m; level, load-bearing non-slip floor.

## 39.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 39.6 Training

Clinical use orientation.

## 39.7 Warranty

Minimum one (1) year against manufacturing defects; storage rack minimum two (2) years.

# 40. Resistance Tubing Set

Quantity: 2

## 40.1 Clinical & Functional Specifications

Latex resistance tubing set with cushioned handles, ankle cuffs, and door anchor attachment for progressive resistance training, muscle strengthening, rehabilitation of upper- and lower-extremity injuries, and home-exercise programs. Color-coded tension grades shall permit progressive loading.

## 40.2 Engineering Technical Specifications

- Tubing: Heavy-duty latex tubing, minimum 5 tension grades (color-coded from light to extra-heavy), each tube minimum 120 cm working length.

- Tension Grades: Progressive resistance across at least 5 levels (typical range: light \~1—2 kg to extra-heavy \~8—10 kg at standard elongation).

- Handles: Cushioned foam-grip handles with heavy-duty metal clip carabiners (1 pair minimum).

- Ankle Cuffs: Adjustable hook-and-loop ankle cuff straps (1 pair minimum).

- Door Anchor: Foam-block door anchor loop attachment.

- Storage: Textile drawstring travel pouch.

- Cleaning: Wipe-clean; keep away from sharp objects and direct sunlight.

- Applicable Standards: EN ISO 20957 where applicable, latex-content declaration provided.

## 40.3 Standard Accessories

- Heavy-duty latex resistance tubing lines with color-coded tension grades (set of 5): Multiple progressive resistances, each minimum 120 cm.

- Cushioned foam-grip handles with heavy-duty metal clip carabiners (1 pair): Ergonomic, non-slip.

- Adjustable ankle cuff straps (1 pair): Hook-and-loop closure, padded, latex-free padding.

- Door anchor block loop attachment: Foam-block anchor for door-frame installation.

- Textile drawstring travel storage pouch: Compact, washable.

## 40.4 Installation Requirements

None. Optional door frame for door-anchor use.

## 40.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 40.6 Training

Clinical use and patient-education orientation.

## 40.7 Warranty

Minimum one (1) year against manufacturing defects (excluding wear).

# 41. TheraBand (Continuous-Roll Elastic Resistance Bands)

Quantity: 4

## 41.1 Clinical & Functional Specifications

Continuous-roll progressive-resistance latex elastic bands for graded resistance training, rehabilitation exercise, and functional strengthening in physiotherapy and occupational therapy. Color-coded grades shall permit progressive clinical loading across the full spectrum of patient strength levels.

## 41.2 Engineering Technical Specifications

- Roll Length: Minimum 5.5 m per roll (continuous, cut-to-length as required).

- Width: Approximately 12—15 cm.

- Resistance Grades: Full color-coded progressive set (typical range: yellow → red → green → blue → black → silver → gold; minimum 6 grades supplied).

- Material: Latex or latex-free (per tender requirement); powder-free finish.

- Storage: Heavy-duty plastic dispensing organizer rack box for multi-roll storage.

- Cutting Tool: Dedicated scissors instrument.

- Handles/Adapters: Plastic locking band clips or handle adapters for hand grip.

- Applicable Standards: RoHS-compliant, latex-content declaration provided.

## 41.3 Standard Accessories

- Continuous-roll progressive resistance latex bands (color-coded set): Full color-graded set (minimum 6 grades), each minimum 5.5 m.

- Heavy-duty plastic storage dispensing organizer rack box: Multi-slot roll storage.

- TheraBand scissor cutting tool: Dedicated heavy-duty scissors.

- Plastic locking band clips / handle adapters: Enable band-end handle attachment.

## 41.4 Installation Requirements

None.

## 41.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 41.6 Training

Clinical use and patient-education orientation.

## 41.7 Warranty

Minimum one (1) year against manufacturing defects (excluding wear).

# 42. Foam Roller

Quantity: 3

## 42.1 Clinical & Functional Specifications

High-density foam cylinders for myofascial release, self-mobilization, muscle recovery, postural exercises, core stabilization, and rehabilitation exercise. Provided in full-round and half-round configurations with a texturized grid-pattern option for deeper tissue mobilization.

## 42.2 Engineering Technical Specifications

- Material: High-density closed-cell EVA foam or equivalent.

- Full-Round Roller Dimensions: Diameter approximately 15 cm, length approximately 90 cm.

- Half-Round Roller Dimensions: Half-cylinder cross-section, length approximately 90 cm.

- Texturized Grid Roller: Full-round with molded/textured grid surface for deep-tissue mobilization.

- Density: High-density, load-tested to minimum 130 kg user weight, retains shape after prolonged use.

- Cleaning: Wipe-clean with mild soap or hospital disinfectant.

- Storage: Protective plastic slip wrap bags.

- Applicable Standards: RoHS-compliant, non-toxic materials.

## 42.3 Standard Accessories

- High-density closed-cell EVA foam cylinders (full round, 90 cm): Standard smooth-surface roller.

- Half-round high-density foam cylinder rollers (90 cm): Half-cylinder for balance/stability exercises.

- Texturized grid-pattern surface roller: Deep-tissue mobilization roller.

- Protective plastic slip wrap bags: Storage protection.

## 42.4 Installation Requirements

None. Requires level non-slip floor for use.

## 42.5 Documentation

Product Information, Exercise Guide, Warranty Certificate.

## 42.6 Training

Clinical use orientation.

## 42.7 Warranty

Minimum one (1) year against manufacturing defects.

# 43. Foam Wedge Set

Quantity: 1

## 43.1 Clinical & Functional Specifications

Structural foam positioning wedge set for patient positioning, postural support, elevation of limbs, spinal alignment during exercise, and paediatric therapeutic positioning. Provided in various inclination angles for versatile clinical use. Wipe-clean vinyl covers permit infection-control compatibility.

## 43.2 Engineering Technical Specifications

- Wedge Angles: Minimum 3 inclination angles supplied (typical range: 15°, 30°, 45°).

- Standard Dimensions: Approximately 50 × 50 × 15—30 cm (varying per angle).

- Foam: High-density structural foam, load-tested to minimum 130 kg user weight, shape-retentive.

- Cover: Wipe-clean fluid-proof antimicrobial vinyl with hidden zippers, seam-sealed.

- Interlocking: Hook-and-loop stabilization panels on base edges for multi-wedge assembly.

- Cleaning: Wipe-clean with hospital disinfectants.

- Applicable Standards: Antimicrobial cover material, RoHS-compliant, phthalate-free.

## 43.3 Standard Accessories

- High-density structural foam positioning wedges (various inclination angles): Minimum 3 wedges, load-tested, shape-retentive.

- Wipe-clean fluid-proof antimicrobial vinyl cover jackets with hidden zippers: Seam-sealed, hospital-grade.

- Interlocking hook-and-loop stabilization panels on base edges: For multi-wedge assembly and positioning.

## 43.4 Installation Requirements

None. Storage shelf recommended.

## 43.5 Documentation

Product Information, Cleaning Instructions, Warranty Certificate.

## 43.6 Training

Clinical use orientation (positioning techniques, contraindications).

## 43.7 Warranty

Minimum one (1) year against manufacturing defects.

# 44. Dumbbell Set with Hand Straps (1 kg)

Quantity: 1

## 44.1 Clinical & Functional Specifications

Vinyl or neoprene coated fixed-weight dumbbell pair with integrated hand safety straps for upper-extremity resistance training in geriatric, neurological, and orthopedic rehabilitation. Hand straps shall permit secure retention for patients with weak grip.

## 44.2 Engineering Technical Specifications

- Weight: 1 kg per dumbbell (fixed weight, 1 pair).

- Weight Tolerance: ±5%.

- Construction: Solid iron core with vinyl or neoprene coating (non-slip, non-marring).

- Hand Straps: Adjustable integrated nylon webbing safety straps with hook-and-loop closure.

- Grip: Ergonomic contoured or straight grip.

- Storage: Compact plastic desktop pyramid storage rack.

- Applicable Standards: EN ISO 20957 where applicable, RoHS-compliant coating.

## 44.3 Standard Accessories

- Vinyl or neoprene coated solid iron dumbbells (1 kg, 1 pair): Non-slip coating, non-marring, weight-tolerance ±5%.

- Adjustable nylon webbing hand safety straps with hook-and-loop closure: For weak-grip retention.

- Compact plastic desktop pyramid storage rack: Holds supplied dumbbell pair.

## 44.4 Installation Requirements

None.

## 44.5 Documentation

Product Information, Warranty Certificate.

## 44.6 Training

Clinical use orientation.

## 44.7 Warranty

Minimum one (1) year against manufacturing defects.

# 45. Wrist Weight (1 kg)

Quantity: 1

## 45.1 Clinical & Functional Specifications

Flexible neoprene weighted wrist wrap pair for upper-extremity resistance training, functional strengthening, and rehabilitation. Iron-sand shot filling permits conformal fit around the wrist during dynamic exercise.

## 45.2 Engineering Technical Specifications

- Weight: 1 kg per wrap (1 pair).

- Weight Tolerance: ±5%.

- Filling: Iron-sand shot, evenly distributed, sealed within compartmentalized pockets.

- Cover: Flexible neoprene, moisture-wicking interior lining.

- Closure: Extra-long nylon webbing strap with heavy-duty D-ring buckle and hook-and-loop closure.

- Adjustable Fit: Wrist circumference range approximately 12—25 cm.

- Cleaning: Wipe-clean, spot-washable.

- Applicable Standards: RoHS-compliant, latex-free.

## 45.3 Standard Accessories

- Flexible neoprene weighted wrist wraps (1 kg, 1 pair): Iron-sand shot filling, compartmentalized, moisture-wicking interior.

- Extra-long nylon webbing strap with heavy-duty D-ring buckle: Adjustable fit, hook-and-loop closure.

- Double-stitched structural seam reinforcement: Filler containment.

## 45.4 Installation Requirements

None.

## 45.5 Documentation

Product Information, Warranty Certificate.

## 45.6 Training

Clinical use orientation.

## 45.7 Warranty

Minimum one (1) year against manufacturing defects.

# 46. Ankle Weight (1 kg)

Quantity: 1

## 46.1 Clinical & Functional Specifications

Anatomically contoured neoprene weighted ankle wrap pair for lower-extremity resistance training, gait re-training, hip and knee strengthening, and rehabilitation exercise.

## 46.2 Engineering Technical Specifications

- Weight: 1 kg per wrap (1 pair).

- Weight Tolerance: ±5%.

- Filling: Iron-sand shot, evenly distributed, sealed within compartmentalized pockets.

- Cover: Anatomically contoured neoprene, moisture-wicking interior lining.

- Closure: Heavy-duty hook-and-loop closure with secure pull-through loops.

- Adjustable Fit: Ankle circumference range approximately 20—35 cm.

- Cleaning: Wipe-clean, spot-washable.

- Applicable Standards: RoHS-compliant, latex-free.

## 46.3 Standard Accessories

- Anatomically contoured neoprene ankle weight wraps (1 kg, 1 pair): Iron-sand shot filling, compartmentalized.

- Heavy-duty hook-and-loop closure straps with pull-through loops: Secure retention during dynamic exercise.

- Soft interior moisture-wicking fabric linings: Comfort during prolonged use.

## 46.4 Installation Requirements

None.

## 46.5 Documentation

Product Information, Warranty Certificate.

## 46.6 Training

Clinical use orientation.

## 46.7 Warranty

Minimum one (1) year against manufacturing defects.

# 47. Mobilization Belt

Quantity: 3

## 47.1 Clinical & Functional Specifications

Heavy-duty woven manual therapy mobilization belt for therapist-assisted joint mobilization, mulligan mobilization-with-movement techniques, spinal mobilization, and manual therapy techniques requiring sustained directional force through the therapist's body weight. The belt shall enable ergonomic therapist positioning while delivering controlled, sustained mobilizing forces to the patient.

## 47.2 Engineering Technical Specifications

- Length: Minimum 2.5 m working length.

- Width: Approximately 5—7.5 cm.

- Material: Heavy-duty woven nylon or polyester webbing.

- Load Rating: Minimum 250 kg working load.

- Buckle: Quick-release industrial-grade composite side-squeeze buckle mechanism.

- Padding: Sliding soft-padded neoprene comfort sleeves for bony-prominence protection.

- Cleaning: Wipe-clean or hand-wash.

- Applicable Standards: Load-tested per institutional safety requirements, RoHS-compliant materials.

## 47.3 Standard Accessories

- Heavy-duty woven nylon manual therapy belt strap (2.5 m): Load-rated, tear-resistant, colorfast.

- Quick-release industrial-grade composite side-squeeze buckle: Rapid disengagement under load.

- Sliding soft-padded neoprene comfort sleeves: For bony-prominence protection, repositionable.

## 47.4 Installation Requirements

None.

## 47.5 Documentation

Product Information, Load-Test Certificate, Warranty Certificate.

## 47.6 Training

Clinical technique training (manual therapy applications).

## 47.7 Warranty

Minimum one (1) year against manufacturing defects.

# 48. Finger Wall Ladder (Shoulder Finger Ladder)

Quantity: 1

## 48.1 Clinical & Functional Specifications

Wall-mounted vertical hardwood finger-ladder panel with graduated finger-step notches for shoulder range-of-motion exercise, forward-flexion and abduction re-training, and post-operative shoulder rehabilitation. Enables patient self-directed progressive range-of-motion training with visual measurement of achieved reach.

## 48.2 Engineering Technical Specifications

- Panel Material: Hardwood (birch, beech, or equivalent) or reinforced composite.

- Panel Dimensions: Approximately 15—20 cm width × 130—160 cm length.

- Finger Steps: Minimum 20 graduated finger-step notches or ridges, evenly spaced.

- Graduation: Numbered or indexed for progress tracking (cm or increments).

- Finish: Glossy clear anti-bacterial varnish coating, wipe-cleanable.

- Mounting: Wall-mount vertical orientation, structural anchoring hardware supplied.

- Applicable Standards: RoHS-compliant, non-toxic finish.

## 48.3 Standard Accessories

- Hardwood vertical wall-mount panel board with 20+ graduated finger steps: Numbered/indexed steps, anti-bacterial varnish finish.

- Heavy-duty wall-anchoring screw kit and structural spacers: Wall-fastening hardware for concrete or drywall mounting.

- Glossy clear anti-bacterial varnish coating layer: Wipe-cleanable, hygienic.

## 48.4 Installation Requirements

Wall-mount location with structural anchor points; mount height configurable per patient population (typical center-line at approximately 150 cm from floor); ambient 10—35 °C. Professional installation recommended.

## 48.5 Documentation

Installation Manual, Product Information, Warranty Certificate.

## 48.6 Training

Clinical use orientation.

## 48.7 Warranty

Minimum two (2) years against manufacturing defects.

ASSESSMENT & DIAGNOSTIC INSTRUMENTS

# 49. Chiropractic / Percussion Massage Gun

Quantity: 1

## 49.1 Clinical & Functional Specifications

Handheld high-torque percussion massage device for deep-tissue myofascial mobilization, muscle recovery, trigger-point release, and adjunctive treatment of musculoskeletal pain and muscle tightness. Shall provide adjustable percussion frequency and stroke depth, multiple interchangeable therapy heads for varied anatomical regions, and rechargeable battery operation for cordless use.

## 49.2 Engineering Technical Specifications

- Motor: Brushless DC motor, high-torque.

- Percussion Frequency: Adjustable, minimum 3 speed levels, range approximately 1200—3200 percussions/minute.

- Stroke Amplitude: Approximately 10—16 mm.

- Stall Force: Minimum 20 kg.

- Noise Level: Approximately 40—55 dBA (typical operating).

- Battery: Rechargeable lithium-ion, minimum 2000 mAh; runtime minimum 2 hours continuous per charge; charging time maximum 4 hours.

- Auto Shut-Off: Timer-controlled, typically 10-minute session limit.

- Display: LCD indicator for speed level and battery status.

- Weight: Approximately 0.8—1.5 kg.

- Applicable Standards: IEC 60601-1 (if classified as medical device) or IEC 60335 (consumer appliance), IEC 60601-1-2 EMC, RoHS-compliant.

## 49.3 Standard Accessories

- High-torque brushless motor percussion massager handpiece: Ergonomic housing, non-slip grip, integrated speed controls.

- Interchangeable therapy head attachments (Ball, Bullet, Fork, Flat): Minimum 4 heads, silicone or high-density foam construction, tool-free attachment.

- Rechargeable internal lithium-ion battery pack: Pre-installed, protected against overcharge and short-circuit.

- Desktop battery charging cradle station with power adapter: Universal input 100—240 VAC, 50/60 Hz.

- Rigid zippered foam-lined protective storage case: Foam-cut compartments for handpiece and all attachments.

## 49.4 Installation Requirements

100—240 VAC, 50/60 Hz outlet for charger; no fixed installation.

## 49.5 Documentation

User Manual, Charging Instructions, Factory Test Certificate, Warranty Certificate.

## 49.6 Training

Clinical use orientation (application zones, contraindications, safety).

## 49.7 Warranty

Minimum two (2) years comprehensive warranty on motor and battery.

# 50. Goniometry Set

Quantity: 3

## 50.1 Clinical & Functional Specifications

Set of manual goniometers for measurement of joint range of motion (ROM) in physical assessment, rehabilitation progress tracking, and clinical documentation. Set shall include large, medium, and small (finger) goniometers to accommodate all joint sizes.

## 50.2 Engineering Technical Specifications

- Standard 12-inch Goniometer: Clear plastic (transparent polycarbonate or equivalent), 360° scale, 1° increments, arm length approximately 30 cm.

- Medium 8-inch Goniometer: Clear plastic, 180° scale, 1° increments, arm length approximately 20 cm.

- Finger Goniometer: Stainless steel, small joint measurement, 180° scale, 1° increments, arm length approximately 6—8 cm.

- Accuracy: ±1° across full scale.

- Scale Marking: Permanent laser-etched or screen-printed, non-fading.

- Applicable Standards: ISO 13485 where applicable, general medical device requirements.

## 50.3 Standard Accessories

- Standard 360° clear plastic 12-inch goniometer: Full-circle scale, 1° increments, ±1° accuracy.

- Medium 180° clear plastic 8-inch goniometer: Half-circle scale, 1° increments.

- Small stainless steel finger goniometer: For small joints (fingers, toes, wrist), 180° scale.

- Snap-closure transparent plastic storage sleeve: Protects instruments during transport.

## 50.4 Installation Requirements

None.

## 50.5 Documentation

Product Information, Warranty Certificate.

## 50.6 Training

Clinical use orientation.

## 50.7 Warranty

Minimum one (1) year against manufacturing defects.

# 51. Inclinometer

Quantity: 1

## 51.1 Clinical & Functional Specifications

Fluid-filled mechanical circular dial inclinometer for measurement of spinal range of motion, joint angle, and postural inclination during clinical assessment and rehabilitation documentation. The gravity-weighted pointer shall provide gravity-referenced angle measurement independent of instrument orientation.

## 51.2 Engineering Technical Specifications

- Scale: 360° circular dial, 1° increments (or 2° per tender preference).

- Measurement Principle: Fluid-damped weighted gravity pointer.

- Accuracy: ±1° or better.

- Zero Adjustment: Adjustable outer baseline ring for zero-offset calibration to any reference position.

- Housing: Rigid plastic or aluminum, transparent front lens, sealed fluid chamber.

- Applicable Standards: ISO 13485 where applicable.

## 51.3 Standard Accessories

- Fluid-filled mechanical circular dial inclinometer: 360° scale, weighted gravity pointer, sealed fluid damping.

- Weighted gravity pointer needle element: Fluid-damped for smooth reading.

- Adjustable zero-offset outer baseline ring: Tool-free rotation for zeroing.

- Rigid padded pocket storage container: Protects instrument during transport.

## 51.4 Installation Requirements

None.

## 51.5 Documentation

Product Information, Calibration Certificate, Warranty Certificate.

## 51.6 Training

Clinical use orientation.

## 51.7 Warranty

Minimum one (1) year against manufacturing defects.

# 52. Dynamometer (Hand-Grip)

Quantity: 1

## 52.1 Clinical & Functional Specifications

Hydraulic hand-grip dynamometer (Jamar-style) for standardized assessment of grip strength in rehabilitation, occupational therapy, hand surgery follow-up, and geriatric functional assessment. Shall provide peak-hold cursor for retention of maximum grip force after release.

## 52.2 Engineering Technical Specifications

- Measurement Principle: Sealed hydraulic system.

- Measurement Range: 0—90 kg (0—200 lb) minimum.

- Accuracy: ±5% of reading or ±2 kg, whichever is greater.

- Handle: Five-position adjustable grip spacing, spring-loaded lock pin, standardized positions per Jamar protocol.

- Display: Dual-scale analog dial (kg + lb) with peak-hold cursor tracking maximum force.

- Reset: Manual peak-cursor reset knob.

- Applicable Standards: ISO 13485 where applicable, general medical device requirements.

## 52.3 Standard Accessories

- Hydraulic hand-grip dynamometer (Jamar-style, 90 kg capacity): Sealed hydraulic, dual-scale.

- Adjustable five-position handle grip spacing bar with spring lock pin: Standardized positions.

- Dual-scale analog readout dial with peak-hold cursor: Kg and lb scales, peak retention.

- Fitted velvet-lined rigid plastic protective storage case: Foam-cut, secure transport.

## 52.4 Installation Requirements

None.

## 52.5 Documentation

User Manual, Factory Calibration Certificate, Warranty Certificate.

## 52.6 Training

Clinical use orientation (standardized grip protocol).

## 52.7 Warranty

Minimum one (1) year against manufacturing defects; hydraulic seal warranted for two (2) years.

# 53. Algometer (Pressure Pain Threshold)

Quantity: 1

## 53.1 Clinical & Functional Specifications

Digital pressure algometer for quantitative measurement of pressure pain threshold in the assessment of myofascial trigger points, fibromyalgia tender points, and quantitative sensory testing during clinical assessment and research.

## 53.2 Engineering Technical Specifications

- Measurement Principle: Digital load cell.

- Measurement Range: 0—10 kg/cm² (approximately 0—100 N/cm²) minimum.

- Accuracy: ±5% of reading or better.

- Resolution: 0.1 kg/cm² or better.

- Display: Digital LCD with peak-hold force display.

- Peak-Hold: Retains maximum applied force until manual reset.

- Applicator Tip: Standard 1 cm² rubber-tipped applicator (calibration standard).

- Power: Long-life coin-cell or standard battery, minimum 100 hours operation.

- Applicable Standards: ISO 13485 where applicable.

## 53.3 Standard Accessories

- Digital pressure pain threshold algometer: Microprocessor-based load cell, digital LCD.

- Rubber-tipped applicator tips (1 cm²): Calibration-standard surface area.

- Peak force memory tracking software: Integrated in display.

- Long-life coin-cell battery pack with soft transport pouch: Pre-installed battery.

## 53.4 Installation Requirements

None.

## 53.5 Documentation

User Manual, Factory Calibration Certificate, Warranty Certificate.

## 53.6 Training

Clinical use orientation (standardized testing protocol).

## 53.7 Warranty

Minimum one (1) year against manufacturing defects.

# 54. Scoliometer

Quantity: 1

## 54.1 Clinical & Functional Specifications

Contoured spinal inclination measurement device for screening of scoliosis and postural deviation through Adam's forward-bend test. Shall provide rapid clinical measurement of the angle of trunk rotation (ATR) during postural screening.

## 54.2 Engineering Technical Specifications

- Measurement Principle: Fluid-filled level tube with bubble indicator.

- Measurement Range: ±30° angle of trunk rotation minimum.

- Accuracy: ±1° or better.

- Scale: Millimeter-graduated posture deviation markings; degree indicators.

- Body Guide: Contoured plastic profile matching spinal geometry, non-invasive contact.

- Applicable Standards: ISO 13485 where applicable, general medical device requirements.

## 54.3 Standard Accessories

- Contoured plastic alignment body guide profile: Matches spinal contour, non-invasive.

- Fluid-filled horizontal level tube with air bubble: Sealed, damped.

- Millimeter posture deviation tracking rule markings: Permanent, non-fading.

- Protective synthetic leather storage pouch: Slide-in storage.

## 54.4 Installation Requirements

None.

## 54.5 Documentation

Product Information, Calibration Certificate, Warranty Certificate.

## 54.6 Training

Clinical use orientation (Adam's forward-bend protocol).

## 54.7 Warranty

Minimum one (1) year against manufacturing defects.

# 55. Stopwatch

Quantity: 2

## 55.1 Clinical & Functional Specifications

Digital quartz electronic stopwatch for timing of clinical assessments, exercise protocols, timed functional tests (6-minute walk test, Timed Up and Go, gait velocity, etc.), and therapeutic session timing.

## 55.2 Engineering Technical Specifications

- Timing Range: 0 to minimum 99 hours 59 minutes 59.99 seconds.

- Resolution: 0.01 second (1/100 sec).

- Accuracy: ±0.001% (quartz crystal reference).

- Functions: Start/Stop, Lap/Split time, Reset.

- Display: Multi-line LCD with backlight (optional).

- Battery: Long-life lithium coin cell, minimum 2 years typical life.

- Water Resistance: IPX4 minimum (splash-resistant).

- Housing: Impact-resistant plastic with protective rubber perimeter.

- Lanyard: Woven neck cord.

- Applicable Standards: RoHS-compliant.

## 55.3 Standard Accessories

- Digital quartz electronic stopwatch: Lap/split tracking, multi-line LCD.

- Woven neck lanyard cord: Anti-loss attachment.

- Long-life lithium coin-cell battery: Pre-installed.

- Impact-resistant rubber perimeter casing: Protects during drops.

## 55.4 Installation Requirements

None.

## 55.5 Documentation

User Manual, Warranty Certificate.

## 55.6 Training

None required.

## 55.7 Warranty

Minimum one (1) year against manufacturing defects.

# 56. Measuring Tape (Anthropometric)

Quantity: 2

## 56.1 Clinical & Functional Specifications

Flexible retractable measuring tape for clinical anthropometry, limb girth/circumference measurement, wound measurement, and functional assessment.

## 56.2 Engineering Technical Specifications

- Length: Minimum 1.5 m (approximately 60 inches).

- Material: Flexible fiberglass or coated fabric tape, non-stretch.

- Scale: Dual metric (cm/mm) and imperial (inches) markings.

- Graduation: 1 mm minimum resolution.

- Retraction: Spring-assisted automatic push-button retraction.

- End Hook: Spring-steel end clip for anchoring.

- Case: Compact plastic housing.

- Cleaning: Wipe-clean.

- Applicable Standards: RoHS-compliant.

## 56.3 Standard Accessories

- Flexible fiberglass anthropometric measuring tape (1.5 m): Dual scale, non-stretch.

- Spring-assisted retraction mechanism case: Push-button retract.

- Spring-steel end-clip hook: Anchor point for one-person measurement.

## 56.4 Installation Requirements

None.

## 56.5 Documentation

Product Information, Warranty Certificate.

## 56.6 Training

None required.

## 56.7 Warranty

Minimum one (1) year against manufacturing defects.

# 57. Reflex Hammer Kit

Quantity: 2

## 57.1 Clinical & Functional Specifications

Set of clinical reflex hammers for neurological examination including deep tendon reflex testing, plantar reflex testing, and general neurological assessment. Set shall include multiple hammer types for versatile clinical use.

## 57.2 Engineering Technical Specifications

- Taylor Hammer: Weighted metal handle with solid rubber triangular ("tomahawk") wedge head; overall length approximately 18—22 cm.

- Babinski Hammer: Telescoping percussion handle with rotatable circular brass disc head; extendable design; overall extended length approximately 25—35 cm.

- Buck Hammer: Metal handle with dual rubber heads; integrated unscrewing needle and bristle brush tool for sensory testing.

- Case: Zippered synthetic leather organizer wallet with individual instrument slots.

- Applicable Standards: ISO 13485 where applicable, RoHS-compliant.

## 57.3 Standard Accessories

- Taylor Tomahawk weighted reflex hammer: Rubber triangular head, metal handle.

- Babinski telescoping percussion hammer with rotatable brass disc: Telescoping handle.

- Buck hammer with integrated needle and bristle brush: Multi-tool neurological hammer.

- Zippered synthetic leather organizing wallet case: Individual slots per instrument.

## 57.4 Installation Requirements

None.

## 57.5 Documentation

Product Information, Warranty Certificate.

## 57.6 Training

Clinical use orientation.

## 57.7 Warranty

Minimum one (1) year against manufacturing defects.

# 58. Tuning Fork Kit

Quantity: 2

## 58.1 Clinical & Functional Specifications

Set of medical tuning forks for neurological vibration sensation testing, hearing assessment (Rinne and Weber tests), and general sensory evaluation. Set shall include multiple standard frequencies for versatile clinical use.

## 58.2 Engineering Technical Specifications

- 128 Hz Tuning Fork: Weighted foot base for vibration sense testing; frequency accuracy ±5%.

- 256 Hz Tuning Fork: Standard clinical hearing/sensation testing; frequency accuracy ±5%.

- 512 Hz Tuning Fork: Rinne/Weber hearing assessment; frequency accuracy ±5%.

- Material: Non-magnetic stainless steel or aluminum alloy.

- Striking Block: Hard rubber mallet block for consistent activation.

- Case: Velvet-lined rigid wooden or metal organizer.

- Applicable Standards: ISO 13485 where applicable, RoHS-compliant, frequency accuracy per international acoustic standards.

## 58.3 Standard Accessories

- Tuning fork 128 Hz with weighted foot base: For vibration sensation testing.

- Tuning fork 256 Hz: Standard sensory/hearing frequency.

- Tuning fork 512 Hz: Hearing assessment frequency.

- Hard rubber striking block mallet: Consistent activation tool.

- Velvet-lined rigid case organizer: Individual instrument slots.

## 58.4 Installation Requirements

None.

## 58.5 Documentation

Product Information, Frequency Certification, Warranty Certificate.

## 58.6 Training

Clinical use orientation.

## 58.7 Warranty

Minimum one (1) year against manufacturing defects.

# 59. Sensory Examination Kit

Quantity: 2

## 59.1 Clinical & Functional Specifications

Comprehensive sensory examination kit for neurological sensory testing including sharp/dull discrimination, thermal sensation, and stereognostic assessment. Shall support standardized bedside neurological evaluation.

## 59.2 Engineering Technical Specifications

- Sharp/Blunt Pins: Neurological dual-tip testing pins, disposable or sterilizable per infection-control policy; sharp tip and blunt tip on opposite ends.

- Wartenberg Pinwheel: Rolling spur wheel with ergonomic handle, spring-loaded rotation; stainless steel construction.

- Thermal Tubes: Sealed metal tubes (1 pair) for hot/cold water filling, corrosion-resistant.

- Case: Rigid plastic organizer with individual instrument slots.

- Applicable Standards: ISO 13485 where applicable, RoHS-compliant.

## 59.3 Standard Accessories

- Neurological dual-tip sharp/blunt testing pins: Set of pins, dual-purpose.

- Wartenberg pinwheel diagnostic instrument: Stainless steel, rolling spur.

- Thermal evaluation metal tubes (1 pair): Corrosion-resistant, for hot/cold testing.

- Rigid plastic containment organizing case: Individual instrument slots.

## 59.4 Installation Requirements

None.

## 59.5 Documentation

Product Information, Warranty Certificate.

## 59.6 Training

Clinical use orientation.

## 59.7 Warranty

Minimum one (1) year against manufacturing defects.

# 60. Monofilaments (Semmes-Weinstein Sensory Testing)

Quantity: 1

## 60.1 Clinical & Functional Specifications

Standardized calibrated monofilament for quantitative sensory testing of protective sensation in diabetic peripheral neuropathy assessment, foot ulcer risk screening, and general peripheral nerve function evaluation.

## 60.2 Engineering Technical Specifications

- Standard Monofilament: 10 g calibrated force (Semmes-Weinstein 5.07 or equivalent), the international standard for diabetic foot screening.

- Force Accuracy: ±10% of nominal force at calibrated buckling point.

- Filament: Nylon fiber, calibrated bend point.

- Handle: Rigid handle with permanent force marking.

- Replacement Filaments: Additional nylon fiber lines (pack of 5 minimum).

- Storage: Fitted slide pocket protection sleeve.

- Applicable Standards: ISO 13485 where applicable, meets Semmes-Weinstein calibration standard.

## 60.3 Standard Accessories

- Standard 10 g sensory evaluation monofilament wand: Calibrated buckling force.

- Replacement nylon fiber monofilament lines (pack of 5): For handle-based systems.

- Fitted slide protection pocket storage sleeve: Prevents filament damage.

## 60.4 Installation Requirements

None.

## 60.5 Documentation

Product Information, Calibration Certificate, Warranty Certificate.

## 60.6 Training

Clinical use orientation (diabetic foot screening protocol).

## 60.7 Warranty

Minimum one (1) year against manufacturing defects.

# 61. Cotton Balls (Sensory Kit)

Quantity: 3

## 61.1 Clinical & Functional Specifications

Medical-grade cotton balls for light-touch sensation testing during neurological examination and for general clinical use in rehabilitation departments.

## 61.2 Engineering Technical Specifications

- Material: Medical-grade 100% pure cotton fiber.

- Quantity: Minimum 100 pieces per pack.

- Diameter: Approximately 2 cm per ball.

- Storage Jar: Clear glass or stainless steel cylindrical container with lid.

- Forceps: Surgical dressing extraction forceps for sterile handling.

- Applicable Standards: Medical-grade cotton per applicable pharmacopoeia; RoHS-compliant storage materials.

## 61.3 Standard Accessories

- Medical-grade pure cotton fiber balls (minimum 100 pieces): Per pack.

- Clear glass or stainless steel cylindrical storage jar with lid: Reusable, sterilizable.

- Surgical dressing extraction forceps: Stainless steel, autoclave-compatible.

## 61.4 Installation Requirements

None.

## 61.5 Documentation

Product Information, Certificate of Analysis (medical-grade cotton).

## 61.6 Training

None required.

## 61.7 Warranty

Not applicable to consumable; storage jar and forceps minimum one (1) year.

# 62. Stereognosis Kit

Quantity: 2

## 62.1 Clinical & Functional Specifications

Assessment kit for stereognostic (tactile object recognition) sensory testing in occupational therapy, cortical sensory assessment, and cognitive-perceptual evaluation. Shall include an assortment of common geometric shapes and everyday objects for standardized blind-identification testing.

## 62.2 Engineering Technical Specifications

- Blindfold: Opaque fabric patient mask with adjustable elastic head strap; wipe-cleanable.

- Test Objects: Assortment of common geometric shapes and everyday plastic tokens (set of 20 minimum), varying in shape, size, and texture; non-toxic materials; rounded edges for safety.

- Sorting Grid: Wooden or plastic modular sorting tray with individual object slots.

- Storage: Drawstring cloth bag for blind object retrieval.

- Applicable Standards: EN 71 (safety), RoHS-compliant materials.

## 62.3 Standard Accessories

- Opaque fabric patient eye-blindfold mask with elastic strap: Wipe-cleanable, adjustable.

- Assortment of geometric shapes and plastic tokens (set of 20): Varied shapes and textures.

- Wooden or plastic sorting grid tray: Individual slots for organized storage.

- Drawstring cloth blind bag: For blind object retrieval during testing.

## 62.4 Installation Requirements

None.

## 62.5 Documentation

Product Information, Warranty Certificate.

## 62.6 Training

Clinical use orientation.

## 62.7 Warranty

Minimum one (1) year against manufacturing defects.

# 63. Barognosis Kit

Quantity: 2

## 63.1 Clinical & Functional Specifications

Assessment kit for barognostic (weight discrimination) sensory testing in occupational therapy and cortical sensory evaluation. Identical-appearance cylinders of graduated weights shall permit standardized blind weight-discrimination testing.

## 63.2 Engineering Technical Specifications

- Cylinders: Identical-size plastic cylinders (set of 8 minimum) with graduated internal weights, externally indistinguishable.

- Weight Range: Graduated across a standardized clinical range (typical 5—100 g or equivalent), with color-coded or bottom-stamped markers for therapist identification.

- Storage: Wooden desk organizer tray.

- Applicable Standards: RoHS-compliant materials, general medical device requirements.

## 63.3 Standard Accessories

- Identical-size geometric plastic cylinder blocks with graduated weights (set of 8): Externally indistinguishable, weight-graded.

- Color-coded or bottom-stamped weight markers: Therapist identification only.

- Wooden desk organizer tray: Individual slots per cylinder.

## 63.4 Installation Requirements

None.

## 63.5 Documentation

Product Information, Warranty Certificate.

## 63.6 Training

Clinical use orientation.

## 63.7 Warranty

Minimum one (1) year against manufacturing defects.

# 64. Cranial Nerve Examination Kit

Quantity: 2

## 64.1 Clinical & Functional Specifications

Comprehensive cranial nerve examination kit for bedside neurological assessment including olfactory testing (CN I), visual acuity (CN II), pupillary reactions and extraocular movements (CN III/IV/VI), and lingual/oropharyngeal examination (CN IX/X/XII).

## 64.2 Engineering Technical Specifications

- Olfactory Test Vials: Scented oil vials in standard test odors (typical set: peppermint, vanilla, coffee, or equivalent), sealed dropper bottles.

- Diagnostic Penlight: Miniature high-intensity LED penlight with pocket clip; battery-operated; adjustable or fixed focus; brightness minimum 30 lumens; pupil-gauge markings on handle.

- Snellen Visual Acuity Card: Pocket-size Snellen chart, standard optotype, printed on non-fading rigid card.

- Tongue Depressors: Disposable plastic tongue depressor blades, minimum 50 pieces per box, standard length.

- Applicable Standards: ISO 13485 where applicable, RoHS-compliant, disposable items medical-grade.

## 64.3 Standard Accessories

- Scented oil vials for olfactory evaluation (starter set): Peppermint, vanilla, coffee, or equivalent standard odors.

- Miniature high-intensity LED diagnostic penlight: Pocket clip, pupil gauge, battery-operated.

- Snellen pocket visual acuity eye card: Standard optotype chart.

- Disposable plastic tongue depressor blades (box of 50): Medical-grade, single-use.

## 64.4 Installation Requirements

None.

## 64.5 Documentation

Product Information, Warranty Certificate.

## 64.6 Training

Clinical use orientation.

## 64.7 Warranty

Minimum one (1) year against manufacturing defects (excluding consumables).

OCCUPATIONAL THERAPY & SENSORY INTEGRATION EQUIPMENT

# 65. Sensory Integration Kit

Quantity: 1

## 65.1 Clinical & Functional Specifications

Comprehensive sensory integration therapy kit for paediatric neurodevelopmental therapy, autism spectrum disorder, sensory processing disorders, and general occupational therapy. Shall provide a variety of tactile, proprioceptive, vestibular, and auditory sensory input tools for structured sensory-integration therapy sessions.

## 65.2 Engineering Technical Specifications

- Sensory Floor Tiles: Minimum 6 tiles with varied textured surfaces (smooth, bumpy, ridged, soft, hard, and equivalent), edge dimensions approximately 30 × 30 cm each; non-toxic materials; interlocking or standalone configuration.

- Weighted Blanket: Approximately 2 kg weight profile, evenly distributed filling in compartmentalized pockets, wipe-clean or washable cover, hypoallergenic materials.

- Vibrating Massage Ball: Rechargeable handheld capsule with adjustable vibration intensity, ergonomic non-slip grip, minimum 60 minutes runtime per charge.

- Noise-Dampening Earmuffs: Passive hearing protection, noise reduction rating (NRR) minimum 20 dB, adjustable padded headband, one pair.

- Storage: Heavy-duty plastic storage bin with lid, sized to contain all kit components.

- Applicable Standards: EN 71 (paediatric safety), RoHS-compliant, non-toxic materials, phthalate-free.

## 65.3 Standard Accessories

- Texturized sensory floor tile mats (set of 6): Varied surface geometries, non-toxic, wipe-cleanable.

- Weighted therapy lap blanket (2 kg): Evenly distributed filling, hypoallergenic cover.

- Vibrating handheld massage ball capsule: Rechargeable, adjustable intensity.

- Noise-dampening acoustic earmuff headsets (1 pair): NRR ≥ 20 dB, adjustable.

- Heavy-duty plastic storage bin with lid: Kit containment.

## 65.4 Installation Requirements

None. Sensory therapy room recommended, minimum 3 × 3 m clear floor space.

## 65.5 Documentation

Product Information, Cleaning Instructions, Warranty Certificate.

## 65.6 Training

Clinical application orientation (sensory-integration protocols).

## 65.7 Warranty

Minimum one (1) year against manufacturing defects.

# 66. Fine Motor Equipment

Quantity: 1

## 66.1 Clinical & Functional Specifications

Occupational therapy fine-motor skills training kit for paediatric and adult rehabilitation of dexterity, hand-eye coordination, precision grasp, and cognitive-motor integration. Suitable for hand therapy, developmental delay, stroke rehabilitation, and general occupational therapy.

## 66.2 Engineering Technical Specifications

- Pegboard: Graduated multi-hole wooden pegboard tray with thin metal or wooden pegs; minimum 25 hole positions; graduated peg diameters for progressive precision demand.

- Nut-and-Bolt Board: Matching board with various thread diameters (minimum 5 sizes) for bilateral precision motor training.

- Tweezers & Bead Sorting: Precision tweezers with graduated bead-placement sorting grid; minimum 3 bead sizes; multi-compartment grid.

- Storage: Heavy-duty plastic organization case with individual compartments.

- Applicable Standards: EN 71 (paediatric safety), non-toxic finishes, RoHS-compliant.

## 66.3 Standard Accessories

- Graduated multi-hole wooden pegboard tray with pegs: Multiple peg diameters, minimum 25 positions.

- Nut-and-bolt matching tracking board: Multiple thread sizes, bilateral coordination training.

- Tweezers and precision bead-placement sorting grids: Multiple bead sizes, multi-compartment grid.

- Heavy-duty plastic organization cabinet case: Individual compartments per component.

## 66.4 Installation Requirements

None. Requires tabletop workspace.

## 66.5 Documentation

Product Information, Warranty Certificate.

## 66.6 Training

Clinical application orientation.

## 66.7 Warranty

Minimum one (1) year against manufacturing defects.

# 67. Finger Exerciser

Quantity: 3

## 67.1 Clinical & Functional Specifications

Handheld spring-loaded independent finger-piston exerciser for individual finger strengthening, hand rehabilitation, and grip endurance training. Independent piston design shall permit isolated exercise of each finger.

## 67.2 Engineering Technical Specifications

- Configuration: Handheld device with independent spring-loaded piston buttons for each finger (minimum 4 fingers).

- Resistance Levels: Minimum 3 supplied resistances (typical range: light \~1 kg, medium \~2 kg, heavy \~3 kg per finger).

- Base: Molded ergonomic rubber hand rest cushion.

- Storage: Clear plastic blister storage sleeves.

- Applicable Standards: RoHS-compliant materials, non-toxic finish.

## 67.3 Standard Accessories

- Spring-loaded independent finger button piston devices (3 resistances): Minimum 3 units per set.

- Molded ergonomic rubber hand rest base cushion: Comfort during prolonged use.

- Clear plastic blister storage container sleeves: Protective storage.

## 67.4 Installation Requirements

None.

## 67.5 Documentation

Product Information, Warranty Certificate.

## 67.6 Training

Clinical use orientation.

## 67.7 Warranty

Minimum one (1) year against manufacturing defects.

# 68. Hand Exercise Ball

Quantity: 3

## 68.1 Clinical & Functional Specifications

Solid silicone hand-therapy squeeze ball set for progressive grip strengthening, hand rehabilitation, and stress-relief exercise. Color-coded graduated resistances shall permit clinical progression.

## 68.2 Engineering Technical Specifications

- Shape: Egg-shaped or spherical, ergonomic hand-fit dimensions.

- Material: Solid medical-grade silicone or equivalent, latex-free.

- Resistance Levels: Minimum 3 supplied resistances (soft, medium, firm), color-coded.

- Diameter: Approximately 6—7 cm.

- Cleaning: Wipe-clean, spot-washable.

- Storage: Mesh textile drawstring pouch.

- Applicable Standards: RoHS-compliant, latex-free, non-toxic materials.

## 68.3 Standard Accessories

- Solid silicone egg-shaped hand squeeze balls (set of 3, color-coded resistances): Latex-free, graduated resistance.

- Microfiber fabric cleansing towel: For wipe-cleaning.

- Mesh textile drawstring storage pouch: Compact storage.

## 68.4 Installation Requirements

None.

## 68.5 Documentation

Product Information, Warranty Certificate.

## 68.6 Training

Clinical use orientation.

## 68.7 Warranty

Minimum one (1) year against manufacturing defects.

# 69. Hand Gripper

Quantity: 2

## 69.1 Clinical & Functional Specifications

A-frame mechanical spring hand gripper with adjustable tension for grip strengthening, forearm endurance training, and hand rehabilitation. Integrated digital counter shall record repetitions for training documentation.

## 69.2 Engineering Technical Specifications

- Design: A-frame mechanical spring gripper.

- Adjustable Tension: Continuous adjustment via screw turn-dial; resistance range approximately 5—60 kg.

- Digital Counter: LCD squeeze counter embedded in handle, auto-reset button, battery-operated (long-life coin cell).

- Grip: Texturized non-slip rubber handle skins.

- Frame: Corrosion-resistant metal.

- Applicable Standards: RoHS-compliant materials.

## 69.3 Standard Accessories

- A-frame mechanical spring hand gripper with adjustable tension screw: Continuously adjustable resistance.

- Digital squeeze counter LCD display: Embedded in handle, auto-reset.

- Texturized non-slip rubber grip handle skins: Ergonomic non-slip.

## 69.4 Installation Requirements

None.

## 69.5 Documentation

Product Information, Warranty Certificate.

## 69.6 Training

Clinical use orientation.

## 69.7 Warranty

Minimum one (1) year against manufacturing defects.

# 70. Stack and Sort Board

Quantity: 2

## 70.1 Clinical & Functional Specifications

Wooden geometric stacking and sorting board for paediatric fine-motor development, hand-eye coordination, shape/color recognition, and cognitive-perceptual training in occupational therapy.

## 70.2 Engineering Technical Specifications

- Base: Solid wooden base plate with vertical geometric peg shafts (minimum 3 shafts).

- Sorting Blocks: Multi-color perforated wooden blocks in multiple geometric shapes (triangles, circles, squares); minimum 15 blocks total, graduated sizes.

- Finish: Clear non-toxic polyurethane polish, wipe-cleanable.

- Safety: Rounded edges and corners, non-toxic paint per paediatric safety standards.

- Applicable Standards: EN 71 (paediatric safety), non-toxic finishes.

## 70.3 Standard Accessories

- Solid wooden base plate with vertical geometric peg shafts: Multi-shaft configuration.

- Multi-color perforated wooden sorting blocks (triangles, circles, squares): Minimum 15 pieces, graduated sizes.

- Clear polyurethane non-toxic polish finish: Wipe-cleanable, paediatric-safe.

## 70.4 Installation Requirements

None. Requires tabletop workspace.

## 70.5 Documentation

Product Information, Safety Certificate (paediatric), Warranty Certificate.

## 70.6 Training

Clinical use orientation.

## 70.7 Warranty

Minimum one (1) year against manufacturing defects.

# 71. Velcro Board

Quantity: 1

## 71.1 Clinical & Functional Specifications

Wall-mounted hook-and-loop fabric target board with attachable geometric target pieces for paediatric fine-motor training, upper-extremity range-of-motion exercise, hand-eye coordination, and gamified therapeutic activity.

## 71.2 Engineering Technical Specifications

- Board: Vertical wooden panel covered with heavy-duty loop-fabric surface; minimum dimensions approximately 80 × 60 cm.

- Target Blocks: Geometric shapes with hook-fabric backing; minimum 12 pieces; multiple colors and shapes.

- Mounting Hardware: Heavy-duty wall-anchoring bracket kit.

- Safety: Rounded edges, non-toxic finishes.

- Applicable Standards: EN 71 (paediatric safety), RoHS-compliant.

## 71.3 Standard Accessories

- Wall-mount vertical wooden panel with loop-fabric surface: Full-surface loop coverage.

- Target tracking blocks and geometric shapes with hook-fabric strips (set of 12): Multi-color, multi-shape.

- Heavy-duty wall anchoring hardware: Structural mounting bracket kit.

## 71.4 Installation Requirements

Wall-mount location with structural anchor points; mounting height configurable per patient population; ambient 10—35 °C. Professional installation recommended.

## 71.5 Documentation

Installation Manual, Product Information, Warranty Certificate.

## 71.6 Training

Clinical use orientation.

## 71.7 Warranty

Minimum one (1) year against manufacturing defects.

# 72. ADL Board (Activities of Daily Living Training Board)

Quantity: 1

## 72.1 Clinical & Functional Specifications

Activities of Daily Living training panel for practice and re-education of functional hand skills including manipulation of door hardware, light switches, locks, clothing fasteners, and other everyday mechanical interfaces. Suitable for occupational therapy, hand rehabilitation, paediatric developmental training, and geriatric/neurological rehabilitation.

## 72.2 Engineering Technical Specifications

- Panel: Rigid wooden or high-density plastic board, dimensions approximately 60 × 80 cm.

- Hardware Fixtures: Integrated functional hardware including door lock, door knob, light switch, slide bolt, hasp, and equivalent (minimum 6 mechanical fixtures).

- Clothing Fasteners: Integrated fabric panels with zippers, buttons, buckles, hook-and-loop closures, shoelaces (minimum 5 fastener types).

- Presentation: Sturdy desktop presentation easel fold-stand, or wall-mount configuration.

- Finish: Non-toxic, wipe-cleanable.

- Applicable Standards: EN 71 (safety), RoHS-compliant.

## 72.3 Standard Accessories

- Activities of Daily Living training panel board: Rigid wooden or plastic base.

- Integrated functional hardware fixtures (door locks, knobs, switches, bolts): Minimum 6 mechanical fixtures.

- Integrated clothing fastener panels (zippers, buttons, buckles, shoelaces): Minimum 5 fastener types on fabric panels.

- Desktop presentation easel fold-stand: Sturdy foldable stand for tabletop presentation.

## 72.4 Installation Requirements

None. Tabletop or wall-mount workspace.

## 72.5 Documentation

Product Information, Warranty Certificate.

## 72.6 Training

Clinical use orientation.

## 72.7 Warranty

Minimum one (1) year against manufacturing defects.

# 73. Ring Tower & Ring Toss

Quantity: 1

## 73.1 Clinical & Functional Specifications

Vertical peg-post ring stacking and ring-toss game set for paediatric upper-extremity coordination training, hand-eye coordination, fine and gross motor skills development, and gamified therapeutic activity.

## 73.2 Engineering Technical Specifications

- Central Peg Tower: Vertical target peg post mounted on wide stable circular base plate; total height approximately 30—50 cm; non-toxic finish.

- Stacking Rings: Graduated colorful plastic doughnut rings (set of 5), graduated diameters for progressive stacking.

- Toss Rings: Woven sisal twine or soft rope throwing rings (set of 5), color-coded, safe for throwing.

- Base: Wide stable circular base plate, non-tip design.

- Applicable Standards: EN 71 (paediatric safety), non-toxic materials.

## 73.3 Standard Accessories

- Central vertical target peg post tower on wide base plate: Non-tip stable base.

- Graduated colorful plastic doughnut rings (set of 5): Graduated diameters, stacking.

- Woven sisal twine throwing rings (set of 5): Color-coded, safe throwing rings.

## 73.4 Installation Requirements

None. Requires level floor and tabletop or floor placement.

## 73.5 Documentation

Product Information, Safety Certificate, Warranty Certificate.

## 73.6 Training

Clinical use orientation.

## 73.7 Warranty

Minimum one (1) year against manufacturing defects.

# 74. Blocks, Cones & Musical Toys

Quantity: 3

## 74.1 Clinical & Functional Specifications

Assorted paediatric therapeutic toy set including geometric wooden blocks, stackable therapy cones, and handheld percussion musical instruments for paediatric fine and gross motor skill development, auditory processing, rhythm training, and sensory-motor integration.

## 74.2 Engineering Technical Specifications

- Wooden Blocks: Geometric colored blocks, minimum 30 pieces per set, non-toxic finish, multiple shapes and colors.

- Therapy Cones: Stackable plastic cones with smooth finger-grip finishes, minimum 10 cones per set, graduated colors.

- Percussion Instruments: Handheld therapy instruments including tambourine, maracas, and xylophone (or equivalent set), safe rounded edges, non-toxic finishes.

- Storage: Heavy-duty canvas equipment storage basket bin.

- Applicable Standards: EN 71 (paediatric safety), non-toxic materials, RoHS-compliant.

## 74.3 Standard Accessories

- Geometric colored wooden building blocks (set of 30): Multiple shapes and colors, non-toxic finish.

- Stackable plastic therapy cones with smooth grip finishes (set of 10): Graduated colors.

- Handheld percussion therapy instruments (tambourine, maracas, xylophone): Safe rounded edges.

- Heavy-duty canvas equipment storage basket bin: Consolidated storage.

## 74.4 Installation Requirements

None.

## 74.5 Documentation

Product Information, Safety Certificate, Warranty Certificate.

## 74.6 Training

Clinical use orientation.

## 74.7 Warranty

Minimum one (1) year against manufacturing defects.

# 75. Squeezing Ball for Hand Activities

Quantity: 2

## 75.1 Clinical & Functional Specifications

Set of flexible polyurethane foam squeezing balls with child-friendly graphics for paediatric hand strengthening, tactile stimulation, and playful therapeutic exercise. Multi-ball set permits varied resistance and grip challenges.

## 75.2 Engineering Technical Specifications

- Material: Flexible polyurethane foam, non-toxic, hypoallergenic, latex-free.

- Quantity: Minimum 6 balls per set, varied colors and graphics.

- Diameter: Approximately 6—8 cm each.

- Surface Graphics: High-visibility child-friendly graphic designs, non-toxic pigments.

- Storage: Net fabric travel organizer bag.

- Applicable Standards: EN 71 (paediatric safety), non-toxic pigments, phthalate-free, latex-free.

## 75.3 Standard Accessories

- Flexible polyurethane foam squeezing balls (set of 6): Non-toxic, hypoallergenic.

- High-visibility child-friendly design graphics: Non-toxic pigments.

- Net fabric travel organizing bag: Compact storage.

## 75.4 Installation Requirements

None.

## 75.5 Documentation

Product Information, Safety Certificate, Warranty Certificate.

## 75.6 Training

Clinical use orientation.

## 75.7 Warranty

Minimum one (1) year against manufacturing defects.

# 76. Needle Board (Threading Board)

Quantity: 1

## 76.1 Clinical & Functional Specifications

Perforated multi-hole matrix threading board for paediatric and adult fine-motor training, hand-eye coordination, bimanual coordination, and precision-grasp rehabilitation. Blunt plastic laces and colorful yarn threads shall be used for threading exercises without needle-injury risk.

## 76.2 Engineering Technical Specifications

- Board: Perforated high-density plastic multi-hole matrix, minimum 100 holes, hole diameter approximately 4—6 mm.

- Board Dimensions: Approximately 25 × 25 cm.

- Laces: Blunt plastic threading laces (minimum 5), thick colorful yarn threads (minimum 5 colors).

- Storage: Safety storage sorting tray drawer.

- Safety: Rounded edges, no sharp needles, blunt-tip laces only.

- Applicable Standards: EN 71 (paediatric safety), non-toxic materials.

## 76.3 Standard Accessories

- Perforated high-density plastic multi-hole matrix grid board: Rounded edges, safe hole diameters.

- Blunt plastic threading laces and colorful yarn threads: No sharp needles, minimum 5 each.

- Safety storage sorting tray drawer: Component containment.

## 76.4 Installation Requirements

None. Tabletop workspace.

## 76.5 Documentation

Product Information, Safety Certificate, Warranty Certificate.

## 76.6 Training

Clinical use orientation.

## 76.7 Warranty

Minimum one (1) year against manufacturing defects.

# 77. Beads and Cotton Balls

Quantity: 1

## 77.1 Clinical & Functional Specifications

Combined threading beads and cotton-ball manipulation set for paediatric fine-motor development, precision-grasp training, tool-use skills, and bilateral hand coordination in occupational therapy.

## 77.2 Engineering Technical Specifications

- Threading Beads: Large-diameter colorful perforated wooden beads, minimum 50 pieces per box, hole diameter suitable for supplied laces; non-toxic finishes.

- Cotton Balls: Sterile white medical cotton balls, minimum 100 pieces per pack.

- Sorting Bowls: Plastic sorting bowls (minimum 4), assorted colors, wipe-cleanable.

- Grasping Tools: Wooden chopsticks or grasping tools for tool-use training.

- Safety: Bead sizes chosen to minimize choking risk per applicable paediatric safety standards.

- Applicable Standards: EN 71 (paediatric safety), non-toxic finishes, RoHS-compliant.

## 77.3 Standard Accessories

- Large-diameter colorful perforated wooden threading beads (box of 50): Non-toxic finish.

- Sterile white medical cotton balls (100 pieces): Medical-grade cotton.

- Plastic sorting bowls (set of 4): Assorted colors, wipe-cleanable.

- Wooden grasping chopsticks / grasping tools: For tool-use training.

## 77.4 Installation Requirements

None. Tabletop workspace with supervision (paediatric use).

## 77.5 Documentation

Product Information, Safety Certificate, Warranty Certificate.

## 77.6 Training

Clinical use orientation.

## 77.7 Warranty

Minimum one (1) year against manufacturing defects.

# 78. Mat Set (Therapy Mats)

Quantity: 2

## 78.1 Clinical & Functional Specifications

Sectional folding high-density foam therapy mats for floor-based paediatric and adult rehabilitation exercises, neurodevelopmental therapy, tumbling and gross-motor activities, and floor-level therapeutic positioning. Interlocking edge design shall permit multi-mat assembly into extended floor coverage.

## 78.2 Engineering Technical Specifications

- Configuration: Sectional folding, multi-panel design.

- Standard Dimensions (deployed): Approximately 2 × 1 m per mat.

- Panel Thickness: Approximately 4—5 cm.

- Foam Density: High-density polyurethane or EVA foam, load-tested, shape-retentive.

- Cover: Heavy-duty wipe-clean reinforced vinyl, stitched or welded seams, fluid-proof antimicrobial finish.

- Handles: Integrated heavy-duty webbing carrying handles.

- Interlocking: Hook-and-loop edge strips for multi-mat locking alignment.

- User Weight Capacity: Minimum 150 kg.

- Applicable Standards: EN 71 where paediatric-applicable, antimicrobial cover material, RoHS-compliant, phthalate-free.

## 78.3 Standard Accessories

- Sectional folding high-density foam therapy mats (2 × 1 m): Multi-panel folding, shape-retentive.

- Heavy-duty wipe-clean reinforced vinyl exterior: Fluid-proof, antimicrobial, seam-sealed.

- Integrated heavy-duty webbing carrying handles: For transport.

- Hook-and-loop edge connection strips: For multi-mat locking alignment.

## 78.4 Installation Requirements

None. Storage floor space or vertical folding storage recommended.

## 78.5 Documentation

Product Information, Cleaning Instructions, Warranty Certificate.

## 78.6 Training

Clinical use orientation.

## 78.7 Warranty

Minimum two (2) years on foam core and cover workmanship.

# 79. Exercise Mats — Floor-Based (Yoga/Pilates Style)

Quantity: 2

## 79.1 Clinical & Functional Specifications

Non-slip closed-cell floor exercise mats for individual patient use during floor-based rehabilitation exercises, yoga, Pilates, stretching, and general therapeutic floor exercise. Roll-up design shall permit compact storage and transport.

## 79.2 Engineering Technical Specifications

- Material: Non-slip closed-cell thermoplastic elastomer (TPE) or equivalent; latex-free, phthalate-free.

- Thickness: Approximately 10 mm.

- Dimensions: Approximately 180 × 60 cm per mat.

- Surface: Texturized double-sided high-friction traction grid.

- Storage: Integrated elastic roll-up bundling strap with shoulder loop.

- Cleaning: Wipe-clean with mild soap or hospital disinfectant.

- Applicable Standards: RoHS-compliant, latex-free, phthalate-free.

## 79.3 Standard Accessories

- Non-slip closed-cell TPE cushioning floor yoga mats (10 mm thick): Latex-free, phthalate-free.

- Integrated elastic roll-up bundling straps with shoulder loop: For compact storage/transport.

- Texturized double-sided high-friction traction grid surface: Non-slip both faces.

## 79.4 Installation Requirements

None.

## 79.5 Documentation

Product Information, Cleaning Instructions, Warranty Certificate.

## 79.6 Training

None required.

## 79.7 Warranty

Minimum one (1) year against manufacturing defects.