HOSPITAL TECHNICAL TENDER SPECIFICATIONS — RESPIRATORY, ANAESTHESIA & AIRWAY EQUIPMENT

Technical Tender Document

Table of Contents

1. ICU-GRADE VENTILATOR (INVASIVE + NON-INVASIVE)	6

1.1 Clinical & Functional Specifications	6

1.2 Engineering Technical Specifications	6

1.3 Standard Accessories	8

1.4 Installation Requirements	10

1.5 Documentation	10

1.6 Training	10

1.7 Warranty	10

2. PORTABLE TRANSPORT VENTILATOR	11

2.1 Clinical & Functional Specifications	11

2.2 Engineering Technical Specifications	11

2.3 Standard Accessories	11

2.4 Installation Requirements	12

2.5 Documentation	12

2.6 Training	12

2.7 Warranty	13

3. ANAESTHESIA WORKSTATION	14

3.1 Clinical & Functional Specifications	14

3.2 Engineering Technical Specifications	14

3.3 Standard Accessories	15

3.4 Installation Requirements	16

3.5 Documentation	16

3.6 Training	16

3.7 Warranty	16

4. HIGH-FLOW NASAL CANNULA (HFNC) UNIT	17

4.1 Clinical & Functional Specifications	17

4.2 Engineering Technical Specifications	17

4.3 Standard Accessories	17

4.4 Installation Requirements	18

4.5 Documentation	18

4.6 Training	18

4.7 Warranty	18

5. CPAP UNIT	19

5.1 Clinical & Functional Specifications	19

5.2 Engineering Technical Specifications	19

5.3 Standard Accessories	19

5.4 Installation Requirements	20

5.5 Documentation	20

5.6 Training	20

5.7 Warranty	20

6. BiPAP UNIT	21

6.1 Clinical & Functional Specifications	21

6.2 Engineering Technical Specifications	21

6.3 Standard Accessories	21

6.4 Installation Requirements	22

6.5 Documentation	22

6.6 Training	22

6.7 Warranty	22

7. BAG-VALVE-MASK (BVM) RESUSCITATOR SET	23

7.1 Clinical & Functional Specifications	23

7.2 Engineering Technical Specifications	23

7.3 Standard Accessories	23

7.4 Installation Requirements	23

7.5 Documentation	24

7.6 Training	24

7.7 Warranty	24

8. OXYGEN THERAPY ACCESSORIES (VENTURI, NON-REBREATHER, NEBULIZER MASKS)	25

8.1 Clinical & Functional Specifications	25

8.2 Engineering Technical Specifications	25

8.3 Standard Accessories	25

8.4 Installation Requirements	26

8.5 Documentation	26

8.6 Training	26

8.7 Warranty	26

9. THERAPEUTIC NEBULIZER	27

9.1 Clinical & Functional Specifications	27

9.2 Engineering Technical Specifications	27

9.3 Standard Accessories	27

9.4 Installation Requirements	28

9.5 Documentation	28

9.6 Training	28

9.7 Warranty	28

10. DIFFICULT-AIRWAY MANAGEMENT TROLLEY	29

10.1 Clinical & Functional Specifications	29

10.2 Engineering Technical Specifications	29

10.3 Standard Accessories	29

10.4 Installation Requirements	30

10.5 Documentation	30

10.6 Training	30

10.7 Warranty	30

11. VIDEO LARYNGOSCOPE	31

11.1 Clinical & Functional Specifications	31

11.2 Engineering Technical Specifications	31

11.3 Standard Accessories	31

11.4 Installation Requirements	32

11.5 Documentation	32

11.6 Training	32

11.7 Warranty	32

12. FIBREOPTIC INTUBATION LARYNGOSCOPE / BRONCHOSCOPE	33

12.1 Clinical & Functional Specifications	33

12.2 Engineering Technical Specifications	33

12.3 Standard Accessories	33

12.4 Installation Requirements	34

12.5 Documentation	34

12.6 Training	34

12.7 Warranty	34

13. CONVENTIONAL LARYNGOSCOPE SET	35

13.1 Clinical & Functional Specifications	35

13.2 Engineering Technical Specifications	35

13.3 Standard Accessories	35

13.4 Installation Requirements	36

13.5 Documentation	36

13.6 Training	36

13.7 Warranty	36

14. PERIPHERAL NERVE STIMULATOR (REGIONAL ANAESTHESIA)	37

14.1 Clinical & Functional Specifications	37

14.2 Engineering Technical Specifications	37

14.3 Standard Accessories	37

14.4 Installation Requirements	38

14.5 Documentation	38

14.6 Training	38

14.7 Warranty	38

# 1. ICU-GRADE VENTILATOR (INVASIVE + NON-INVASIVE)

Quantity: 14

## 1.1 Clinical & Functional Specifications

The equipment shall be a microprocessor-controlled, electronically and pneumatically driven, critical-care ventilator suitable for continuous invasive and non-invasive ventilatory support of adult, pediatric, and (where configured) neonatal patients in intensive care, high-dependency, and step-down units. It shall support the complete range of conventional and advanced ventilation therapies, integrated monitoring of respiratory mechanics, lung-protective ventilation strategies, weaning protocols, and intra-hospital transport with battery operation.

The ventilator shall provide:

- Fully closed-loop electronic control of inspiratory and expiratory flow, pressure, tidal volume, and FiO₂

- Comprehensive alarm system with configurable priority levels

- Automatic patient-circuit compliance and leak compensation

- Automatic self-test and pre-operational check routines

- Trend memory and event log for a minimum of 72 hours

- Data connectivity to hospital patient monitoring, PDMS, and HIS systems

Ventilation Modes (minimum required): \ Volume Control (VCV), Pressure Control (PCV), Pressure Regulated Volume Control (PRVC) or equivalent dual-controlled mode, SIMV (Volume and Pressure), Pressure Support Ventilation (PSV), CPAP, Non-Invasive Ventilation (NIV) with leak compensation, NIV-PS, Apnea Backup Ventilation, Manual/Sigh breath, and an advanced spontaneous breathing / weaning mode (e.g. proportional or adaptive support ventilation or equivalent).

Monitored Parameters (minimum): \ Peak, Plateau, Mean and PEEP airway pressures; Inspired and Expired tidal volume; Minute volume (total, spontaneous, mandatory); Respiratory rate (total and spontaneous); I:E ratio and inspiratory/expiratory times; FiO₂; Static and dynamic compliance; Airway resistance; Intrinsic PEEP (auto-PEEP); Rapid Shallow Breathing Index (RSBI/f/Vt); Work of breathing (where applicable); P0.1; Leak percentage in NIV mode.

Displayed Waveforms and Loops: \ Real-time Pressure-Time, Flow-Time, Volume-Time waveforms and Pressure-Volume, Flow-Volume loops with reference-loop storage and comparison.

## 1.2 Engineering Technical Specifications

Display

- Integrated high-resolution color TFT/LED touchscreen, minimum 12-inch diagonal, or equivalent or better

- Minimum resolution 1024 × 768 pixels

- Tiltable and swivelable mounting; anti-glare surface

- Simultaneous display of numerical values, at least three waveforms, and at least two loops

User Interface

- Touchscreen with additional dedicated hard keys and/or rotary encoder for critical functions

- Multi-language selectable interface

Control & Processing

- Microprocessor-controlled with dual-processor architecture or equivalent safety redundancy

- Independent monitoring/alarm processor for safety verification

Gas Delivery System

- Dual gas source: high-pressure oxygen and medical air (2.7 — 6 bar pipeline pressure)

- Integrated electronic gas blender

- FiO₂ range: 21 — 100%, adjustable in ≤1% increments

- Accuracy: within ±3% of set value

Flow and Pressure Ranges (minimum)

- Tidal Volume: 20 — 2000 mL, adjustable

- Respiratory Rate: 1 — 100 breaths/min (adult/pediatric); down to neonatal range where configured

- Inspiratory Flow: up to 180 L/min or higher

- PEEP/CPAP: 0 — 45 cmH₂O

- Pressure Support: 0 — 60 cmH₂O

- Pressure Limit: adjustable up to 100 cmH₂O

- I:E Ratio: 1:9 to 4:1, selectable

- Inspiratory Trigger: Flow (0.5 — 15 L/min) and Pressure (−0.5 to −20 cmH₂O), selectable

- Rise Time / Inspiratory Slope: adjustable

Accuracy

- Tidal volume: within ±10% or ±10 mL, whichever is greater

- Pressure measurement: within ±2 cmH₂O or ±4% of reading

- Flow measurement: within ±5%

Alarms

- Audible and visual, minimum three priority levels (high/medium/low)

- User-adjustable limits for high/low airway pressure, high/low minute volume, high/low tidal volume, high/low respiratory rate, high/low FiO₂, apnea, disconnection, occlusion, low battery, gas supply failure

- Alarm silence with automatic re-activation

Power Supply

- 100 — 240 VAC, 50/60 Hz

- Integrated rechargeable internal battery, minimum 60 minutes of operation at typical settings on a full charge; hot-swappable or dual-battery configuration preferred

- Charging time: maximum 8 hours to full charge

Connectivity

- Minimum: RS-232/USB and Ethernet (RJ-45) port

- HL7 compatibility for HIS/PDMS integration

- Wi-Fi capability (optional/configurable)

- Nurse-call interface output

Environmental

- Operating temperature: 10 — 40 °C

- Storage temperature: −20 to +60 °C

- Relative humidity: 15 — 95%, non-condensing

Physical

- Compact, cart/trolley-mounted design

- Approximate weight: 20 — 60 kg with trolley (configuration-dependent)

- Ergonomic mobile trolley with lockable casters (included as accessory)

Safety & Standards Compliance

- IEC 60601-1 (General safety)

- IEC 60601-1-2 (EMC)

- IEC 60601-1-8 (Alarm systems)

- IEC 60601-2-12 (Particular requirements for lung ventilators for critical care)

- ISO 80601-2-12

- ISO 14971 (Risk management)

- IEC 62304 (Medical device software)

- CE MDR marking and/or US FDA 510(k) clearance

- Class II protection; Type BF applied part

## 1.3 Standard Accessories

Heavy-duty mobile trolley with lockable casters and steering handle

- Powder-coated steel or medical-grade aluminum construction

- Minimum four dual-wheel antistatic casters, at least two with independent locking brakes

- Ergonomic push handle; integrated cable management; utility basket

- Load capacity suitable to safely support ventilator and humidifier assembly

Articulated support arm for patient breathing circuit

- Multi-jointed articulating arm with adjustable friction/tension joints

- Ventilator/trolley-mount compatible

- Circuit-holder clamp for standard 22 mm circuit tubing

- Sterilizable/cleanable outer finish

Reusable adult silicone breathing circuit with water traps

- Medical-grade silicone, latex-free, DEHP-free

- Standard 22 mm connectors; length approximately 1.5 — 1.8 m

- Integrated water traps at both limbs

- Autoclavable at 134 °C for minimum 100 cycles

Reusable pediatric silicone breathing circuit with water traps

- Medical-grade silicone, latex-free, DEHP-free

- 15 mm/22 mm connectors, appropriate internal diameter for pediatric flows

- Integrated water traps

- Autoclavable

Reusable adult NIV silicone mask with headgear

- Medical-grade silicone cushion; transparent shell

- Full-face and/or oronasal configuration as required

- Anti-asphyxia valve; adjustable multi-point headgear

- Standard 22 mm connector; autoclavable/high-level disinfection compatible

Reusable pediatric NIV silicone mask with headgear

- Medical-grade silicone cushion, pediatric sizing

- Anti-asphyxia valve; adjustable multi-point headgear

- Compatible with standard 22 mm ventilator connectors

Integrated electronic or ultrasonic humidifier base unit with reusable chamber

- Servo-controlled heated humidifier

- Temperature range: 31 — 40 °C, adjustable

- Dual heated-wire circuit compatibility

- Reusable autoclavable water chamber with automatic or manual water feed

- Alarms for high/low temperature and empty chamber

- IEC 60601-2-74 compliant

High-pressure oxygen and medical air inlet hoses with color-coded country-specific gas connectors

- Reinforced, kink-resistant medical gas hoses, minimum 3 m length

- Color-coded per ISO 32 / country-specific standards

- NIST or DISS end-fittings compatible with hospital wall outlets (specify at contract stage)

- Pressure rating minimum 10 bar

Reusable flow sensor (proximal or internal, per manufacturer's design)

- Compatible with adult and pediatric flow ranges

- Autoclavable or high-level disinfectable per manufacturer's IFU

- Accuracy: within ±5%

Reusable expiratory valve assembly with integrated membrane

- Autoclavable expiratory valve assembly

- Silicone membrane, latex-free

- Tool-free removal and reassembly

Built-in rechargeable backup battery pack

- Sealed lithium-ion or equivalent

- Minimum 60 minutes runtime at typical adult settings

- Maximum charging time 8 hours

- Battery status indication on main display

## 1.4 Installation Requirements

- Electrical: 100 — 240 VAC, 50/60 Hz, dedicated grounded outlet; UPS-compatible

- Medical gases: Oxygen and Medical Air pipeline outlets, 2.7 — 6 bar, country-specific fittings

- Environmental: standard ICU conditions; ambient 10 — 40 °C

- Space: allowance for trolley footprint plus humidifier and circuit swing radius

- Network: RJ-45 drop for HIS/PDMS integration (where connectivity is used)

## 1.5 Documentation

Supplier shall provide User Manual, Service Manual, Parts Catalogue, Calibration Certificate, Factory Test Certificate, Electrical Safety Test Report (IEC 62353), Preventive Maintenance Schedule, Installation Report, and Commissioning Report — all in English, one hard copy and one soft copy per unit.

## 1.6 Training

Clinical training for ICU physicians, nurses, and respiratory therapists; biomedical engineering training for the hospital technical team covering operation, PPM, troubleshooting, and calibration; on-site application training; refresher session at 6 months post-commissioning.

## 1.7 Warranty

Minimum two (2) years comprehensive on-site warranty from date of commissioning, covering all labour, spare parts, consumable calibration items (excluding disposables), preventive maintenance (minimum twice yearly), breakdown maintenance, software updates, and technical support. Minimum 10-year availability of spare parts guaranteed in writing.

# 2. PORTABLE TRANSPORT VENTILATOR

Quantity: 5

## 2.1 Clinical & Functional Specifications

Compact, lightweight, ruggedized ventilator for intra-hospital and inter-hospital transport of adult and pediatric patients, suitable for use in ambulances, aircraft, and MRI-adjacent areas (where declared). The unit shall provide invasive and non-invasive ventilation with a full alarm suite and internal battery operation, suitable for emergency, ICU-transfer, and pre-hospital use.

Ventilation Modes: Volume Control, Pressure Control, SIMV, PSV, CPAP, NIV with leak compensation, Apnea Backup, and Manual/Assist breath.

Monitored Parameters: Peak/Plateau/Mean/PEEP airway pressures, Tidal Volume (inspired/expired), Minute Volume, Respiratory Rate, I:E ratio, FiO₂, Leak.

## 2.2 Engineering Technical Specifications

- Display: Sunlight-readable color TFT/LED, minimum 5-inch diagonal, resolution minimum 800 × 480 pixels

- Ventilation Ranges: Tidal Volume 20 — 2000 mL; Rate 1 — 80 bpm; PEEP 0 — 20 cmH₂O; Pressure Support 0 — 40 cmH₂O; Pressure limit adjustable to 80 cmH₂O

- FiO₂: 21 — 100% (with air-entrainment/turbine) or 40 — 100% (oxygen-driven), configurable

- Pneumatic Design: Internal turbine/blower preferred to allow operation without medical air supply; supplemental low-pressure and high-pressure oxygen inputs

- Accuracy: Volume ±10%, Pressure ±2 cmH₂O or ±4%

- Alarms: IEC 60601-1-8 compliant, multi-priority

- Battery: Internal rechargeable lithium-ion, minimum 4 hours runtime at typical settings; hot-swappable pack strongly preferred

- Charging Time: Maximum 6 hours

- Power: 100 — 240 VAC 50/60 Hz; 12 — 28 VDC vehicle input

- Weight: Approximately 3 — 8 kg with battery

- Ingress Protection: Minimum IP54

- Environmental: Operating −5 to +50 °C; humidity 15 — 95% non-condensing; altitude tolerance up to at least 3000 m

- Shock/Vibration: Certified to relevant EN 1789 / RTCA DO-160 or equivalent transport standards

- Connectivity: USB and/or RS-232; optional Bluetooth

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-8; IEC 60601-1-12 (for use in EMS environment); ISO 80601-2-12 and/or ISO 80601-2-84 (as applicable); EN 794-3 or equivalent; CE MDR / FDA 510(k).

## 2.3 Standard Accessories

Bedside rail mounting bracket and ambulance pole-clamp assembly

- Corrosion-resistant metal alloy construction; quick-release lever

- Compatible with standard hospital bed rails and 25 — 38 mm ambulance/EMS poles

Reusable adult transport silicone breathing circuit

- Medical-grade silicone, lightweight, 22 mm connectors, autoclavable

- Length appropriate for transport (approximately 1.2 — 1.5 m)

Reusable pediatric transport silicone breathing circuit

- Medical-grade silicone, 15/22 mm connectors, autoclavable

- Reduced internal volume suitable for pediatric flows

Rechargeable lithium-ion battery pack with integrated charging module

- Minimum 4 hours runtime; charge-state indicator; hot-swap capability preferred

- Compliant with IEC 62133

AC power supply adapter and 12V DC vehicle power cable

- Medical-grade AC adapter, 100 — 240 VAC 50/60 Hz

- 12 — 28 VDC vehicle cable with standard EMS plug and reverse-polarity/over-current protection

High-pressure oxygen input hose with quick-connect adapter

- Reinforced medical hose, minimum 2 m, 4 — 6 bar operating pressure

- Country-specific quick-connect

Lightweight oxygen cylinder regulator and pin-index adapter

- Pin-index safety system (PISS) compliant per ISO 407

- Adjustable outlet pressure, integrated pressure gauge, cylinder contents indicator

Rugged, weather-resistant padded carrying bag with shoulder strap

- Water- and dust-resistant outer shell

- Padded protective interior; access panels for controls/display; oxygen cylinder pocket; reinforced shoulder strap

## 2.4 Installation Requirements

No fixed installation; user training on mounting bracket use, vehicle power connection, and battery management.

## 2.5 Documentation

User Manual, Service Manual, Parts Catalogue, Factory Test Certificate, Electrical Safety Report, and PPM Schedule (English, hard and soft copies).

## 2.6 Training

Clinical training for transport teams (ICU, ER, anesthesia); technical training on battery care, cleaning, and first-line troubleshooting.

## 2.7 Warranty

Minimum two (2) years comprehensive warranty covering parts, labour, PPM, software, and technical support; minimum 10-year spare-parts availability.

# 3. ANAESTHESIA WORKSTATION

Quantity: 5

## 3.1 Clinical & Functional Specifications

Fully integrated anaesthesia workstation for adult and pediatric patients (and neonatal where configured) suitable for general anaesthesia in operating theatres. Shall integrate an advanced electronic anaesthesia ventilator, precision electronic or mechanical gas delivery, dual-agent vaporizer manifold, circle absorber system, integrated monitoring of respiratory mechanics and gases, and safety features including hypoxic guard, gas failure protection, and comprehensive alarms.

Ventilation Modes (minimum): Volume Control, Pressure Control, Pressure Regulated Volume Control or equivalent, SIMV (Vol & Pres), Pressure Support, Manual, and Spontaneous with PS backup. Suitable for tidal volumes down to at least 20 mL for pediatric/neonatal use.

Anaesthetic Gases: O₂, N₂O, Medical Air pipeline supplies with cylinder backup mounting; hypoxic guard limiting FiO₂ to a safe minimum (e.g. ≥25%).

## 3.2 Engineering Technical Specifications

- Display: Color TFT touchscreen, minimum 12-inch diagonal, resolution ≥1024 × 768

- Waveforms: Real-time P-T, F-T, V-T waveforms; PV and FV loops

- Ventilator Ranges: Tidal Volume 20 — 1500 mL; Rate 4 — 80 bpm (extended for pediatric); PEEP 0 — 20 cmH₂O; Pmax adjustable to 80 cmH₂O; I:E ratio 4:1 to 1:8

- Fresh Gas Flow: Electronic or mechanical flowmeters, 0 — 15 L/min per gas; low-flow anaesthesia (≤1 L/min) capable

- Vaporizer Manifold: Dual-position, interlock preventing simultaneous use of two agents; agent-specific keyed fill preferred

- Circle System: Autoclavable circle absorber with soda-lime canister, unidirectional valves, APL valve (0 — 70 cmH₂O adjustable)

- Safety Features: Hypoxic guard, O₂ failure alarm and N₂O cut-off, oxygen flush (35 — 75 L/min), pipeline and cylinder pressure gauges, anti-hypoxia mixing control

- Integrated Monitoring: Airway pressure, expired tidal volume, minute volume, respiratory rate, FiO₂ (galvanic or paramagnetic), and provision for agent gas (multi-gas) monitoring module — either integrated or via connectivity

- Alarms: IEC 60601-1-8 multi-priority

- Power: 100 — 240 VAC 50/60 Hz; internal battery backup minimum 60 minutes

- Storage: Integrated writing surface, drawer storage, and side rails

- Environmental: 10 — 40 °C operating; 15 — 95% humidity non-condensing

- Connectivity: RS-232/USB/Ethernet; HL7 compatible; AIMS integration capability

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-8; ISO 80601-2-13 (anaesthetic workstations); ISO 8835 series; ISO 5358; ISO 8359 (where applicable); CE MDR / FDA 510(k).

## 3.3 Standard Accessories

Integrated multi-mode anaesthesia ventilator with color display screen

- Electronically controlled ventilator as part of the workstation; specifications as above

Selectatec-compatible dual-position vaporizer mounting manifold

- Industry-standard dual-position interlocking manifold

- Mechanical interlock preventing simultaneous vapor delivery from two vaporizers

Isoflurane vaporizer (key-fill or pour-fill)

- Agent-specific temperature- and flow-compensated

- Concentration range: 0 — 5%, calibrated

- Accuracy: within ±10% or ±0.2 vol%

- Compliant with ISO 80601-2-13

Sevoflurane vaporizer (key-fill or pour-fill)

- Agent-specific temperature- and flow-compensated

- Concentration range: 0 — 8%, calibrated

- Accuracy: within ±10% or ±0.2 vol%

- Compliant with ISO 80601-2-13

Integrated ascending or descending breathing bag bellows assembly

- Transparent bellows housing with graduated volume scale

- Adult and pediatric bellows capacity or interchangeable

Reusable autoclavable circle absorber system with integrated soda lime canister

- Fully autoclavable circle system

- Transparent soda-lime canister with visible fill indicator; tool-free removal

- Unidirectional inspiratory/expiratory valves; APL valve

Reusable silicone patient breathing circuit with 2-litre reservoir bag

- Medical-grade silicone, latex-free, autoclavable

- 22 mm standard connectors; 2 L silicone reservoir bag

Anesthetic Gas Scavenging System (AGSS) active receiving assembly and transfer hose

- Active scavenging system with reservoir and flow indicator

- Compliant with ISO 80601-2-13 / ISO 8835-3

Pipeline inlet hoses for Oxygen, Nitrous Oxide, and Medical Air with NIST/DISS fittings

- Color-coded per ISO 32; NIST or DISS non-interchangeable end-fittings

- Length minimum 3 m; pressure rating 10 bar

Auxiliary oxygen flowmeter with dedicated outlet port

- Independent flowmeter, 0 — 10 L/min, with dedicated outlet for supplementary oxygen

Integrated backup emergency oxygen flush valve

- Non-latching push-button; delivery 35 — 75 L/min

- Direct connection to fresh-gas outlet, bypassing vaporizer

## 3.4 Installation Requirements

- Electrical: dedicated grounded 100 — 240 VAC outlet; UPS-compatible

- Medical gases: O₂, N₂O, and Medical Air pipeline terminals with country-specific fittings; vacuum outlet for AGSS

- Space: standard OR footprint with clearance for cable/hose routing; scavenging duct connection

- HVAC: standard OR HVAC with anaesthetic gas exchange rate

## 3.5 Documentation

User Manual, Service Manual, Parts Catalogue, Vaporizer Calibration Certificates, Factory Test Certificate, Electrical Safety Report (IEC 62353), PPM Schedule, Installation and Commissioning Reports.

## 3.6 Training

Clinical training for anaesthesiologists and OR nurses; biomedical training including PPM, vaporizer service intervals, soda-lime replacement, and basic troubleshooting; application refresher at 6 months.

## 3.7 Warranty

Minimum two (2) years comprehensive on-site warranty; twice-yearly PPM; software updates included; 10-year spares availability.

# 4. HIGH-FLOW NASAL CANNULA (HFNC) UNIT

Quantity: 13

## 4.1 Clinical & Functional Specifications

Standalone high-flow humidified oxygen therapy system for adult and pediatric patients requiring non-invasive respiratory support with heated, humidified, precisely blended gas at high flow rates. Suitable for hypoxemic respiratory failure, post-extubation support, and pre-oxygenation.

## 4.2 Engineering Technical Specifications

- Flow Range: Minimum 2 — 60 L/min (adult); pediatric mode down to 2 L/min

- FiO₂ Range: 21 — 100%, adjustable in ≤1% increments, accuracy within ±3%

- Temperature Settings: Selectable, typically 31 / 34 / 37 °C at patient outlet

- Humidification: Servo-controlled heated humidifier with reusable auto-fill chamber; heated-wire circuit

- Display: Color touchscreen, minimum 7-inch diagonal, resolution ≥800 × 480

- Alarms: Flow, FiO₂, temperature, water level, circuit disconnect, power failure — IEC 60601-1-8 compliant

- Gas Inputs: High-pressure oxygen (2.7 — 6 bar) with integrated electronic blender; internal turbine for air (allowing operation without wall air) preferred

- Power: 100 — 240 VAC, 50/60 Hz

- Battery Backup: Internal rechargeable, minimum 30 minutes

- Environmental: 10 — 40 °C operating; 15 — 95% humidity non-condensing

- Weight: Approximately 5 — 12 kg

- Connectivity: RS-232/USB minimum; optional Ethernet

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-8; IEC 60601-2-74 (humidifiers); ISO 80601-2-90 (respiratory high-flow therapy) or equivalent; CE MDR / FDA.

## 4.3 Standard Accessories

Mobile pole stand with lockable casters and IV pole adapter

- Powder-coated steel or aluminum, minimum four antistatic casters (two locking)

- Adjustable height; universal IV pole adapter

Heated patient breathing tube with integrated heating wire

- Medical-grade, low-compliance tubing with dual-limb or single-limb heated wire

- Standard 22 mm connectors; single-patient or reusable per manufacturer's IFU

Reusable automatic-filling humidification chamber

- Autoclavable or high-level disinfectable per IFU; automatic water feed

- Compatible with sterile water supply system

Reusable adult high-flow nasal cannula — Medium

- Medical-grade silicone, latex-free; sized for adult medium

- Soft, non-occlusive prongs; adjustable head strap

Reusable adult high-flow nasal cannula — Large

- Medical-grade silicone, sized for adult large

- Soft prongs; adjustable head strap

Oxygen inlet connecting hose with flowmeter adapter

- Reinforced, kink-resistant, minimum 3 m, color-coded, country-specific fitting

Integrated electronic blender system for FiO₂ control

- Precision electronic blender; FiO₂ 21 — 100%, ±3% accuracy

Built-in rechargeable backup battery pack

- Minimum 30 minutes; charge indicator on display

## 4.4 Installation Requirements

Standard grounded outlet 100 — 240 VAC; medical oxygen pipeline outlet; ward-level ambient conditions.

## 4.5 Documentation

User Manual, Service Manual, Parts List, Factory Test Certificate, Electrical Safety Report, PPM Schedule.

## 4.6 Training

Clinical training for ICU/ward nursing and respiratory therapy staff; technical training on humidifier care, filter changes, and troubleshooting.

## 4.7 Warranty

Minimum two (2) years comprehensive on-site warranty; 10-year spares availability.

# 5. CPAP UNIT

Quantity: 2

## 5.1 Clinical & Functional Specifications

Continuous Positive Airway Pressure device for the treatment of obstructive sleep apnea and related conditions in adult patients. Provides fixed or auto-titrating positive pressure with integrated humidification, data logging, and compliance reporting.

## 5.2 Engineering Technical Specifications

- Pressure Range: 4 — 20 cmH₂O, adjustable in ≤0.5 cmH₂O increments; accuracy ±1 cmH₂O

- Modes: Fixed CPAP and Auto-CPAP (APAP)

- Ramp Function: 0 — 45 minutes adjustable

- Expiratory Pressure Relief: Selectable feature

- Humidification: Integrated heated humidifier, adjustable levels

- Display: Color LCD, minimum 2.5-inch, backlit

- Data Logging: Minimum 12 months compliance data; therapy events, leak, AHI/AI indices

- Connectivity: USB and/or SD card; Bluetooth/Wi-Fi optional for cloud/PC reporting

- Sound Level: ≤30 dBA at typical pressure

- Power: 100 — 240 VAC 50/60 Hz; 12/24 VDC compatible (via adapter)

- Weight: Approximately 1 — 2 kg (main unit)

- Environmental: 5 — 35 °C operating

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-11 (home healthcare); ISO 80601-2-70 (sleep apnoea therapy); ISO 17510; CE MDR / FDA.

## 5.3 Standard Accessories

Integrated heated humidifier unit with water chamber

- Adjustable heating; dishwasher-safe or manually-cleanable removable chamber

Flexible heated breathing tube

- Medical-grade, standard 22 mm; heated wire; length approximately 1.8 m

Reusable silicone nasal mask with adjustable headgear

- Medical-grade silicone cushion; multiple sizes; adjustable four-point headgear

Reusable silicone full-face mask with adjustable headgear

- Medical-grade silicone, full-face oronasal configuration; anti-asphyxia valve; adjustable headgear

AC to DC power supply adapter with main power cord

- Medical-grade, 100 — 240 VAC input; country-appropriate mains cord

Cleanable inlet air filters (assembled in unit)

- Reusable foam/washable filter plus disposable fine-particle filter option

Lightweight padded travel carrying bag

- Padded interior with compartments for unit, humidifier chamber, tube, mask, and adapter

## 5.4 Installation Requirements

Standard grounded electrical outlet; bedside table space with distilled-water access for humidifier.

## 5.5 Documentation

User Manual, Service Manual (as applicable), Factory Test Certificate, Electrical Safety Report, and Care/Cleaning Guide.

## 5.6 Training

Clinical training for prescribing/pulmonology staff; patient education material and setup training; technical training for biomedical staff.

## 5.7 Warranty

Minimum two (2) years comprehensive warranty covering unit and accessories (except consumables); software updates included.

# 6. BiPAP UNIT

Quantity: 2

## 6.1 Clinical & Functional Specifications

Bi-level positive airway pressure device providing distinct inspiratory and expiratory pressures for treatment of complex sleep apnea, hypoventilation syndromes, COPD, and NIV support. Adult use; pediatric optional.

## 6.2 Engineering Technical Specifications

- Modes: CPAP, Bi-level Spontaneous (S), Spontaneous/Timed (S/T), Timed (T), and Auto Bi-level

- IPAP Range: 4 — 30 cmH₂O

- EPAP Range: 4 — 25 cmH₂O

- Pressure Support: Adjustable up to 25 cmH₂O

- Back-up Rate: 4 — 40 bpm

- Rise Time / Ti min / Ti max: Adjustable

- Accuracy: ±1 cmH₂O

- Ramp: 0 — 45 minutes

- Humidification: Integrated heated humidifier with adjustable levels

- Leak Compensation: Automatic

- Display: Color LCD minimum 2.5-inch, backlit

- Data Storage: Minimum 12 months; AHI, leak, adherence

- Connectivity: USB/SD card; Bluetooth/Wi-Fi optional

- Sound Level: ≤30 dBA

- Power: 100 — 240 VAC 50/60 Hz; DC adapter compatibility

- Weight: Approximately 1.5 — 3 kg

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-11; ISO 80601-2-72 or ISO 80601-2-70 as applicable; CE MDR / FDA.

## 6.3 Standard Accessories

Integrated heated humidifier with removable water tub

- Adjustable heat, cleanable/dishwasher-safe removable chamber

Flexible heated patient delivery tube

- Medical-grade, 22 mm, heated wire, approximately 1.8 m

Reusable silicone full-face mask with headgear strap

- Medical-grade silicone; anti-asphyxia valve; adjustable headgear

Reusable silicone nasal mask with headgear strap

- Medical-grade silicone; adjustable headgear

AC power adapter with main electrical cable

- Medical-grade, 100 — 240 VAC, country-appropriate cord

Fine particle inlet air filters (assembled in unit)

- Combined coarse reusable and fine disposable filter set

Rigid transit storage bag

- Rigid or semi-rigid protective case with fitted compartments

## 6.4 Installation Requirements

Grounded electrical outlet; bedside space; distilled water for humidifier.

## 6.5 Documentation

User Manual, Service Manual, Factory Test Certificate, Electrical Safety Report, PPM Schedule.

## 6.6 Training

Clinical training for pulmonology/ICU staff; patient education; biomedical training on filter care and troubleshooting.

## 6.7 Warranty

Minimum two (2) years comprehensive on-site warranty; software updates included.

# 7. BAG-VALVE-MASK (BVM) RESUSCITATOR SET

Quantity: 10

## 7.1 Clinical & Functional Specifications

Manual self-inflating resuscitator for emergency ventilation of adult patients in cardiac/respiratory arrest, transport, and pre-oxygenation. Complete reusable set with pressure limiting, oxygen enrichment via reservoir, and standard connectors.

## 7.2 Engineering Technical Specifications

- Bag Volume (Adult): 1500 mL (approximate)

- Delivered Tidal Volume: ≥600 mL single-hand, ≥1000 mL two-hand

- Material: Medical-grade silicone, latex-free, DEHP-free; fully autoclavable at 134 °C

- Pressure Limiting Valve: 40 ± 5 cmH₂O with over-ride option

- Connectors: Standard 22 mm/15 mm ISO connectors

- Oxygen Inlet: Nipple connector for oxygen tubing; reservoir attachment

- Oxygen Delivery with Reservoir at 15 L/min: ≥85% FiO₂

- Operating Temperature: −20 to +50 °C

- Reusable Cycles: Minimum 100 autoclave cycles

Standards: ISO 10651-4 (manual resuscitators); ISO 5356-1 (conical connectors); CE MDR / FDA.

## 7.3 Standard Accessories

Reusable adult silicone resuscitator bag (1500 mL) with pressure limiting valve

- Fully autoclavable silicone bag; visible transparent or translucent construction preferred

- Integrated 40 cmH₂O pressure limiting valve with over-ride

Reusable silicone face mask size 4 (Adult Medium)

- Transparent silicone dome with cushioned rim; ISO 22 mm connector; autoclavable

Reusable silicone face mask size 5 (Adult Large)

- Transparent silicone dome, cushioned rim; ISO 22 mm connector; autoclavable

Oxygen reservoir bag with inlet valve assembly

- Minimum 2500 mL reservoir capacity; latex-free; inlet non-rebreathing valve

Oxygen star-lumen connecting tube (2-metre length)

- Anti-crush star-lumen construction; universal connectors; medical-grade PVC or silicone

Rigid protective clear plastic storage case

- Transparent, impact-resistant, wipeable case with foam insert

## 7.4 Installation Requirements

No installation; storage at defined emergency locations.

## 7.5 Documentation

User Manual, Cleaning/Sterilization Instructions, Factory Test Certificate, Certificate of Conformity (ISO 10651-4).

## 7.6 Training

Clinical training on adult BVM technique, oxygen enrichment, and cleaning/sterilization procedures.

## 7.7 Warranty

Minimum two (2) years against manufacturing defects; consumable elastomeric parts (valves, membranes) covered per IFU cycle life.

# 8. OXYGEN THERAPY ACCESSORIES (VENTURI, NON-REBREATHER, NEBULIZER MASKS)

Quantity: TBC

## 8.1 Clinical & Functional Specifications

Complete oxygen therapy accessory set for controlled-concentration oxygen delivery (Venturi principle), high-concentration delivery (non-rebreather with reservoir), and aerosolized medication delivery (nebulizer mask) to adult and pediatric patients.

## 8.2 Engineering Technical Specifications

- Materials: Medical-grade PVC or silicone, latex-free, DEHP-free where possible

- Connectors: Standard universal oxygen tubing connectors

- Venturi Kit: Color-coded dilution jets covering FiO₂ 24%, 28%, 31%, 35%, 40%, 50% (per ISO 407 color coding where applicable); accuracy of delivered FiO₂ within ±2% at recommended flow

- Non-Rebreather Mask: Integrated one-way valve preventing exhaled gas re-entry into reservoir; reservoir bag minimum 600 mL; delivers ≥85% FiO₂ at 10 — 15 L/min

- Nebulizer Mask: Compatible with standard nebulizer chamber; provides aerosol delivery with adequate entrainment

- Sizing: Adult and pediatric sizes as ordered

- Sterility: Single-use, individually packaged, ethylene-oxide or gamma sterilized as applicable

Standards: ISO 8185; ISO 5367; ISO 407; ISO 10993 (biocompatibility); CE MDR.

## 8.3 Standard Accessories

Variable-select Venturi kit including dilution color-coded valves (24% to 50%)

- Full color-coded jet set covering FiO₂ 24 — 50%

- Compatible mask with integrated jet housing

Non-rebreather oxygen mask assembly with check-valves and reservoir bag

- One-way valves on mask side ports and between mask and reservoir

- Reservoir bag minimum 600 mL; kink-resistant tubing minimum 2 m

- Adult and pediatric sizes

High-output therapeutic nebulizer mask assembly with medication chamber

- Aerosol mask compatible with standard nebulizer cup

- Medication chamber capacity minimum 6 mL; particle size within therapeutic range (MMAD 1 — 5 µm)

Anti-crush star-lumen oxygen tubing with universal connectors

- Star-lumen construction; medical-grade PVC; length minimum 2 m; universal connectors

Adjustable elastic head straps and moldable aluminum nose clips per mask unit

- Latex-free elastic strap; adjustable moldable metal nose bridge

## 8.4 Installation Requirements

Bedside oxygen source; standard flowmeter with metered outlet.

## 8.5 Documentation

Certificate of Conformity; Instructions for Use; Biocompatibility documentation per ISO 10993.

## 8.6 Training

Basic clinical orientation on selection of appropriate device per prescribed FiO₂ and flow.

## 8.7 Warranty

Minimum two (2) years against manufacturing defects for reusable components; disposables per single-use designation.

# 9. THERAPEUTIC NEBULIZER

Quantity: 17

## 9.1 Clinical & Functional Specifications

Compressor-driven jet nebulizer for aerosolized medication delivery to adult and pediatric patients in wards, outpatient clinics, and emergency use. Delivers therapeutic aerosol particles in the respirable range.

## 9.2 Engineering Technical Specifications

- Compressor Type: Oil-free piston or diaphragm

- Maximum Pressure: ≥30 psi (≥2 bar)

- Operating Pressure: 8 — 16 psi at nebulizer

- Flow Rate at Nebulizer: ≥6 L/min

- Nebulization Rate: ≥0.15 mL/min

- Particle Size (MMAD): 1 — 5 µm respirable range; ≥70% particles \<5 µm

- Medication Cup Capacity: Minimum 6 mL

- Sound Level: ≤65 dBA at 1 m

- Duty Cycle: Continuous or minimum 30 minutes ON per 60 minutes

- Power: 100 — 240 VAC 50/60 Hz

- Weight: Approximately 1.5 — 3 kg

- Environmental: 10 — 40 °C operating; 15 — 95% humidity non-condensing

- Filter: Replaceable inlet air filter

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-1-11 (home use where applicable); ISO 27427 (aerosol therapy equipment); CE MDR / FDA.

## 9.3 Standard Accessories

Heavy-duty oil-free compressor base unit

- Oil-free piston/diaphragm; thermal overload protection; carrying handle

- Rubber feet; ventilated housing

Reusable nebulizer cup / medication chamber kit

- Medical-grade polymer, autoclavable or top-rack dishwasher-safe per IFU

- Minimum 6 mL medication capacity

Reusable adult aerosol face mask

- Medical-grade PVC or silicone, latex-free; adjustable elastic strap and nose clip

Reusable pediatric aerosol face mask

- Medical-grade PVC or silicone, pediatric size; latex-free; adjustable strap

Angled mouthpiece attachment

- Rigid polymer mouthpiece; angled ergonomic design; latex-free

Flexible air connecting tube (2-metre length)

- Kink-resistant PVC or silicone; universal connectors

Replacement inlet air filter set (5 pieces)

- Manufacturer-specified filter media; visual change indicator preferred

## 9.4 Installation Requirements

Standard grounded electrical outlet; ward or outpatient bedside space.

## 9.5 Documentation

User Manual, Cleaning Instructions, Factory Test Certificate, Electrical Safety Report.

## 9.6 Training

Clinical training on nebulizer therapy technique; cleaning/disinfection SOP for reusable parts; biomedical filter-change protocol.

## 9.7 Warranty

Minimum two (2) years comprehensive warranty for compressor and accessories (except disposables); minimum 5-year spare filter availability.

# 10. DIFFICULT-AIRWAY MANAGEMENT TROLLEY

Quantity: 1

## 10.1 Clinical & Functional Specifications

Dedicated mobile crash-cart—style trolley for organized storage and rapid deployment of difficult-airway equipment (video/fiberoptic scopes, blades, LMAs, ETTs, cricothyroidotomy kits, drugs) per international difficult-airway algorithms. For OR, ICU, ER, and rapid-response use.

## 10.2 Engineering Technical Specifications

- Construction: Stainless steel (AISI 304 grade or better) or high-impact ABS/polymer, seamless welded/moulded body, corrosion-resistant, wipeable finish

- Drawers: Minimum 4 — 6 drawers with silent ball-bearing telescopic slides; central locking system with key

- Drawer Configuration: Color-coded modular dividers/organizers per airway-algorithm stage

- Top Surface: Anti-slip, wipe-clean, with raised edges to prevent items from sliding off; transparent protective mat

- Handles: Ergonomic push handles on at least two sides

- Casters: Minimum four antistatic conductive dual-wheel casters (minimum 125 mm diameter), at least two with brakes and one directional (fifth-wheel) if possible

- Load Capacity: Minimum 100 kg static, 60 kg dynamic

- Accessories Mounting: Side rails for IV pole, oxygen cylinder holder, sharps container, waste bin, defibrillator platform (as configured)

- Dimensions: Compact footprint suitable for corridor and door passage (approximately 600 — 900 mm width)

- Weight: Approximately 60 — 90 kg unloaded

Standards: Applicable ISO/EN standards for medical furniture and mobile equipment; anti-static compliance where OR/anaesthetic gas exposure applies; CE / FDA.

## 10.3 Standard Accessories

Stainless steel or ABS plastic mobile cart with multi-drawer lockable configuration

- As specified above; central key lock covering all drawers

Top surface transparent protective mat and integrated push handles

- Silicone or PVC transparent mat, wipeable; ergonomic push handles

Side-mounted sharps container bracket and biohazard waste bin

- Universal bracket accepting standard sharps container; hinged-lid waste bin

Heavy-duty IV pole attachment

- Telescoping stainless-steel IV pole, minimum two hooks, height-adjustable, load rating suitable for infusion pumps

Integrated oxygen cylinder holder bracket

- Bracket accepting standard portable oxygen cylinder (D or E size), secure retention strap

Internal modular drawer organizer dividers

- Adjustable, removable, autoclavable or wipeable dividers to configure per DAM algorithm

## 10.4 Installation Requirements

No installation; assembly per manufacturer instruction; storage at designated OR/ER/ICU location.

## 10.5 Documentation

User/Assembly Manual, Certificate of Conformity, Cleaning/Disinfection Instructions.

## 10.6 Training

Orientation on trolley layout aligned with DAM algorithm; cleaning/disinfection SOP.

## 10.7 Warranty

Minimum two (2) years against manufacturing and structural defects; caster and lock mechanism warranty as per IFU cycle life.

# 11. VIDEO LARYNGOSCOPE

Quantity: 6

## 11.1 Clinical & Functional Specifications

Portable video laryngoscope for direct and indirect visualization of the glottis during endotracheal intubation in adult and pediatric patients. Suitable for routine and difficult airway management in OR, ICU, ER, and pre-hospital settings.

## 11.2 Engineering Technical Specifications

- Display: Integrated color LCD/TFT, minimum 3-inch diagonal, resolution ≥800 × 480; anti-glare, tiltable/rotatable preferred

- Camera: High-resolution CMOS image sensor; minimum resolution 640 × 480; anti-fog technology

- Illumination: High-intensity LED, minimum 20,000 lux at blade tip; color temperature approximating daylight

- Field of View: Minimum 60°

- Recording: Onboard video/still image recording with internal memory or removable card (minimum 8 GB) and USB export

- Battery: Rechargeable lithium-ion; minimum 60 minutes continuous use per full charge; charging time maximum 4 hours

- Reprocessing: Handle/display unit wipeable; blades autoclavable or high-level disinfectable per IFU

- Ingress Protection: Minimum IPX4 on handle unit

- Weight: Approximately 300 — 700 g (handle with blade)

- Environmental: 10 — 40 °C operating

Standards: IEC 60601-1; IEC 60601-1-2; ISO 7376 (laryngoscopes); IEC 62366 (usability); CE MDR / FDA.

## 11.3 Standard Accessories

Reusable video laryngoscope core handle display unit with high-resolution LCD

- As specified above; ergonomic handle with intuitive controls

Reusable Macintosh blade Size 3 with integrated camera and LED

- Autoclavable or high-level disinfectable per IFU

- Integrated distal camera and LED illumination

- Anti-fog design

Reusable Macintosh blade Size 4 with integrated camera and LED

- As above, adult large size

Rigid stylet optimized for video intubation tube tracking

- Malleable/rigid stylet contoured to match blade curvature

- Medical-grade stainless steel or polymer; autoclavable

Desktop charging cradle with AC power adapter unit

- Charging dock; medical-grade adapter 100 — 240 VAC 50/60 Hz; charge status indicator

Waterproof protective storage and disinfection case

- Rigid, waterproof case with foam insert; suitable for transport and controlled disinfection storage

## 11.4 Installation Requirements

Grounded electrical outlet for charging cradle; storage on airway trolley or ready-access shelf.

## 11.5 Documentation

User Manual, Service Manual, Reprocessing/Disinfection Guide, Factory Test Certificate, Electrical Safety Report.

## 11.6 Training

Clinical training for intubators on device use and image interpretation; biomedical training on cleaning, disinfection, and battery care.

## 11.7 Warranty

Minimum two (2) years comprehensive on-site warranty for handle/display; blade warranty per IFU cycle life; minimum 5-year spares availability.

# 12. FIBREOPTIC INTUBATION LARYNGOSCOPE / BRONCHOSCOPE

Quantity: 1

## 12.1 Clinical & Functional Specifications

Flexible fiberoptic laryngoscope/bronchoscope for awake and asleep intubation, difficult-airway management, and diagnostic/therapeutic bronchoscopy in adult (and pediatric where sized) patients. Provides high-resolution visualization with steerable tip.

## 12.2 Engineering Technical Specifications

- Insertion Tube Outer Diameter: Adult intubation scope approximately 3.8 — 5.5 mm; smaller diameters for pediatric configurations

- Working Length: Minimum 600 mm

- Working Channel Diameter: Minimum 1.2 mm (adult intubation configuration); larger where bronchoscopy is intended

- Tip Angulation: Minimum 180° up / 130° down bidirectional

- Field of View: Minimum 85°

- Depth of Field: Approximately 3 — 50 mm

- Illumination: LED light source (handle-integrated battery or desktop console)

- Image Type: Fiberoptic bundle or CMOS chip-on-tip (video), per configuration

- Reprocessing: Fully immersible; compatible with standard high-level chemical disinfection and/or automated endoscope reprocessors per IFU

- Leak Testing: Watertight body with dedicated leak-test port

- Weight: Approximately 300 — 700 g

- Battery (portable style): Minimum 60 minutes continuous illumination

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-2-18 (endoscopic equipment); ISO 8600 series; ISO 15883 (reprocessing); CE MDR / FDA.

## 12.3 Standard Accessories

Flexible fiberoptic insertion tube with steerable bidirectional tip controls

- As specified above; ergonomic thumb-lever tip control

LED light source unit (integrated handle battery-operated or desktop console)

- High-intensity LED; adjustable brightness; low heat generation

- Console version: 100 — 240 VAC 50/60 Hz

Detachable battery charger unit with rechargeable lithium batteries (for handle style)

- Rechargeable lithium-ion pack; charger with 100 — 240 VAC adapter; charge indicator

Light guide cable adapter

- Universal adapter compatible with common light-source connections

Suction valve cleaning brush and irrigation adapters

- Cleaning brushes sized to working channel; standard luer/irrigation adapters for reprocessing

ETT fixation/intubation bite blocks

- Medical-grade polymer bite blocks; adult sizes

Watertight leak testing kit

- Pressure-source with gauge; complete leak-test setup per manufacturer IFU

Padded rigid transport case

- Rigid, foam-lined transport and storage case with contoured insert protecting insertion tube

## 12.4 Installation Requirements

Standard grounded electrical outlet for charger/console; endoscope reprocessing area with high-level disinfectant supply and drying capability.

## 12.5 Documentation

User Manual, Service Manual, Reprocessing Guide, Factory Test Certificate, Electrical Safety Report, Leak-Test Procedure.

## 12.6 Training

Clinical training on fiberoptic intubation and (if applicable) bronchoscopy; comprehensive reprocessing training per ISO 15883; biomedical training on leak testing and preventive care.

## 12.7 Warranty

Minimum two (2) years comprehensive warranty covering scope, light source, and battery; fiber-bundle warranty per manufacturer's standard; minimum 5-year spares availability.

# 13. CONVENTIONAL LARYNGOSCOPE SET

Quantity: 12

## 13.1 Clinical & Functional Specifications

Complete reusable set of conventional direct laryngoscopes with interchangeable curved (Macintosh) and straight (Miller) blades for adult and pediatric endotracheal intubation. Robust, all-metal construction for continuous OR, ICU, ER, and ward use.

## 13.2 Engineering Technical Specifications

- Construction: Full stainless steel or equivalent corrosion-resistant medical alloy, fully autoclavable at 134 °C

- Blade Fitment: ISO 7376 "Green Standard" fitting between handle and blade — universal cross-brand compatibility

- Illumination: LED light source at fiber-optic bundle or bulb-in-blade configuration; minimum 2000 lux at blade tip; color temperature approximating daylight

- Handle Types: Adult (standard) and pediatric (stubby) sizes

- Power Source: Standard C or AA alkaline / rechargeable battery within handle; battery life minimum 2 hours continuous illumination

- Blade Cycle Life: Minimum 4000 autoclave cycles

- Standards: ISO 7376; IEC 60601-1 where powered; CE MDR / FDA where applicable

## 13.3 Standard Accessories

Heavy-duty stainless steel battery handle (Standard/Adult size)

- Full stainless steel construction; knurled non-slip grip

- ISO 7376 green-standard fitting; autoclavable

Heavy-duty stainless steel battery handle (Stubby/Pediatric size)

- Compact stubby handle for pediatric/obstetric use

- Full stainless steel; autoclavable; ISO 7376 fitting

Curved Macintosh blade Size 1

- Neonatal/infant curved blade; stainless steel; ISO 7376; autoclavable

Curved Macintosh blade Size 2

- Pediatric curved blade; as above

Curved Macintosh blade Size 3

- Adult standard curved blade; as above

Curved Macintosh blade Size 4

- Adult large curved blade; as above

Straight Miller blade Size 0

- Neonatal straight blade; stainless steel; ISO 7376; autoclavable

Straight Miller blade Size 1

- Infant straight blade; as above

Spare xenon or LED replacement lamps (2 pieces)

- Manufacturer-specified spare bulbs/LEDs matched to handle/blade type

Rigid protective storage case with fitted foam slots

- Rigid, wipeable case with contoured foam holding handles, all blade sizes, and spare bulbs

## 13.4 Installation Requirements

No installation; storage on airway/anaesthesia trolley; batteries as per handle specification.

## 13.5 Documentation

User Manual, Cleaning/Sterilization Instructions, Certificate of Conformity (ISO 7376), Factory Test Certificate.

## 13.6 Training

Clinical training on selection and use of blades; cleaning and autoclave protocol; bulb/LED replacement.

## 13.7 Warranty

Minimum two (2) years against manufacturing defects for handles and blades; bulb/LED per IFU life.

# 14. PERIPHERAL NERVE STIMULATOR (REGIONAL ANAESTHESIA)

Quantity: 1

## 14.1 Clinical & Functional Specifications

Portable peripheral nerve stimulator for nerve localization during regional anaesthesia (peripheral nerve blocks) and for monitoring depth of neuromuscular blockade during general anaesthesia. Delivers precise constant-current impulses via insulated stimulating needles or surface electrodes.

## 14.2 Engineering Technical Specifications

- Output Current Range: 0 — 5 mA, adjustable in ≤0.01 mA increments at low ranges

- Current Accuracy: Within ±5% of set value

- Pulse Width: Selectable, typically 0.1 / 0.3 / 1.0 ms

- Stimulation Frequency: Selectable 1 Hz, 2 Hz; Train-of-Four (TOF) at 2 Hz for 4 impulses; Tetanic (50 Hz); Post-Tetanic Count (PTC); Double-Burst Stimulation (DBS) — modes for neuromuscular monitoring configuration

- Constant Current Output: Maintained against load impedance up to at least 10 kΩ

- Display: Digital LCD/OLED, backlit; simultaneous display of set and delivered current

- Alarms/Indicators: Open-circuit (disconnect) indication, low battery, high impedance

- Power: Standard 9V alkaline battery or rechargeable equivalent; battery life minimum 40 hours typical use

- Weight: Approximately 200 — 400 g

- Ingress Protection: Minimum IPX4

- Environmental: 10 — 40 °C operating

Standards: IEC 60601-1; IEC 60601-1-2; IEC 60601-2-10 (nerve and muscle stimulators); CE MDR / FDA.

## 14.3 Standard Accessories

Main pocket-sized control unit with digital parameter display

- Compact ergonomic housing; single-hand operation; digital display of current, mode, and status

Patient extension cable with crocodile/alligator clip connectors

- Insulated medical-grade cable, minimum 1.5 m; color-coded (black/red) crocodile clips

- Compatible with insulated stimulating needles (standard connector)

Surface electrode interface adapter

- Adapter cable accepting standard ECG-type surface electrodes for TOF/neuromuscular monitoring

Removable protective rubber boot enclosure

- Silicone or medical-grade rubber protective sleeve; splash-resistant; wipeable

Rechargeable or standard alkaline 9V battery set

- Standard 9V PP3 alkaline batteries and/or rechargeable equivalents with charger

Rigid transport pouch

- Padded protective pouch/case with belt loop and internal compartments for unit, cables, and spare batteries

## 14.4 Installation Requirements

No fixed installation; belt-carried or trolley-stored.

## 14.5 Documentation

User Manual, Service Manual, Certificate of Conformity, Factory Test Certificate, Electrical Safety Report.

## 14.6 Training

Clinical training for anaesthesiologists on peripheral nerve block localization and neuromuscular monitoring interpretation; biomedical care and battery management.

## 14.7 Warranty

Minimum two (2) years comprehensive warranty; cable/consumable clips per IFU cycle life; minimum 5-year spares availability