HOSPITAL EQUIPMENT — TECHNICAL TENDER SPECIFICATIONS

Surgical Support Equipment Package

Table of Contents

1. OPERATING TABLE — ELECTRO-HYDRAULIC, MULTI-POSITION	5

1.1  Clinical & Functional Specifications	5

1.2  Engineering Technical Specifications	5

1.3  Standard Accessories	6

1.4  Installation Requirements	6

1.5  Documentation	6

1.6  Training	7

1.7  Warranty	7

2. ELECTRO-HYDRAULIC OT TABLE — CARDIAC DEDICATED	8

2.1  Clinical & Functional Specifications	8

2.2  Engineering Technical Specifications	8

2.3  Standard Accessories	8

2.4  Installation Requirements	9

3. FULLY HYDRAULIC OT TABLE — UROLOGY / GENERAL SURGERY	10

3.1  Clinical & Functional Specifications	10

3.2  Engineering Technical Specifications	10

3.3  Standard Accessories	10

3.4  Installation Requirements	11

4. OT LIGHTS — LED, CEILING-MOUNTED, DUAL-HEAD	12

4.1  Clinical & Functional Specifications	12

4.2  Engineering Technical Specifications	12

4.3  Standard Accessories	12

4.4  Installation Requirements	13

5. MOBILE OT LIGHT	14

5.1  Clinical & Functional Specifications	14

5.2  Engineering Technical Specifications	14

5.3  Standard Accessories	14

5.4  Installation Requirements	14

6. SURGICAL MICROSCOPE (GENERAL)	15

6.1  Clinical & Functional Specifications	15

6.2  Engineering Technical Specifications	15

6.3  Standard Accessories	15

6.4  Installation Requirements	16

7. SURGEON CONTROL PANEL (OR INTEGRATION SYSTEM)	17

7.1  Clinical & Functional Specifications	17

7.2  Engineering Technical Specifications	17

7.3  Standard Accessories	17

7.4  Installation Requirements	18

8. SURGICAL PENDANT (ANAESTHESIA + SURGICAL + ELECTRICAL)	19

8.1  Clinical & Functional Specifications	19

8.2  Engineering Technical Specifications	19

8.3  Standard Accessories	19

8.4  Installation Requirements	20

9. ELECTROSURGICAL UNIT — MONOPOLAR + BIPOLAR DIATHERMY	21

9.1  Clinical & Functional Specifications	21

9.2  Engineering Technical Specifications	21

9.3  Standard Accessories	21

9.4  Installation Requirements	22

10. ADVANCED BIPOLAR VESSEL-SEALING GENERATOR	23

10.1  Clinical & Functional Specifications	23

10.2  Engineering Technical Specifications	23

10.3  Standard Accessories	23

10.4  Installation Requirements	23

11. HIGH-VACUUM SURGICAL SUCTION UNIT	24

11.1  Clinical & Functional Specifications	24

11.2  Engineering Technical Specifications	24

11.3  Standard Accessories	24

11.4  Installation Requirements	24

12. HIGH-VACUUM SUCTION MACHINE (BEDSIDE/FLOOR)	26

12.1  Clinical & Functional Specifications	26

12.2  Engineering Technical Specifications	26

12.3  Standard Accessories	26

12.4  Installation Requirements	26

13. PORTABLE / HANDHELD SUCTION UNIT	27

13.1  Clinical & Functional Specifications	27

13.2  Engineering Technical Specifications	27

13.3  Standard Accessories	27

13.4  Installation Requirements	27

14. HEAVY-DUTY DUAL-BOTTLE SUCTION MACHINE	28

14.1  Clinical & Functional Specifications	28

14.2  Engineering Technical Specifications	28

14.3  Standard Accessories	28

14.4  Installation Requirements	28

15. PNEUMATIC TOURNIQUET SYSTEM	29

15.1  Clinical & Functional Specifications	29

15.2  Engineering Technical Specifications	29

15.3  Standard Accessories	29

15.4  Installation Requirements	29

16. SURGICAL HEADLIGHT (LED)	30

16.1  Clinical & Functional Specifications	30

16.2  Engineering Technical Specifications	30

16.3  Standard Accessories	30

16.4  Installation Requirements	30

# 1. OPERATING TABLE — ELECTRO-HYDRAULIC, MULTI-POSITION

Quantity: 4

## 1.1  Clinical & Functional Specifications

The operating table shall be a universal, general-purpose, electro-hydraulic surgical table intended for performing a wide range of surgical procedures including general surgery, orthopedic, gynecological, ENT, urological, vascular, and minor cardio-thoracic surgeries in a modern tertiary-care operating theatre. The table shall support full patient positioning including supine, prone, lateral, Trendelenburg, reverse-Trendelenburg, lithotomy, beach-chair, kidney, and lateral tilt positions. It shall be electro-hydraulically driven, motor-controlled, and shall permit motorized adjustment of height, longitudinal tilt, lateral tilt, back section, leg section, and kidney bridge. Positioning shall be smooth, silent, vibration-free, and shall not disturb the surgical field. The table shall be C-arm compatible over the thorax, abdomen, and pelvis regions to permit intraoperative fluoroscopy.

## 1.2  Engineering Technical Specifications

- Table Type: Electro-hydraulic, motorized, mobile operating table with floor-locking mechanism.

- Table Top: Multi-sectional, minimum 4 sections (head, back, seat, leg), radiolucent in specified zones, X-ray translucent, cassette-compatible with full-length tunnel for cassette insertion.

- Table Top Material: Antistatic, seamless, non-porous, chemical-resistant, easy-to-clean synthetic surface compatible with hospital-grade disinfectants.

- Overall Table-Top Length: Minimum 2000 mm (approximately) including head section.

- Table-Top Width: Minimum 500 mm without side rails.

- Patient Safe Working Load: Minimum 250 kg in normal orientation; minimum 200 kg with full articulation.

- Height Adjustment Range: Minimum 600 mm to maximum 1000 mm (approximately), motorized.

- Trendelenburg / Reverse Trendelenburg: Minimum ±30°.

- Lateral Tilt: Minimum ±20°.

- Back Section Adjustment: Minimum -40° to +80°.

- Leg Section Adjustment: Minimum -90° to +80°, split-leg configuration.

- Kidney Bridge Elevation: Minimum 100 mm (where applicable).

- Longitudinal Shift: Minimum 300 mm for extended C-arm coverage (where configured).

- Control System: Microprocessor-controlled with membrane keypad; auto-leveling, auto-zero return, and memory-position recall functions.

- Drive System: Electro-hydraulic actuators, silent operation, leak-free.

- Base: Heavy-duty steel base with antistatic castors and integrated central floor-locking.

- Side Rails: Standard EU/US-profile stainless steel accessory rails on both sides for universal accessory clamping.

- Power Supply: 100–240 VAC, 50/60 Hz, single-phase.

- Battery Backup: Integrated rechargeable sealed maintenance-free battery, minimum 50 full-position cycles per full charge.

- Charging Time: Maximum 8 hours for full charge.

- Protection Class: IPX4 or better for table top.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-46 (particular requirements for operating tables), ISO 13485 manufacturing, CE MDR or US FDA cleared.

## 1.3  Standard Accessories

- High-density, multi-section, antistatic, memory-foam mattress pad set: Minimum 70 mm thickness, seamless, antistatic, radiolucent, pressure-distributing viscoelastic foam; waterproof, seamless PU cover; latex-free; fluid-impermeable; disinfectant-compatible; matched to table-top sections with anti-slip fixation.

- Handheld wired remote control unit: Ergonomic, waterproof (minimum IPX7), membrane keypad, backlit function keys, minimum 3 m coiled cable, all-position control with emergency stop.

- Auxiliary override control panel: Column-integrated membrane panel providing full manual override in case of hand-control failure; sealed against fluid ingress.

- Adjusting swivel armboards with fixation straps and clamps (1 pair): Radiolucent, padded, swivel range minimum 180°, standard side-rail compatible clamps, load capacity minimum 15 kg per board, autoclavable straps.

- Padded lateral support braces (1 pair): Radiolucent core with high-density foam padding, adjustable ball-joint clamps, standard rail compatible, fluid-resistant covers.

- Padded shoulder support braces (1 pair): Adjustable, atraumatic, contoured design, side-rail mounted, load-rated for Trendelenburg position.

- Adjusting anesthesia screen frame with clamp: Stainless steel, telescopic, height-adjustable, universal rail-clamp mount, foldable for storage.

- Reusable body restraint strap with metal buckle: Minimum 2 m length, wide contoured design, chemical/disinfectant-compatible fabric, non-slip patient interface.

- Integrated rechargeable backup battery with power cord: Sealed maintenance-free type, integrated charge-status indicator, medical-grade detachable power cord.

## 1.4  Installation Requirements

Level, non-porous, seamless OT flooring; single-phase 100–240 VAC, 50/60 Hz supply with dedicated equipotential earthing; minimum 500 mm clearance around table for accessory movement; ambient temperature 15–30 °C, humidity 30–75% non-condensing; compatible with mobile C-arm imaging systems.

## 1.5  Documentation

User manual, service manual, illustrated spare-parts catalogue, factory test certificate, electrical safety test report (IEC 62353), preventive maintenance schedule, installation and commissioning reports, calibration certificate, CE MDR / FDA certificate, and ISO 13485 certificate.

## 1.6  Training

Clinical training for OT nursing and surgical staff; biomedical engineering training on electro-hydraulic system, actuator servicing, and troubleshooting; user-level PM training; on-site application training after installation.

## 1.7  Warranty

Minimum two years comprehensive warranty covering labour, spare parts, preventive maintenance (minimum two PM visits per year), breakdown maintenance, software updates, and technical support.

# 2. ELECTRO-HYDRAULIC OT TABLE — CARDIAC DEDICATED

Quantity: 1

## 2.1  Clinical & Functional Specifications

Dedicated cardiac surgical operating table intended for open-heart surgery, coronary artery bypass grafting, valve procedures, and hybrid cardiac interventions requiring full-length imaging. The table shall provide full-body C-arm access, longitudinal shift for cranial-to-femoral fluoroscopic coverage, radiolucent carbon-fiber top, and high load capacity for adult and bariatric cardiac patients. All positioning functions shall be motorized and controllable via wired and wireless remotes.

## 2.2  Engineering Technical Specifications

- Table Type: Electro-hydraulic, motorized, cardiac-dedicated with full-length longitudinal shift.

- Table Top: Segmented, radiolucent carbon-fiber, seamless, X-ray translucent over full-body length.

- Overall Length: Minimum 2100 mm.

- Table Top Width: Minimum 500 mm.

- Patient Safe Working Load: Minimum 300 kg.

- Height Range: Minimum 600–1050 mm.

- Trendelenburg / Reverse: Minimum ±30°.

- Lateral Tilt: Minimum ±20°.

- Longitudinal Shift: Minimum 400 mm.

- Back Section: Minimum -40° to +80°.

- Control: Microprocessor with memory positions, auto-zero return, wired + wireless remote.

- Power: 100–240 VAC, 50/60 Hz.

- Battery: Integrated backup, minimum 50 cycles per full charge.

- Applicable Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-46, CE MDR / FDA cleared, ISO 13485.

## 2.3  Standard Accessories

- Segmented radiolucent carbon-fiber mattress platform: Continuous X-ray translucency, seamless, chemical-resistant.

- High-density anti-shear antistatic memory foam mattress set: Minimum 80 mm thickness, pressure-relieving viscoelastic core, seamless PU cover, radiolucent.

- Handheld wired remote + wireless backup: IPX7, membrane keypad, encrypted wireless pairing, minimum 5 m operating range.

- Auxiliary override panel: Column-mounted, sealed manual override.

- Extra-wide pivoting armboards (1 pair): Multi-axis clamps, radiolucent, load ≥15 kg each.

- Heavy-duty sternal/lateral support pads (1 pair): Padded, universal rail-clamp mount.

- Telescopic anesthesia screen with dual clamps: Stainless steel, height-adjustable.

- Integrated longitudinal shift mechanism: Minimum 400 mm motorized shift.

- High-capacity rechargeable backup battery: Sealed, maintenance-free.

## 2.4  Installation Requirements

Dedicated 100–240 VAC medical-grade supply with equipotential earthing; C-arm-compatible OT layout; ambient 15–30 °C, humidity 30–75%.

5--7. Documentation, Training, Warranty

Same as Section 1. Minimum two years comprehensive warranty.

# 3. FULLY HYDRAULIC OT TABLE — UROLOGY / GENERAL SURGERY

Quantity: 1

## 3.1  Clinical & Functional Specifications

Fully hydraulic, manually-operated operating table for urology, endoscopic urological procedures, cystoscopy, TURP, and general surgery. The table shall include a kidney elevator/bridge, lithotomy configuration, and drainage system for wet procedures. No electrical power required for positioning.

## 3.2  Engineering Technical Specifications

- Type: Fully hydraulic, foot-pedal-operated, manual positioning.

- Table Top: Multi-sectional, radiolucent zones, seamless antistatic surface, waterproof.

- Length: Minimum 2000 mm.

- Width: Minimum 500 mm.

- Safe Working Load: Minimum 200 kg.

- Height Range: Minimum 750–1000 mm, foot-pedal hydraulic.

- Trendelenburg / Reverse: Minimum ±25°.

- Lateral Tilt: Minimum ±15°.

- Back Section: Minimum -30° to +80°.

- Leg Section: Split-leg, minimum -90° to +30°.

- Kidney Bridge: Mechanical elevator, minimum 100 mm lift.

- Drainage: Integrated drainage channel with removable collection tray.

- Base: Heavy-duty stainless-steel/chrome-plated base with central floor-lock and antistatic castors.

- Applicable Standards: IEC 60601-2-46, CE MDR / FDA cleared, ISO 13485.

## 3.3  Standard Accessories

- Foot-pedal hydraulic lift + manual positioning levers: Bilateral pedal access, corrosion-resistant.

- High-density waterproof antistatic mattress pad set: Minimum 60 mm thickness, seamless waterproof PU cover, disinfectant-compatible.

- Integrated mechanical kidney elevator/bridge: Minimum 100 mm elevation, manual crank/lever operation.

- Padded Goepel-type lithotomy leg crutches (1 pair): Heavy-duty ball-joint clamps, adjustable height and abduction, padded knee support, load-rated ≥15 kg each.

- Drainage basin holder with removable stainless-steel urology tray: Autoclavable AISI 304 stainless steel.

- Multi-axis adjustable armboards (1 pair): Radiolucent, padded, universal clamps.

- Adjustable anesthesia screen frame with clamp: Stainless steel telescopic.

- Heavy-duty body restraint strap: Wide contoured, disinfectant-compatible fabric.

## 3.4  Installation Requirements

Level OT floor with drainage compatibility; no electrical power required for positioning; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive warranty.

# 4. OT LIGHTS — LED, CEILING-MOUNTED, DUAL-HEAD

Quantity: 4

## 4.1  Clinical & Functional Specifications

Ceiling-mounted, dual-head, high-intensity LED surgical lighting system providing shadow-free, cold, true-color illumination of the surgical field for all major surgical procedures. Both lightheads shall be independently maneuverable through spring-balanced articulated arms with 360-degree rotation and shall support integration of a medical-grade HD video camera.

## 4.2  Engineering Technical Specifications

- Light Source: High-efficiency medical-grade LED with minimum 30,000 hours rated life.

- Central Illuminance (Ec) at 1 m per head: Minimum 130,000 lux, adjustable.

- Color Temperature: Selectable within 3500–5000 K, minimum 3 selectable settings.

- Color Rendering Index (CRI / Ra): Minimum 95.

- R9 (Red rendering): Minimum 90.

- Light Field Diameter (d10): Adjustable, minimum 200–300 mm.

- Depth of Illumination (L1+L2): Minimum 1000 mm at 60% Ec.

- Heat Emission at Surgical Field: Maximum 500 mW/m²/klx.

- Shadow Dilution: Automatic shadow management with multi-reflector or multi-LED technology; performance maintained with obstruction and cavity.

- Control: Touch-screen wall panel + on-lighthead sterile control.

- Suspension: Multi-axis spring-balanced arms, minimum 360° rotation at each joint.

- Camera Preparation: Integrated HD/Full-HD camera port in main head, hot-swappable.

- Power Supply: 100–240 VAC, 50/60 Hz; integrated central power supply unit with redundancy.

- Backup Power: Compatible with hospital UPS/emergency supply, automatic switchover.

- Standards: IEC 60601-1, IEC 60601-2-41 (surgical luminaires), IEC 60601-1-2 (EMC), CE MDR / FDA, ISO 13485.

## 4.3  Standard Accessories

- Dual-head LED assembly (Main + Satellite): Main head minimum 130,000 lux, satellite head minimum 100,000 lux, both with independent controls.

- Multi-axis spring-balanced ceiling suspension arms: 360° rotation at each joint, stable positioning without drift.

- Autoclavable detachable sterile handles (4 pieces): Steam-autoclavable at 134 °C, ergonomic grip, quick-release.

- Wall-mounted touchscreen control panel: Minimum 5-inch capacitive touch, membrane-sealed, all-function control.

- Integrated medical-grade HD camera preparation: Minimum Full-HD (1920×1080), sterile handle control, video output HD-SDI/DVI/HDMI.

- Automatic shadow management system: Sensor-based or multi-source shadow compensation.

- Central synchronization power supply unit: Redundant PSU, UPS-compatible, wall/ceiling-cabinet mounted.

## 4.4  Installation Requirements

Reinforced ceiling with minimum load rating 4× system weight; concealed ceiling flange with pre-installed anchor plate; false-ceiling coordination minimum 300 mm clearance; 100–240 VAC dedicated supply with equipotential earthing; UPS backup recommended; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive warranty including LED module.

# 5. MOBILE OT LIGHT

Quantity: 3

## 5.1  Clinical & Functional Specifications

Mobile, battery-backed LED surgical light for emergency operating rooms, minor surgery, and as backup light during ceiling-light failure.

## 5.2  Engineering Technical Specifications

- Light Source: Medical-grade LED, minimum 30,000 hours.

- Central Illuminance (Ec) at 1 m: Minimum 100,000 lux.

- Color Temperature: 4000–5000 K.

- CRI (Ra): Minimum 90.

- Light Field Diameter: Adjustable, minimum 180–250 mm.

- Depth of Illumination: Minimum 700 mm at 60% Ec.

- Focus: Manual, integrated in sterile handle.

- Stand: Heavy-duty mobile with 5-caster base, minimum 4 with antistatic lockable wheels.

- Arm: Spring-balanced articulated arm with minimum 3 joints.

- Battery Backup: Minimum 2 hours at full intensity.

- Charging Time: Maximum 8 hours.

- Power: 100–240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-2-41, IEC 60601-1-2, CE MDR / FDA.

## 5.3  Standard Accessories

- High-luminance LED head with manual focus: Integrated in sterile handle.

- Heavy-duty mobile stand with articulated spring-balanced arm: Stable, drift-free.

- Four antistatic lockable casters: Minimum 100 mm diameter.

- Autoclavable sterile handles (2 pieces): 134 °C steam sterilization.

- Integrated digital control panel: Intensity and mode indicator on head housing.

- Built-in rechargeable battery pack: Minimum 2-hour runtime; Li-ion or sealed lead-acid; integrated charge indicator; medical-grade power cord.

## 5.4  Installation Requirements

100–240 VAC standard socket; no fixed installation required.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 6. SURGICAL MICROSCOPE (GENERAL)

Quantity: 4

## 6.1  Clinical & Functional Specifications

High-performance surgical operating microscope for ENT, ophthalmic, plastic, hand, and general microsurgical procedures. It shall provide high-magnification, stereoscopic, coaxially illuminated imaging with motorized zoom, focus, and X-Y positioning, and shall support HD video documentation.

## 6.2  Engineering Technical Specifications

- Optical System: Apochromatic, stereoscopic, coaxial illumination.

- Magnification: Motorized zoom, minimum range 0.4× to 2.4× (total system magnification 4× to 25× with standard eyepieces and objective).

- Zoom Ratio: Minimum 1:6.

- Working Distance: Variable, minimum 200–400 mm (Varioskop-type or interchangeable objectives).

- Eyepieces: Widefield 10× or 12.5×, individual diopter adjustment ±5.

- Binocular Tube: Tiltable 0–180°, inclinable.

- Illumination: High-intensity LED or Xenon, minimum 40,000 lux at focal point, coaxial with red-reflex.

- Backup Illumination: Automatic switchover backup lamp.

- Focus / Zoom Control: Motorized via multi-function foot pedal.

- Stand: Floor-standing with electromagnetic braking on all axes.

- Casters: Lockable antistatic.

- Video Output: Integrated HD (minimum 1920×1080) camera with medical-grade display, HD-SDI/HDMI/DVI output.

- Power: 100–240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-41 (where applicable), CE MDR / FDA, ISO 13485.

## 6.3  Standard Accessories

- High-stability floor stand with electromagnetic brakes and lockable casters.

- Binocular tube, tilt 0–180°.

- Widefield 10× or 12.5× eyepieces with diopter adjustment.

- Motorized zoom/focus system with multi-function foot pedal: Waterproof, minimum 6 functions.

- Integrated LED or Xenon illumination with backup lamp module.

- Variable working-distance objective (Varioskop or multi-focal): 200–400 mm.

- HD camera + medical-grade monitor: Minimum 24-inch display, Full-HD, wall/arm mounted.

- Autoclavable protective caps/covers for manual adjustment knobs.

## 6.4  Installation Requirements

Level OT floor, 100–240 VAC dedicated socket, ambient 15–30 °C, humidity 30–75%.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 7. SURGEON CONTROL PANEL (OR INTEGRATION SYSTEM)

Quantity: 4

## 7.1  Clinical & Functional Specifications

Fully integrated OR control and communication system that allows the surgeon or circulating nurse to control OT lights, table position, HVAC, music, video routing, and timers from a single wall-mounted touch panel. The system shall provide centralized digital routing of audio/video signals between imaging sources (endoscopy, C-arm, microscope) and displays within the OR and to remote conferencing endpoints.

## 7.2  Engineering Technical Specifications

- Enclosure: Wall-flush medical-grade stainless-steel housing, IP54 front.

- Display: Minimum 15-inch multi-touch capacitive medical-grade, minimum 1920×1080 resolution, antibacterial glass, brightness minimum 300 cd/m².

- Processor: Multi-core medical-grade embedded processor.

- Memory: Minimum 8 GB RAM.

- Storage: Minimum 256 GB SSD.

- Video Inputs/Outputs: Minimum 4 HD inputs and 4 HD outputs (HD-SDI/DVI/HDMI/DisplayPort), 4K-capable inputs preferred.

- Video Routing: Full any-to-any matrix routing with picture-in-picture and side-by-side.

- Audio: Bi-directional intercom, VoIP-capable.

- Environmental Control Interface: Interfaces for OT light dimming, HVAC set-point, table positioning (subject to OEM interface availability), pendant lighting.

- Timers/Clock: Digital real-time clock, elapsed-time, count-down, procedure timer.

- Connectivity: Ethernet (minimum 1 Gbps), Wi-Fi optional, DICOM/HL7 interface for image export.

- UPS Backup: Integrated or wall-cabinet UPS, minimum 15-minute runtime.

- Cybersecurity: IEC 80001 compliance, role-based access, encrypted data at rest and in transit.

- Standards: IEC 60601-1, IEC 60601-1-2, IEC 62304, IEC 62366, IEC 80001, DICOM, HL7, CE MDR / FDA, ISO 13485.

## 7.3  Standard Accessories

- Wall-flush stainless-steel enclosure: Medical-grade AISI 304, sealed IP54 front, cleanable seams.

- Multi-touch display: Minimum 15-inch, capacitive, antibacterial.

- Central control processing module: Audio, video, environment routing.

- Digital interfaces for OT lights, table, HVAC: Standard protocol adapters (DALI, KNX, BACnet, or manufacturer OEM APIs as configured).

- Medical-grade digital clock/timer display: Elapsed time, count-down, count-up.

- UPS for processor backup: Minimum 15-minute runtime, sealed maintenance-free.

## 7.4  Installation Requirements

Recessed wall cavity minimum 100 mm depth; structured cabling (Cat 6A minimum) to all AV endpoints; 100–240 VAC UPS-backed supply; equipotential earthing; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Additionally: cybersecurity documentation, software update policy, and system architecture drawings. Minimum two years comprehensive.

# 8. SURGICAL PENDANT (ANAESTHESIA + SURGICAL + ELECTRICAL)

Quantity: 8

## 8.1  Clinical & Functional Specifications

Ceiling-mounted articulated service pendant providing medical gases, electrical outlets, low-voltage data ports, and equipment mounting for anesthesia workstations and surgical equipment. Pendants shall provide ergonomic access, cable management, and shall permit smooth articulation without drift.

## 8.2  Engineering Technical Specifications

- Configuration: Single or dual articulated arms, ceiling-mounted, minimum 2 joints per arm with 330–360° rotation each.

- Load Capacity: Minimum 150 kg per arm (equipment + accessories).

- Braking: Pneumatic or mechanical braking at each joint.

- Vertical Travel: Minimum 600 mm (motorized where specified) for anesthesia pendant.

- Service Head: Fully loaded with medical gas terminal units, electrical outlets, data ports, and equipment rails.

- Medical Gas Outlets: Minimum O₂, N₂O, Medical Air (4 bar), Surgical Suction (vacuum), AGSS; compliant with ISO 7396-1 and BS 5682 / DIN / NIST / Ohmeda standard as specified.

- Electrical Sockets: Minimum 8 medical-grade universal sockets per pendant, IEC 60601-1 compliant, individually fused.

- Data Ports: Minimum 2× RJ45, 2× HDMI, 2× BNC, 2× USB.

- Shelves / Drawers: Minimum 2 height-adjustable shelves, integrated drawer unit.

- IV Pole: Heavy-duty stainless-steel IV pole with clamp, load minimum 10 kg.

- Power: 100–240 VAC, 50/60 Hz, isolated medical supply compatible.

- Standards: IEC 60601-1, IEC 60601-1-2, ISO 11197 (medical supply units), ISO 7396-1 (medical gas pipelines), CE MDR / FDA, ISO 13485.

## 8.3  Standard Accessories

- Ceiling-mounted single or dual articulated arms: Heavy-duty, 2-joint minimum.

- 360° rotation joints with braking: Pneumatic or mechanical, drift-free.

- Fully loaded service head console.

- Medical gas outlets: O₂, N₂O, Medical Air, Surgical Suction, AGSS; ISO 7396-1 compliant; color-coded per HTM 02-01 or equivalent.

- Universal electrical sockets: Minimum 8 per pendant, individually fused/breakered.

- Low-voltage data ports: RJ45 ×2, HDMI ×2, BNC ×2, USB ×2 minimum.

- Height-adjustable equipment shelves with drawer unit (2 shelves/pendant): Load ≥30 kg per shelf.

- Heavy-duty IV pole with clamp: Stainless steel, telescopic.

## 8.4  Installation Requirements

Reinforced ceiling with load rating minimum 4× pendant weight; coordinated ceiling flange with concealed anchoring; medical gas pipelines to ISO 7396-1; UPS-backed 100–240 VAC electrical supply; equipotential earthing; structured data cabling (Cat 6A).

5--7. Documentation, Training, Warranty

As per Section 1. Additionally: medical gas commissioning and pressure test reports per ISO 7396-1. Minimum two years comprehensive.

# 9. ELECTROSURGICAL UNIT — MONOPOLAR + BIPOLAR DIATHERMY

Quantity: 8

## 9.1  Clinical & Functional Specifications

Microprocessor-controlled electrosurgical generator providing monopolar cutting, monopolar coagulation, and bipolar coagulation for open and endoscopic surgery. Continuous contact-quality monitoring (return-electrode monitoring, REM/CQM) shall be provided for patient safety.

## 9.2  Engineering Technical Specifications

- Output Power (Monopolar Cut): Minimum 300 W into 300 Ω.

- Output Power (Monopolar Coag): Minimum 120 W into 300 Ω.

- Output Power (Bipolar): Minimum 70 W into 100 Ω.

- Operating Frequency: Within 350 kHz–500 kHz (RF).

- Modes: Minimum pure cut, blend cut (multiple hemostasis levels), spray/fulguration coag, contact/desiccation coag, bipolar coag, bipolar cut.

- Auto-start bipolar: With adjustable delay.

- Return-Electrode Monitoring (REM/CQM): Continuous impedance monitoring with audible/visual alarm and automatic shutoff.

- Display: Minimum 5-inch color LCD/TFT, resolution minimum 800×480, real-time power/mode display.

- User Interface: Membrane keys or touchscreen, sealed against fluid.

- Presets: Minimum 20 user-configurable programs.

- Foot Switches: Waterproof (IPX8), monopolar dual-pedal + bipolar single-pedal.

- Power Supply: 100–240 VAC, 50/60 Hz.

- Safety Class: Class I, Type CF defibrillation-proof applied parts.

- Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-2 (HF surgical equipment), CE MDR / FDA, ISO 13485.

## 9.3  Standard Accessories

- Dual-button monopolar hand-switching pencils (2 pieces): Reusable/autoclavable body with disposable-tip compatibility, ergonomic grip, standard 3-pin connector.

- Reusable adult silicone patient return plates with cable (2 pieces): Split-plate design compatible with REM/CQM, silicone construction, autoclavable, minimum contact area 110 cm².

- Reusable bipolar forceps — straight and angled tips — with connection cable: Stainless steel, insulated shafts, autoclavable, standard bipolar connector.

- Waterproof dual-pedal foot switch (monopolar): IPX8, color-coded yellow (cut) / blue (coag).

- Waterproof single-pedal foot switch (bipolar): IPX8, color-coded blue.

- Continuous contact-quality monitoring system: Integrated in generator.

- Mobile equipment trolley: Lockable antistatic casters, accessory drawers, cable management.

## 9.4  Installation Requirements

100–240 VAC medical-grade socket with dedicated earthing; ambient 15–30 °C; EMC-compliant OT environment.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 10. ADVANCED BIPOLAR VESSEL-SEALING GENERATOR

Quantity: 1

## 10.1  Clinical & Functional Specifications

Advanced bipolar generator for permanent sealing and division of vessels and tissue bundles up to and including large-diameter vessels (up to 7 mm). The generator shall provide automated feedback-controlled energy delivery with tissue-impedance sensing, automatic end-tone, and instrument recognition. Intended for open, laparoscopic, and thoracoscopic surgery.

## 10.2  Engineering Technical Specifications

- Sealing Capability: Vessels and tissue bundles up to minimum 7 mm diameter.

- Feedback Control: Real-time tissue impedance measurement with automatic energy modulation and end-of-cycle tone.

- Instrument Recognition: Automatic recognition of connected instrument with parameter auto-selection.

- Compatibility: Compatible with vendor-supplied family of reusable/single-use vessel-sealing instruments (open, laparoscopic).

- Display: Touchscreen minimum 7-inch, resolution minimum 800×480.

- Foot Switch: Waterproof multi-function pedal, IPX8.

- Modes: Minimum vessel seal, seal-and-cut, standard bipolar coag.

- Power Supply: 100–240 VAC, 50/60 Hz.

- Safety Class: Class I, Type CF defibrillation-proof.

- Data: USB export of procedure logs.

- Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-2, CE MDR / FDA, ISO 13485.

## 10.3  Standard Accessories

- Generator interface module with instrument recognition software.

- High-resolution touchscreen control panel.

- Waterproof multi-function foot pedal (IPX8).

- Instrument connection cable ports: Minimum 2.

- Mobile utility cart with lockable casters and cable management hooks.

- Main power cable with medical-grade plug.

## 10.4  Installation Requirements

100–240 VAC medical-grade socket with dedicated earthing; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 11. HIGH-VACUUM SURGICAL SUCTION UNIT

Quantity: 4

## 11.1  Clinical & Functional Specifications

Heavy-duty, oil-free, high-flow surgical suction unit for use in operating theatres requiring uninterrupted suction of blood, body fluids, and irrigation solutions. Dual-jar change-over allows continuous operation.

## 11.2  Engineering Technical Specifications

- Maximum Vacuum: Minimum -90 kPa (-675 mmHg).

- Free-Air Flow: Minimum 60 L/min.

- Pump Type: Oil-free, maintenance-free.

- Jar Capacity: 2 × 4 L, autoclavable polycarbonate.

- Change-over Valve: Automatic or manual, allowing seamless switching.

- Overflow Protection: Mechanical float + hydrophobic anti-bacterial filter.

- Vacuum Regulator: With analog/digital gauge, range 0 to -90 kPa.

- Foot Switch: Pneumatic or waterproof (IPX8).

- Noise Level: Maximum 60 dB(A) at 1 m.

- Chassis: Heavy-duty mobile, antistatic lockable casters, push handle.

- Power: 100–240 VAC, 50/60 Hz.

- Safety Class: Class I.

- Standards: IEC 60601-1, IEC 60601-1-2, ISO 10079-1 (high-flow high-vacuum medical suction), CE MDR / FDA.

## 11.3  Standard Accessories

- Heavy-duty mobile chassis with antistatic lockable casters and push handle.

- Autoclavable polycarbonate jars (2 × 4 L): Graduated, 134 °C compatible.

- Change-over valve: For seamless jar switching.

- High-flow oil-free vacuum pump.

- Overflow safety valve + anti-bacterial hydrophobic filter.

- Vacuum regulator with pressure gauge.

- Pneumatic or waterproof foot switch.

## 11.4  Installation Requirements

100–240 VAC socket; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 12. HIGH-VACUUM SUCTION MACHINE (BEDSIDE/FLOOR)

Quantity: 74

## 12.1  Clinical & Functional Specifications

Compact, bedside/floor-standing suction unit for ward, ICU, recovery, and general clinical use for suctioning of respiratory secretions and body fluids.

## 12.2  Engineering Technical Specifications

- Maximum Vacuum: Minimum -80 kPa (-600 mmHg).

- Free-Air Flow: Minimum 30 L/min.

- Pump Type: Direct-drive oil-free piston.

- Jar Capacity: 2 L or 3 L, autoclavable polycarbonate.

- Overflow Protection: Mechanical float shut-off + hydrophobic anti-bacterial inline filter.

- Vacuum Regulator: Analog negative-pressure gauge.

- Noise Level: Maximum 60 dB(A) at 1 m.

- Chassis: Compact mobile with integrated handle and castors.

- Power: 100–240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2, ISO 10079-1, CE MDR / FDA.

## 12.3  Standard Accessories

- Compact mobile casing with transport handle and casters.

- Autoclavable polycarbonate jar (1 × 2 L or 3 L).

- Mechanical overflow shut-off valve.

- Hydrophobic anti-bacterial inline filter.

- Direct-drive oil-free piston pump.

- Vacuum regulator with analog gauge.

- Main power cord with integrated cable wrap.

## 12.4  Installation Requirements

100–240 VAC standard socket; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 13. PORTABLE / HANDHELD SUCTION UNIT

Quantity: 3

## 13.1  Clinical & Functional Specifications

Battery-operated portable suction unit for emergency, ambulance, and patient-transport use.

## 13.2  Engineering Technical Specifications

- Maximum Vacuum: Minimum -70 kPa (-525 mmHg).

- Free-Air Flow: Minimum 25 L/min.

- Jar Capacity: 1 L autoclavable with overflow protection.

- Battery: Rechargeable lithium-ion, minimum 60 minutes continuous operation at maximum vacuum.

- Charging Time: Maximum 4 hours.

- Power Options: 100–240 VAC adapter + 12 V DC vehicle charger.

- Weight: Approximately 3–5 kg with battery.

- Housing: Impact-resistant, ergonomic handle.

- Protection Class: IP34 or better.

- Standards: IEC 60601-1, IEC 60601-1-2, ISO 10079-1, CE MDR / FDA.

## 13.3  Standard Accessories

- Impact-resistant rigid plastic housing with handle.

- Autoclavable 1 L collection canister with overflow protection.

- Built-in rechargeable Li-ion battery pack.

- AC adapter + 12 V DC vehicle charging cable.

- Hydrophobic anti-bacterial filter set (3 pieces).

- Padded transport shoulder bag.

## 13.4  Installation Requirements

Standard 100–240 VAC socket for charging; no fixed installation.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive (battery minimum one year).

# 14. HEAVY-DUTY DUAL-BOTTLE SUCTION MACHINE

Quantity: 1

## 14.1  Clinical & Functional Specifications

Heavy-duty, dual-bottle, high-capacity suction machine for prolonged high-volume aspiration in trauma, obstetric, and major surgical procedures.

## 14.2  Engineering Technical Specifications

- Maximum Vacuum: Minimum -90 kPa (-675 mmHg).

- Free-Air Flow: Minimum 80 L/min.

- Pump: Low-noise oil-free rotary vane.

- Jar Capacity: 2 × 5 L autoclavable polycarbonate.

- Switching: Manual or automatic dual-bottle change-over.

- Overflow Protection: Dual mechanical float + hydrophobic filter.

- Vacuum Regulator: Precision, large high-visibility gauge.

- Noise Level: Maximum 65 dB(A) at 1 m.

- Chassis: Stainless-steel trolley with perimeter rubber bumper and castors.

- Power: 100–240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2, ISO 10079-1, CE MDR / FDA.

## 14.3  Standard Accessories

- Stainless-steel mobile trolley with rubber bumper.

- Autoclavable graduated polycarbonate bottles (2 × 5 L).

- Bottle switching valve (manual/automatic).

- Low-noise oil-free rotary vane pump.

- Dual overflow protection (float + filter).

- Precision vacuum regulator with high-visibility gauge.

- Main power cord with medical-grade plug.

## 14.4  Installation Requirements

100–240 VAC medical-grade socket; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 15. PNEUMATIC TOURNIQUET SYSTEM

Quantity: 4

## 15.1  Clinical & Functional Specifications

Dual-channel electronic pneumatic tourniquet system for orthopedic upper-limb and lower-limb surgery, providing controlled and monitored limb occlusion with independent pressure and time control for two limbs simultaneously.

## 15.2  Engineering Technical Specifications

- Channels: Independent dual channels.

- Pressure Range: 0–600 mmHg, adjustable per channel.

- Pressure Accuracy: Within ±5 mmHg.

- Timer Range: 0–360 minutes per channel with audible alarm.

- Alarms: Over-pressure, under-pressure, hose disconnection, time expiry, low battery.

- Display: Minimum dual digital pressure/time display per channel.

- Pressure Source: Internal compressor or medical air supply.

- Battery Backup: Minimum 2 hours.

- Power: 100–240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2, CE MDR / FDA.

## 15.3  Standard Accessories

- Dual-channel electronic control console with independent displays.

- Mobile pole stand with lockable casters and basket.

- Autoclavable lint-free upper-limb cuffs (S/M/L).

- Autoclavable lint-free lower-limb cuffs (M/L).

- Color-coded dual-line connecting hoses with quick-release couplings.

- Rechargeable backup battery + main power adapter.

## 15.4  Installation Requirements

100–240 VAC socket; medical air outlet where external air source specified; ambient 15–30 °C.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive.

# 16. SURGICAL HEADLIGHT (LED)

Quantity: 7

## 16.1  Clinical & Functional Specifications

High-intensity, cordless LED surgical headlight providing focused, shadow-free spot illumination for surgical procedures in cavities and confined spaces.

## 16.2  Engineering Technical Specifications

- Light Source: Medical-grade LED, minimum 30,000 hours life.

- Illuminance at 400 mm: Minimum 50,000 lux.

- Color Temperature: 4500–5500 K.

- CRI (Ra): Minimum 85.

- Spot Size: Adjustable, minimum 30–100 mm at 400 mm working distance.

- Weight of Head Assembly: Approximately 100–200 g (excluding battery).

- Battery Type: Rechargeable lithium-ion, belt-clip mounted.

- Battery Runtime: Minimum 4 hours per battery at full intensity.

- Charging Time: Maximum 4 hours.

- Headband: Ergonomic, fully adjustable, with rear counterweight.

- Standards: IEC 60601-1, IEC 60601-1-2, IEC 60601-2-41 (where applicable), CE MDR / FDA.

## 16.3  Standard Accessories

- Ergonomic adjustable headband assembly.

- High-intensity LED optic module with adjustable spot control.

- Rechargeable Li-ion battery packs with belt clip (2 per headlight).

- Multi-slot intelligent desktop battery charger.

- Fiber-optic backing/connecting cables (where console-style charging applies).

- Rigid padded storage/transport case.

## 16.4  Installation Requirements

Standard 100–240 VAC socket for battery charger; no fixed installation.

5--7. Documentation, Training, Warranty

As per Section 1. Minimum two years comprehensive (battery minimum one year)