TECHNICAL TENDER SPECIFICATIONS — UROLOGY DEPARTMENT

Technical Tender Document

Table of Contents

1. Complete Lower Tract Endoscopic Set (Cystoscope)	6

1.1 Clinical & Functional Specifications	6

1.2 Engineering Technical Specifications	6

1.3 Standard Accessories	6

1.4 Installation Requirements	7

1.5 Documentation	7

1.6 Training	7

1.7 Warranty	7

2. Resectoscope Set (for TURP / TURBT)	8

2.1 Clinical & Functional Specifications	8

2.2 Engineering Technical Specifications	8

2.3 Standard Accessories	8

2.4 Installation Requirements	9

2.5 Documentation	9

2.6 Training	9

2.7 Warranty	9

3. Semi-Rigid Ureteroscope (6 / 7.5 Fr)	10

3.1 Clinical & Functional Specifications	10

3.2 Engineering Technical Specifications	10

3.3 Standard Accessories	10

3.4 Installation Requirements	10

3.5 Documentation	11

3.6 Training	11

3.7 Warranty	11

4. Optical Internal Urethrotome (Sachse Type)	12

4.1 Clinical & Functional Specifications	12

4.2 Engineering Technical Specifications	12

4.3 Standard Accessories	12

4.4 Installation Requirements	12

4.5 Documentation, 6. Training, 7. Warranty	12

5. Stone Punch (Mauermayer Type)	13

5.1 Clinical & Functional Specifications	13

5.2 Engineering Technical Specifications	13

5.3 Standard Accessories	13

5.4 Installation, Documentation, Training, Warranty	13

6. Dual-Action Intracorporeal Lithotripsy System	14

6.1 Clinical & Functional Specifications	14

6.2 Engineering Technical Specifications	14

6.3 Standard Accessories	14

6.4 Installation Requirements	15

6.5 Documentation	15

6.6 Training	15

6.7 Warranty	15

7. Endoscopic Camera System, Light Source, Light-Guide Cable	16

7.1 Clinical & Functional Specifications	16

7.2 Engineering Technical Specifications	16

7.3 Standard Accessories	16

7.4 Installation Requirements	17

7.5 Documentation	17

7.6 Training	17

7.7 Warranty	17

8. Lithotripter (Extracorporeal Shockwave Lithotripsy — ESWL)	18

8.1 Clinical & Functional Specifications	18

8.2 Engineering Technical Specifications	18

8.3 Standard Accessories	18

8.4 Installation Requirements	19

8.5 Documentation	19

8.6 Training	19

8.7 Warranty	19

9. Major Urology Instrument Set	20

9.1 Clinical & Functional Specifications	20

9.2 Engineering Technical Specifications	20

9.3 Standard Accessories	20

9.4 Installation Requirements	20

9.5 Documentation, Training, Warranty	20

10. Minor Urology Instrument Set	22

10.1 Clinical & Functional Specifications	22

10.2 Engineering Technical Specifications	22

10.3 Standard Accessories	22

10.4 Installation, Documentation, Training, Warranty	22

11. Rib Retractors — Finochietto Type	23

11.1 Clinical & Functional Specifications	23

11.2 Engineering Technical Specifications	23

11.3 Standard Accessories	23

11.4 Installation, Documentation, Training, Warranty	23

12. Vascular Clamps Set (Satinsky, Bulldog, Lahey, DeBakey)	24

12.1 Clinical & Functional Specifications	24

12.2 Engineering Technical Specifications	24

12.3 Standard Accessories	24

12.4 Installation, Documentation, Training, Warranty	24

13. Microvascular Clamp Set	25

13.1 Clinical & Functional Specifications	25

13.2 Engineering Technical Specifications	25

13.3 Standard Accessories	25

13.4 Installation, Documentation, Training, Warranty	25

14. Self-Retaining Retractors — Balfour and Bookwalter	26

14.1 Clinical & Functional Specifications	26

14.2 Engineering Technical Specifications	26

14.3 Standard Accessories	26

14.4 Installation, Documentation, Training, Warranty	26

15. Ring (Sponge-Holding) Forceps Set	27

15.1 Clinical & Functional Specifications	27

15.2 Engineering Technical Specifications	27

15.3 Standard Accessories	27

15.4 Installation, Documentation, Training, Warranty	27

16. Suction Tips and Tubing Set	28

16.1 Clinical & Functional Specifications	28

16.2 Engineering Technical Specifications	28

16.3 Standard Accessories	28

16.4 Installation, Documentation, Training, Warranty	28

17. Kidney Pedicle Clamps Set	29

17.1 Clinical & Functional Specifications	29

17.2 Engineering Technical Specifications	29

17.3 Standard Accessories	29

17.4 Installation, Documentation, Training, Warranty	29

18. Urethral Sounds and Dilators	30

18.1 Clinical & Functional Specifications	30

18.2 Engineering Technical Specifications	30

18.3 Standard Accessories	30

18.4 Installation, Documentation, Training, Warranty	30

19. Metallic Suprapubic Cystostomy Trocar Set	31

19.1 Clinical & Functional Specifications	31

19.2 Engineering Technical Specifications	31

19.3 Standard Accessories	31

19.4 Installation Requirements	31

19.5 Documentation	31

19.6 Training	31

19.7 Warranty	31

# 1. Complete Lower Tract Endoscopic Set (Cystoscope)

Quantity: 1

## 1.1 Clinical & Functional Specifications

The Complete Lower Urinary Tract Endoscopic Set (Cystoscope) shall be intended for diagnostic and minor therapeutic endoscopic examination of the urethra, prostatic urethra, bladder neck, bladder wall, and ureteric orifices in adult patients. The system shall provide direct high-resolution visualization suitable for identification of mucosal lesions, tumors, calculi, strictures, foreign bodies, and inflammation, and shall permit ancillary interventions including biopsy, foreign body removal, ureteric catheterization, and stent placement. The set shall be fully compatible with standard hospital fiber-optic light sources, endoscopic camera systems, and irrigation systems, and shall be reusable, autoclavable/sterilizable, and suitable for repeated clinical use in an operating theatre or endoscopy suite environment.

## 1.2 Engineering Technical Specifications

- Optical System: Rigid rod-lens optical design or equivalent proven optical technology, offering wide-angle field of view of not less than 70° and high-definition image transmission compatibility.

- Telescope Angles Provided: 0°, 30°, and 70° directions of view (one each), interchangeable within the same working sheath.

- Telescope Outer Diameter: Approximately 2.7 mm to 4 mm, compatible with the supplied working sheaths.

- Working Length of Telescope: Minimum 290 mm, suitable for adult cystoscopy.

- Sheath Size Range: 19 Fr to 22 Fr, color-coded per international convention.

- Irrigation: Continuous-flow or single-flow capability through independent inflow/outflow stopcock valves; high-flow design.

- Instrument Channel: Minimum one working channel accepting 5 Fr to 7 Fr semi-rigid instruments; double-channel bridge with Albarran-type deflecting lever for directional control of ureteric catheters and instruments.

- Light Transmission: Fiber-optic light post compatible with universal adapters (Storz/Wolf/Olympus/ACMI equivalent connectors).

- Autoclavability: Fully steam-sterilizable at 134 °C; also compatible with STERRAD/ETO low-temperature sterilization.

- Materials: Medical-grade stainless steel, sapphire distal window, chemically resistant seals.

- Applicable Standards: IEC 60601-1, IEC 60601-2-18, ISO 8600 series (endoscopes — general requirements), ISO 17664 (reprocessing), CE MDR, FDA 510(k) or equivalent regulatory clearance.

## 1.3 Standard Accessories

- Rigid 0° diagnostic telescope — rod-lens optics, working length ≥ 290 mm, autoclavable, fiber-optic light post, compatible with supplied sheaths.

- Rigid 30° diagnostic telescope — identical mechanical specifications as 0°, 30° angle of view for bladder wall inspection.

- Rigid 70° diagnostic telescope — identical mechanical specifications, 70° angle of view for lateral/anterior wall inspection.

- Examination cystoscope sheaths (19 Fr—22 Fr): — Stainless steel, color-coded per Fr size, matched obturator, inflow/outflow ports.

- Universal double-channel bridge with Albarran lever: — Two working channels accepting 5—7 Fr instruments, deflecting lever with minimum 30° deflection.

- Semi-rigid biopsy forceps (5—7 Fr): — Stainless steel, sharp cutting cups, working length matching sheath, autoclavable.

- Semi-rigid grasping forceps (5—7 Fr): — Stainless steel, atraumatic and toothed variants, autoclavable.

- High-flow inflow/outflow stopcock valves: — Metal/PEEK, Luer-lock compatible, autoclavable.

- Heavy-duty fiber-optic light guide cable: — Length 2.5—3 m, minimum 3.5 mm active fiber bundle, universal adapters both ends, autoclavable/soakable.

- Perforated stainless steel sterilization tray: — Silicone matting/holders, secured lid, appropriately sized for all set components.

## 1.4 Installation Requirements

No fixed installation required. Requires access to a standard light source, endoscopic camera system, irrigation stand, and CSSD reprocessing facility with autoclave capability at 134 °C. Storage area to be maintained at 15—30 °C, humidity 30—75 % RH.

## 1.5 Documentation

User manual, reprocessing/sterilization manual, parts catalogue, factory test certificate, CE/FDA certificates, ISO 13485 certificate, warranty documentation.

## 1.6 Training

On-site clinical application training for urologists and OT staff (minimum 2 days); reprocessing and handling training for CSSD staff; biomedical handling and inspection training.

## 1.7 Warranty

Minimum 2 years comprehensive warranty covering optics integrity, mechanical function, and manufacturing defects, including labour, replacement parts, and technical support.

# 2. Resectoscope Set (for TURP / TURBT)

Quantity: 1

## 2.1 Clinical & Functional Specifications

The Resectoscope Set shall be intended for endoscopic transurethral resection of the prostate (TURP) and transurethral resection of bladder tumors (TURBT). It shall permit continuous-flow irrigation, high-frequency monopolar and/or bipolar electrosurgical cutting and coagulation, and clear visualization of the operative field. The system shall be compatible with standard electrosurgical generators and shall allow safe, controlled tissue resection and hemostasis in adult urological surgery.

## 2.2 Engineering Technical Specifications

- Telescope: Rigid wide-angle rod-lens optics, 12° and/or 30° angle of view, working length ≥ 300 mm, autoclavable.

- Working Element: Passive and/or active spring-return mechanism, compatible with monopolar and bipolar electrodes.

- Sheath Sizes: Inner sheath (continuous-flow) and outer sheath 24 Fr to 27 Fr, ceramic beak insulator, quick-release locking.

- Electrical Compatibility: Compatible with monopolar HF generators up to 400 W and bipolar generators up to 320 W, with universal connector configurations.

- Irrigation: Continuous-flow design with independent inflow/outflow stopcocks, high-flow capability.

- Materials: Medical-grade stainless steel, ceramic insulation, high-temperature-resistant polymers.

- Sterilization: Steam autoclavable at 134 °C; compatible with ETO and STERRAD.

- Standards: IEC 60601-1, IEC 60601-2-2 (HF surgical equipment), IEC 60601-2-18, ISO 8600, CE MDR, FDA clearance.

## 2.3 Standard Accessories

- Rigid 12° or 30° wide-angle telescope: — Rod-lens optics, autoclavable, universal light post.

- Passive/active working element: — Spring-loaded return, ergonomic thumb ring, monopolar/bipolar compatibility.

- Continuous-flow inner sheath with ceramic beak: — 24—27 Fr, ceramic insulator, autoclavable.

- Outer sheath with quick-release lock: — Compatible with laser and standard electrodes.

- Obturators (central & visual): — Matching sheath size, atraumatic tips, autoclavable.

- Reusable loop cutting electrodes — pack of 3: — Tungsten wire loop, monopolar/bipolar variants, standard cutting radius.

- Reusable roller ball coagulation electrodes — pack of 2: — Ball/barrel design, autoclavable.

- Monopolar/bipolar HF cables: — Universal connectors (compatible with common ESU generators), minimum 3 m length, autoclavable.

- Ellik evacuator assembly: — Glass or polycarbonate bulb, ≥ 500 mL, connecting cone, autoclavable.

- Perforated stainless steel sterilization tray: — Silicone holders, lid, adequately sized for the full set.

## 2.4 Installation Requirements

Compatible with standard HF electrosurgical generator, irrigation setup, and endoscopic camera/light source; no fixed installation.

## 2.5 Documentation

User & reprocessing manuals, parts catalogue, CE/FDA/ISO 13485 certificates, electrical safety compliance certificates.

## 2.6 Training

Urology surgeon training, CSSD reprocessing training, biomedical handling training.

## 2.7 Warranty

Minimum 2 years comprehensive.

# 3. Semi-Rigid Ureteroscope (6 / 7.5 Fr)

Quantity: 1

## 3.1 Clinical & Functional Specifications

The Semi-Rigid Ureteroscope shall be intended for endoscopic diagnostic and therapeutic access to the distal, mid, and proximal ureter for the visualization and management of ureteric calculi, strictures, tumors, and foreign bodies. It shall permit passage of laser fibers, lithotripsy probes, guidewires, baskets, and biopsy instruments through integrated working channels while providing continuous irrigation and high-resolution direct visualization.

## 3.2 Engineering Technical Specifications

- Distal Tip Diameter: Approximately 6.0 Fr; proximal shaft ≤ 7.5 Fr.

- Working Length: Minimum 430 mm.

- Optics: Integrated fiber-optic bundle with 5° or 12° direction of view, high-resolution image transmission.

- Working Channels: Two channels (dual) with combined 3.4 Fr to 5 Fr instrument capacity, permitting simultaneous irrigation and instrumentation.

- Laser Fiber Compatibility: Compatible with Holmium:YAG and Thulium laser fibers of 200—550 µm through sealing cap/valve.

- Irrigation: Dedicated stopcocks with Luer-lock adapters.

- Leakage Testing: Watertight, immersion leak-tested design.

- Sterilization: Autoclavable at 134 °C and compatible with ETO/STERRAD.

- Standards: IEC 60601-1, IEC 60601-2-18, ISO 8600, CE MDR, FDA clearance.

## 3.3 Standard Accessories

- Semi-rigid uretero-renoscope — 6/7.5 Fr, working length ≥ 430 mm, 5°/12° optics.

- Integrated dual working/irrigation channels — combined 3.4—5 Fr, sealed proximal ports.

- Detachable inflow stopcocks & Luer adapters — autoclavable.

- Silicone sealing caps/valves — for laser fiber entry, autoclavable, minimum 5 pieces.

- Semi-rigid foreign body forceps (3 Fr) — matching working length.

- Semi-rigid stone biopsy forceps (3 Fr) — matching working length.

- Leakage test adapter plug — manual/automatic, watertight.

- Rigid padded transport/sterilization case — protective interior, autoclavable or ETO-compatible.

## 3.4 Installation Requirements

No fixed installation. Requires standard light source, camera, and irrigation system.

## 3.5 Documentation

User manual, reprocessing manual, parts catalogue, CE/FDA/ISO 13485 certificates.

## 3.6 Training

Clinical and reprocessing training.

## 3.7 Warranty

Minimum 2 years comprehensive (including fiber-optic bundle integrity).

# 4. Optical Internal Urethrotome (Sachse Type)

Quantity: 1

## 4.1 Clinical & Functional Specifications

The Optical Internal Urethrotome shall be intended for the endoscopic visual incision (Sachse-type internal urethrotomy) of urethral strictures under direct optical control. It shall permit passage of cold knife blades and guidewires through a dedicated working channel with continuous irrigation.

## 4.2 Engineering Technical Specifications

- Telescope: Rigid wide-angle 0° or 12° rod-lens optics, working length compatible with sheath.

- Sheath Size: 21 Fr, with supplementary irrigation channel.

- Working Element: Spring-return handle with locking cutting-blade holder.

- Blade Types: Cold-knife straight, wavy (sickle), and hook blades, reusable, precision-ground.

- Guidewire Channel: Accommodating 0.035" / 0.038" standard urological guidewires.

- Sterilization: Autoclavable at 134 °C.

- Standards: IEC 60601-1, IEC 60601-2-18, ISO 8600, CE MDR, FDA clearance.

## 4.3 Standard Accessories

- Rigid 0° or 12° telescope — rod-lens optics, universal light post.

- Urethrotome working element — spring-loaded, ergonomic, blade-locking.

- Sheath 21 Fr — with dedicated irrigation channel, autoclavable.

- Half-round central obturator — matching sheath.

- Cold-knife blades (straight, wavy, hook) — reusable, precision-ground, autoclavable.

- Guidewire channel adapter — 0.035"/0.038" compatible.

- Perforated stainless steel sterilization tray — silicone matting.

## 4.4 Installation Requirements

No fixed installation.

## 4.5 Documentation, 6. Training, 7. Warranty

Standard as above (Item 1); minimum 2-year warranty.

# 5. Stone Punch (Mauermayer Type)

Quantity: 1

## 5.1 Clinical & Functional Specifications

The Stone Punch (Mauermayer-type) shall be intended for endoscopic mechanical crushing and evacuation of bladder calculi under direct visualization. It shall provide heavy-duty mechanical crushing capability through a rigid working sheath with matched optics and integrated flushing.

## 5.2 Engineering Technical Specifications

- Sheath: 24 Fr to 26 Fr, insulated tip, heavy-duty stainless steel.

- Crushing Mechanism: Mechanical screw-drive or lever, capable of crushing calculi up to \~30 mm diameter (typical clinical performance).

- Working Insert: Micro-lock coupling, autoclavable.

- Optics Compatibility: Compatible with standard rigid telescopes (0°/12°/30°).

- Sterilization: Steam autoclavable at 134 °C.

- Standards: IEC 60601-1, ISO 8600, CE MDR, FDA clearance.

## 5.3 Standard Accessories

- Heavy-duty working sheath (24—26 Fr) — insulated tip, autoclavable.

- Central obturator with quick-lock connection.

- Lithotrite working insert — micro-lock coupling.

- Reusable heavy-duty crushing jaws — mechanical screw handle, high-strength stainless steel.

- Direct flushing stopcocks — for irrigation and debris clearance.

- Heavy-duty perforated stainless steel tray — silicone matting.

## 5.4 Installation, Documentation, Training, Warranty

No fixed installation; standard documentation and training; minimum 2-year warranty.

# 6. Dual-Action Intracorporeal Lithotripsy System

Quantity: 1

## 6.1 Clinical & Functional Specifications

The Dual-Action Intracorporeal Lithotripsy System shall be intended for intracorporeal fragmentation of urinary tract calculi (renal, ureteric, bladder) using combined pneumatic (ballistic) and ultrasonic energy modalities, either independently or simultaneously, with integrated suction for continuous fragment evacuation.

## 6.2 Engineering Technical Specifications

- Modalities: Pneumatic (ballistic) and ultrasonic, independently and simultaneously selectable.

- Display: Color TFT LCD or equivalent, minimum 5-inch diagonal, digital indication of energy, frequency, pressure, and suction values.

- Pneumatic Energy: Adjustable single-shot and continuous modes; frequency selectable up to 15 Hz or better; adjustable impact energy.

- Ultrasonic Frequency: Approximately 24—27 kHz, adjustable amplitude/power up to 100 W or better.

- Suction: Integrated regulated vacuum with adjustable pressure and collection reservoir.

- Probes: Reusable ultrasonic (hollow) and ballistic (solid) probes in multiple diameters (0.8—3.8 mm) and lengths compatible with rigid and semi-rigid endoscopes.

- Foot Control: Waterproof dual-pedal control (IPX8 or better).

- Power Supply: 100—240 VAC, 50/60 Hz.

- Environmental: Operating 10—40 °C, 30—75 % RH.

- Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-2 (where applicable), IEC 62353, CE MDR, FDA clearance.

## 6.3 Standard Accessories

- Main control generator console — digital display of pressure/energy, adjustable parameters, alarms.

- Combined pneumatic + ultrasonic handpiece — lightweight, ergonomic, autoclavable or high-level disinfectable.

- Reusable ultrasonic probes — hollow, multiple diameters/lengths, autoclavable.

- Reusable ballistic probes — solid, multiple diameters/lengths, autoclavable.

- Integrated suction module + collection bottle bracket — adjustable vacuum, autoclavable reservoir.

- Probe fixation & torque wrench tools — manufacturer-standard.

- Waterproof dual-pedal foot switch — IPX8, minimum 3 m cable.

- Medical-grade power cable — appropriate national plug configuration.

## 6.4 Installation Requirements

Dedicated 220—240 VAC (or 100—240 VAC), 50/60 Hz single-phase outlet, protective earth, UPS-compatible. Compressed medical air / nitrogen supply per manufacturer specification (typically 4—8 bar dry, clean).

## 6.5 Documentation

User & service manuals, parts catalogue, calibration certificate, factory test certificate, electrical safety report, PM schedule.

## 6.6 Training

Clinical application training, biomedical PM training, troubleshooting training.

## 6.7 Warranty

Minimum 2 years comprehensive including labour, parts (excluding consumable probes after normal wear), PM visits, and software updates.

# 7. Endoscopic Camera System, Light Source, Light-Guide Cable

Quantity: 1

## 7.1 Clinical & Functional Specifications

The system shall be intended to provide high-definition video imaging and high-intensity cold-light illumination for rigid and flexible endoscopic urological procedures, with digital image capture, display, recording, and PACS/HIS connectivity capabilities. It shall support multiple specialties and be mountable on a mobile equipment tower for use in operating theatres and endoscopy suites.

## 7.2 Engineering Technical Specifications

- Camera Head: 3-CCD or 1-CMOS Full-HD or better; minimum resolution Full-HD 1920 × 1080; 4K UHD equivalent or better acceptable.

- Optical Zoom Coupler: Integrated, focal length adjustable, C-mount or equivalent universal coupling.

- Video Processor: Digital signal processing with white balance, image enhancement, noise reduction, freeze, zoom, dual-image, and picture-in-picture; multiple video outputs (HDMI, DVI, 3G-SDI, DisplayPort).

- Recording: Internal or external recording to USB, minimum 500 GB storage or better; still image capture in JPEG/TIFF; video in MPEG-4/H.264 or better.

- Light Source: LED (minimum 250 W equivalent) or Xenon (300 W typical); color temperature approximately 5500—6500 K; lamp life minimum 20,000 hours (LED) or per manufacturer for Xenon; adjustable intensity.

- Light-Guide Cable: Fiber-optic, length 2.5—3 m, active bundle ≥ 3.5 mm, universal endoscope adapters, autoclavable/soakable.

- Display: Minimum 24-inch medical-grade LED/LCD monitor, resolution minimum 1920 × 1080, wide viewing angle.

- Connectivity: DICOM 3.0, HL7, USB, LAN (RJ-45), Wi-Fi optional.

- Power Supply: 100—240 VAC, 50/60 Hz.

- Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-18, IEC 62304 (software), IEC 62366 (usability), DICOM, CE MDR, FDA clearance.

## 7.3 Standard Accessories

- Full-HD camera control head with integrated zoom coupler — autoclavable/soakable head; ergonomic buttons.

- HD digital video processor console — multi-format outputs, image processing, recording.

- LED or Xenon light source console — adjustable intensity, universal turret for endoscope connection.

- Fiber-optic light-guide cable (2.5—3 m) — active bundle ≥ 3.5 mm, autoclavable/soakable.

- Interconnecting cables — BNC, HDMI, DVI, DisplayPort as applicable, medical grade.

- Watertight cleaning cap set — for camera head reprocessing.

- Mobile equipment tower cart — lockable casters (minimum 4), integrated hospital-grade isolation transformer with power strip, monitor arm, shelves for all components.

## 7.4 Installation Requirements

Single-phase 100—240 VAC, 50/60 Hz supply with protective earthing; UPS compatible; hospital LAN drop for DICOM/HL7 connectivity.

## 7.5 Documentation

User & service manuals, parts catalogue, calibration certificate, electrical safety test report, DICOM conformance statement, PM schedule.

## 7.6 Training

Clinical, biomedical PM, application, and basic troubleshooting training.

## 7.7 Warranty

Minimum 2 years comprehensive including software updates.

# 8. Lithotripter (Extracorporeal Shockwave Lithotripsy — ESWL)

Quantity: 1

## 8.1 Clinical & Functional Specifications

The Extracorporeal Shockwave Lithotripter (ESWL) shall be intended for non-invasive fragmentation of renal and ureteric calculi through externally generated focused shockwaves, with real-time integrated localization via ultrasound and/or fluoroscopy. The system shall provide a motorized multi-axis patient table, ergonomic operator workstation, and safety interlocks suitable for outpatient and inpatient lithotripsy services.

## 8.2 Engineering Technical Specifications

- Shockwave Source: Electromagnetic, electrohydraulic, or piezoelectric technology; adjustable focal energy from low to high; energy levels selectable in graduated steps.

- Shockwave Frequency: Selectable, minimum 60—240 shocks/minute.

- Focal Zone: Adjustable focal depth suitable for typical adult renal/ureteric anatomy (approximately 50—170 mm from source).

- Localization: Integrated ultrasound (2—5 MHz sector/convex probe) and/or interface for C-arm fluoroscopy with isocentric alignment.

- Patient Table: Motorized multi-axis (X/Y/Z + tilt); radiolucent section for fluoroscopy; patient load capacity minimum 180 kg.

- Operator Workstation: Dual medical-grade monitors, minimum 19-inch each, resolution minimum 1280 × 1024; touchscreen or keyboard/mouse control.

- Water Cushion / Coupling: Sealed water cushion with degassed water circulation and temperature regulation (approximately 30—37 °C).

- Safety: Emergency stop button, patient/operator ECG synchronization option, radiation shielding (where fluoroscopy integrated), interlocks.

- Power Supply: 220—240 VAC or 380—415 VAC (as configured), 50/60 Hz; integrated isolation transformer.

- Standards: IEC 60601-1, IEC 60601-1-2 (EMC), IEC 60601-2-36 (extracorporeally induced lithotripsy), IEC 60601-2-37 (ultrasound), IEC 60601-2-54 (fluoroscopy, where applicable), CE MDR, FDA clearance.

## 8.3 Standard Accessories

- Motorized multi-axis patient treatment table — radiolucent section, minimum 180 kg capacity, remote control positioning.

- ESWL therapy head with shockwave generator — adjustable focal energy, self-contained coupling.

- Integrated ultrasound localization system and/or C-arm interface — real-time targeting.

- Handheld remote controller — for table and therapy head positioning.

- Central operator workstation console — dual medical-grade monitors.

- Water cooling & degassed water circulation module — with connecting hoses, filters, temperature regulation.

- Patient safety emergency stop switch module — hardwired, clearly labelled.

- Integrated isolation transformer and power distribution unit — hospital-grade, appropriate load rating.

## 8.4 Installation Requirements

Dedicated procedure room with minimum 25 m² floor area; reinforced flooring per equipment weight; three-phase or single-phase supply per system configuration; protective earth ≤ 1 Ω; HVAC-controlled environment (20—24 °C, 40—60 % RH); potable water inlet and drainage; radiation shielding of walls/doors as per national regulatory requirements where fluoroscopy is integrated; UPS optional for control console.

## 8.5 Documentation

User manual, service manual, parts catalogue, factory test certificate, calibration certificate, electrical safety report, radiation safety certificate (where applicable), commissioning and installation report, PM schedule.

## 8.6 Training

Clinical application training (minimum 5 days), biomedical PM training (minimum 3 days), radiation safety training (where applicable), troubleshooting training.

## 8.7 Warranty

Minimum 2 years comprehensive including labour, parts, PM visits (at least twice yearly), software updates, and technical support. Shockwave source (electrode/coil) coverage as per manufacturer's standard shock-count warranty.

# 9. Major Urology Instrument Set

Quantity: 1

## 9.1 Clinical & Functional Specifications

The Major Urology Instrument Set shall provide a comprehensive assortment of hand-held surgical instruments required for major open urological procedures including nephrectomy, cystectomy, prostatectomy, pyeloplasty, and ureteric reconstruction. All instruments shall be reusable, autoclavable, precision-manufactured, and suitable for repeated clinical use.

## 9.2 Engineering Technical Specifications

- Material: Medical-grade AISI 420 / AISI 410 martensitic stainless steel or equivalent, corrosion-resistant, non-magnetic where applicable.

- Finish: Satin/matte finish to reduce glare.

- Sterilization: Steam autoclavable at 134 °C; compatible with ETO/STERRAD.

- Hardness: Working tips heat-treated per ISO 7153-1.

- Standards: ISO 7153-1 (surgical instruments — materials), ISO 13485, CE MDR, FDA-listed.

## 9.3 Standard Accessories

- Mayo tissue scissors — straight & curved: — Lengths 15 cm and 17 cm, one pair each.

- Metzenbaum dissecting scissors — standard & long: — Lengths 18 cm and 23 cm, one pair each.

- DeBakey cardiovascular tissue forceps: — Lengths 16, 20, and 24 cm; atraumatic serrations.

- Tissue forceps with teeth (1×2): — Lengths 15, 20, and 25 cm.

- Crile and Kelly hemostatic forceps: — Straight and curved; lengths 14—16 cm; sets of 6 each.

- Allis and Babcock tissue grasping forceps: — Lengths 15—20 cm; 4 each.

- Rochester-Pean and Heaney hemostatic forceps: — Lengths 20—24 cm; 4 each.

- Mayo-Hegar and Masson needle holders (heavy-duty): — Lengths 18—24 cm with tungsten carbide inserts; 3 each.

- Deaver and Richardson retractors: — Various blade widths (25 mm, 50 mm, 75 mm); 2 each.

- Malleable copper ribbon retractors — set of 3 widths: — 25 mm, 50 mm, 75 mm; annealed copper or malleable stainless steel.

- Heavy-duty perforated stainless steel sterilization trays — with silicone mats and secured lids.

## 9.4 Installation Requirements

CSSD reprocessing facility with steam sterilization.

## 9.5 Documentation, Training, Warranty

Standard documentation; handling and reprocessing training; minimum 2-year warranty against manufacturing defects.

# 10. Minor Urology Instrument Set

Quantity: 1

## 10.1 Clinical & Functional Specifications

The Minor Urology Instrument Set shall provide a comprehensive assortment of hand-held surgical instruments required for minor open urological procedures such as circumcision, hydrocelectomy, orchidopexy, vasectomy, and superficial urological interventions.

## 10.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel (AISI 420/410 or equivalent).

- Finish: Satin/matte, corrosion-resistant.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 10.3 Standard Accessories

- Operating scissors (sharp/blunt, straight & curved): — Lengths 13—15 cm; 2 pairs each.

- Metzenbaum dissecting scissors (short/standard): — Lengths 14—18 cm.

- Adson & Standard tissue forceps (with & without teeth): — Lengths 12—15 cm.

- Halsted Mosquito hemostatic forceps (straight & curved): — Length 12.5 cm; set of 12.

- Crile hemostatic forceps (curved): — Length 14 cm; set of 6.

- Allis tissue forceps (short): — Length 15 cm; 4 pieces.

- Backhaus towel clamps: — Length 9—13 cm; set of 6.

- Mayo-Hegar needle holders (short & medium): — Lengths 14 and 16 cm.

- Senn-Miller and Army-Navy retractors: — 2 each.

- Stainless steel perforated cleaning tray with locking lid.

## 10.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 11. Rib Retractors — Finochietto Type

Quantity: 1

## 11.1 Clinical & Functional Specifications

The Finochietto Rib Retractor shall be intended for controlled retraction and separation of ribs during thoracoabdominal urological procedures and adrenal/renal surgery requiring transthoracic access.

## 11.2 Engineering Technical Specifications

- Material: Anodized aluminum or medical stainless steel; corrosion-resistant.

- Rack Mechanism: Heavy-duty screw-drive with self-locking ratchet.

- Spread Range: Minimum 0—260 mm.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 11.3 Standard Accessories

- Heavy-duty rack bar frame — anodized aluminum or stainless steel.

- Standard adult blades — set of 2: — Blade width approximately 50 mm, depth 50—75 mm.

- Deep profile blades — set of 2: — Blade depth 75—100 mm.

- Removable crank handle with mechanical locking pin.

- Rigid sterilization/containment storage box.

## 11.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 12. Vascular Clamps Set (Satinsky, Bulldog, Lahey, DeBakey)

Quantity: 1

## 12.1 Clinical & Functional Specifications

The Vascular Clamp Set shall be intended for atraumatic occlusion of major and minor vessels during vascular and urological reconstructive surgery, including renal pedicle control, IVC clamping, and tangential vessel occlusion.

## 12.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel with atraumatic DeBakey-pattern serrations on jaws.

- Jaw Design: Atraumatic; non-crushing; parallel closure.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 12.3 Standard Accessories

- Satinsky vena cava clamps (small, medium, large): — One each.

- DeBakey tangential occlusion clamps (various angles): — Straight, curved, angled — 3 pieces.

- Lahey dissecting/hemostatic forceps: — Lengths 18—22 cm; 2 pieces.

- Bulldog clamps (straight & curved, cross-action): — Set of 6, graduated sizes.

- Bulldog clamp applying forceps/carriers: — 2 pieces.

- Perforated stainless steel micro-instrument storage box — with silicone rows.

## 12.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 13. Microvascular Clamp Set

Quantity: 1

## 13.1 Clinical & Functional Specifications

The Microvascular Clamp Set shall be intended for atraumatic occlusion and approximation of small-caliber vessels during microvascular anastomosis, renal transplantation, and reconstructive urological microsurgery.

## 13.2 Engineering Technical Specifications

- Material: Titanium or medical-grade stainless steel; non-magnetic titanium preferred for microvascular use.

- Clamp Pressure: Graduated low-force closure (typically 15—50 g depending on vessel size).

- Vessel Range: 0.5—3 mm.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 13.3 Standard Accessories

- Approximator clamps with sliding bar (1 mm to 3 mm): — Set of 3 sizes.

- Single micro-bulldog clamps (straight & curved, graduated tensions): — Minimum 8 pieces.

- Micro-clamp applying forceps (lockless design): — 2 pieces.

- Ultra-fine micro-vessel dilators and jeweler-type forceps: — 2 pairs each.

- High-density protective sterilization case — with soft silicone slots.

## 13.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 14. Self-Retaining Retractors — Balfour and Bookwalter

Quantity: 1

## 14.1 Clinical & Functional Specifications

The Self-Retaining Retractor set shall be intended for hands-free retraction of abdominal wall and viscera during major open urological and pelvic procedures, providing stable, adjustable exposure of the operative field.

## 14.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel with radiopaque or anodized finish where applicable.

- Frame: Modular oval and segmental ring frames; table-post assembly with universal OR-table clamp.

- Blade Range: Multiple lengths and profiles (Kelly, Harrington, Malleable, Balfour styles).

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 14.3 Standard Accessories

- Balfour main spreader frame — with interchangeable side blades.

- Balfour center bladder blade attachment.

- Bookwalter table-post assembly with heavy-duty table clamp — universal OR-table compatible.

- Bookwalter oval and segmental ring frames — multiple sizes.

- Bookwalter ratchet slide clamps — set of 8.

- Bookwalter blade assortment (Kelly, Harrington, Malleable, Balfour styles) — minimum 8 blades of varied depths and widths.

- Heavy-duty stainless steel sterilization containers — with structural dividers.

## 14.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 15. Ring (Sponge-Holding) Forceps Set

Quantity: 1

## 15.1 Clinical & Functional Specifications

Ring/sponge-holding forceps shall be intended for atraumatic grasping and manipulation of surgical sponges, swabs, and delicate tissues during open urological procedures.

## 15.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel; satin finish.

- Jaws: Serrated ring pattern; ratcheted lock.

- Length: Approximately 9.5 inches (24 cm).

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 15.3 Standard Accessories

- Foerster sponge forceps, straight, serrated, 9.5-inch — 4 pieces.

- Foerster sponge forceps, curved, serrated, 9.5-inch — 4 pieces.

- Heavy-duty stainless steel organizing rack / wire tray holder — autoclavable.

## 15.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 16. Suction Tips and Tubing Set

Quantity: 1

## 16.1 Clinical & Functional Specifications

The Suction Tips and Tubing Set shall be intended for evacuation of blood, irrigation fluids, and debris from the surgical field during open and endoscopic urological procedures.

## 16.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel (tips) and medical-grade autoclavable silicone (hoses).

- Tips: Yankauer, Poole, and Frazier patterns; multiple Fr sizes.

- Sterilization: Autoclavable at 134 °C; silicone tubing repeatedly autoclavable (minimum 100 cycles).

- Standards: ISO 7153-1, ISO 10993 (biocompatibility for silicone), ISO 13485, CE MDR.

## 16.3 Standard Accessories

- Yankauer rigid suction tips with removable rosetip caps — stainless steel, 3 pieces.

- Poole abdominal suction sleeves with outer sheaths — stainless steel, 2 pieces.

- Frazier micro-suction tips (8 Fr, 10 Fr, 12 Fr) with stylets — one set each.

- Reusable autoclavable silicone suction hoses (3 m length) — 2 pieces, wall thickness sufficient for negative-pressure resistance.

- Universal tapered hose connectors and step adapters — stainless steel, autoclavable.

## 16.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 17. Kidney Pedicle Clamps Set

Quantity: 1

## 17.1 Clinical & Functional Specifications

Kidney Pedicle Clamps shall be intended for controlled occlusion of the renal pedicle during nephrectomy and partial nephrectomy, providing secure vascular control with atraumatic tissue handling.

## 17.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel; atraumatic serrations.

- Length: Approximately 22—26 cm.

- Jaw Curvature: Straight, curved, angled variants.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 17.3 Standard Accessories

- Herrick kidney pedicle clamps (long profile) — 2 pieces.

- Guyon kidney pedicle clamps (various jaw curvatures) — 3 pieces (straight, curved, angled).

- Stille-Pullan renal vascular occlusion clamps — 2 pieces.

- Perforated stainless steel instrumentation tray — with silicone matting.

## 17.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 18. Urethral Sounds and Dilators

Quantity: 1

## 18.1 Clinical & Functional Specifications

Urethral sounds and dilators shall be intended for gradual mechanical dilatation of urethral strictures and calibration of urethral caliber in male and female patients.

## 18.2 Engineering Technical Specifications

- Material: Medical-grade solid stainless steel; polished surface finish.

- Size Range: Van Buren curved sounds, graduated from 8 Fr to 36 Fr (French scale).

- Length: Approximately 25—27 cm.

- Otis Bougies à Boule / Dilators: Olive-tipped, graduated sizes.

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 18.3 Standard Accessories

- Van Buren curved urethral sounds — graduated 8 Fr to 36 Fr — minimum 12 sizes.

- Otis urethral dilator set / bougies à boule — graduated olive-tipped, minimum 6 sizes.

- Reusable protective canvas storage roll or perforated stainless steel case.

- Manufacturer-standard calibration verification gauge — for periodic Fr-size verification.

## 18.4 Installation, Documentation, Training, Warranty

Standard; minimum 2-year warranty.

# 19. Metallic Suprapubic Cystostomy Trocar Set

Quantity: 1

## 19.1 Clinical & Functional Specifications

The Metallic Suprapubic Cystostomy Trocar Set shall be intended for percutaneous suprapubic drainage of the urinary bladder, providing controlled introduction of a catheter through the anterior abdominal wall into a distended bladder under sterile conditions.

## 19.2 Engineering Technical Specifications

- Material: Medical-grade stainless steel; corrosion-resistant.

- Trocar Configuration: Curved and straight variants; pyramidal sharp-tip obturator.

- Sheath/Cannula Sizes: Compatible with 12 Fr to 22 Fr Foley catheters (typical clinical range).

- Sterilization: Autoclavable at 134 °C.

- Standards: ISO 7153-1, ISO 13485, CE MDR.

## 19.3 Standard Accessories

- Reusable curved suprapubic trocar cannula with sharp pyramidal-tip obturator.

- Reusable straight suprapubic trocar cannula with matching obturator.

- Catheter introduction sleeve and guide channel adapter — accommodating 12—22 Fr catheters.

- Heavy-duty perforated sterilization container with protective locking mechanism.

## 19.4 Installation Requirements

No fixed installation; requires CSSD reprocessing.

## 19.5 Documentation

User & reprocessing manuals, CE/FDA/ISO 13485 certificates.

## 19.6 Training

Clinical handling and reprocessing training.

## 19.7 Warranty

Minimum 2 years against manufacturing defects.